## Enrich1
id | canonical_name | category | subtype | key_attributes_schema | default_iq | default_oq | default_pq | default_calibration_months | default_maintenance_months | default_requalification_months | risk_level | regulatory_references | search_aliases | locale_overrides | notes
| Medical Refrigerator | Cold Chain | 2°C to 8°C Upright | {"temperature_range":{"type":"range","units":"C"},"volume":{"type":"number","units":"L"},"defrost_type":{"type":"enum","values":["Manual","Auto"]}} | TRUE | TRUE | TRUE | 12 | 12 | 12 | Medium | FDA QMSR;ISO 13485;ISO 14971 | fridge;cold storage;chiller;unit storage | {} | PQ typically requires 24-72 hour temperature mapping study.
| Ultra-Low Freezer | Cold Chain | -80°C Cascade Freezer | {"temp_setpoint":{"type":"number","units":"C"},"backup_battery":{"type":"boolean"},"co2_backup":{"type":"boolean"}} | TRUE | TRUE | TRUE | 12 | 6 | 12 | High | FDA QMSR;ISO 13485;ISO 14971;21 CFR 1271 | ULT;deep freeze;minus 80;cryo storage | {} | Critical for HCT/P and biologics. Requires alarm verification in OQ.
| Controlled Rate Freezer | Cold Chain | Liquid Nitrogen (LN2) | {"cooling_rate":{"type":"number","units":"C/min"},"chamber_volume":{"type":"number","units":"L"},"gas_source":{"type":"string"}} | TRUE | TRUE | TRUE | 6 | 6 | 12 | High | 21 CFR 1271;FDA QMSR;ISO 13485 | cryopreservation;cell freezer;LN2 freezer | {} | Used in HCT/P for minimal cell damage during freezing.
| Analytical Balance | Measurement & Metrology | 4-Place Analytical | {"readability":{"type":"number","units":"mg"},"capacity":{"type":"number","units":"g"},"internal_cal":{"type":"boolean"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Medium | USP <1058>;USP <1251>;ISO 17025 | scale;weighing;mass meter;analytical scale | {} | USP <1058> Group B. Daily check required before use.
| Microbalance | Measurement & Metrology | 6-Place Micro | {"readability":{"type":"number","units":"ug"},"min_weight":{"type":"number","units":"mg"},"anti_static":{"type":"boolean"}} | TRUE | TRUE | FALSE | 6 | 6 |  | High | USP <1058>;USP <1251>;ISO 17025 | precision scale;micro scale;fine balance | {} | Highly sensitive to vibration and airflow. OQ includes corner load tests.
| Biosafety Cabinet (BSC) | Environment & Safety | Class II Type A2 | {"inflow_vel":{"type":"number","units":"fpm"},"downflow_vel":{"type":"number","units":"fpm"},"hepa_eff":{"type":"number","units":"%"}} | TRUE | TRUE | TRUE | 12 | 12 | 12 | High | NSF 49;ISO 14644;FDA QMSR | hood;bio-hood;laminar flow;cell culture hood | {} | Requires annual certification for HEPA integrity and airflow velocity.
| Biosafety Cabinet (BSC) | Environment & Safety | Class II Type B2 | {"exhaust_vol":{"type":"number","units":"cfm"},"static_press":{"type":"number","units":"in_wg"},"total_exhaust":{"type":"boolean"}} | TRUE | TRUE | TRUE | 12 | 6 | 12 | High | NSF 49;ISO 14644;FDA QMSR | total exhaust hood;chemical bio hood;B2 BSC | {} | Mandatory for volatile chemicals or radionuclides. Hard-ducted to exhaust.
| Steam Autoclave | Sterilization & Decontamination | Pre-Vacuum Large Batch | {"chamber_vol":{"type":"number","units":"L"},"max_press":{"type":"number","units":"psi"},"max_temp":{"type":"number","units":"C"}} | TRUE | TRUE | TRUE | 6 | 6 | 12 | High | ISO 17665;ISO 11138;ANSI/AAMI ST79 | sterilizer;steam kettle;high-pressure steam | {} | PQ must include biological indicator (BI) kill studies (3 consecutive runs).
| Laboratory Incubator | Analytical Laboratory | CO2/Cell Culture | {"co2_range":{"type":"range","units":"%"},"temp_range":{"type":"range","units":"C"},"humidity":{"type":"boolean"}} | TRUE | TRUE | TRUE | 6 | 6 | 12 | High | FDA QMSR;ISO 13485;USP <1058> | cell incubator;CO2 chamber;culture unit | {} | Critical for cell growth. PQ involves CO2 and temp mapping with load.
| HPLC System | Analytical Laboratory | Quaternary Gradient | {"pump_pressure":{"type":"number","units":"bar"},"flow_precision":{"type":"number","units":"%"},"detector":{"type":"string"}} | TRUE | TRUE | TRUE | 12 | 6 | 12 | Medium | USP <1058>;21 CFR Part 11;ISO 13485 | chromatography;liquid chromatograph;Agilent;Waters | {} | USP <1058> Group C. Requires computerized system validation (CSV).
| Injection Molding Machine | Production: Material Forming | All-Electric Precision | {"clamp_force":{"type":"number","units":"ton"},"shot_size":{"type":"number","units":"oz"},"screw_dia":{"type":"number","units":"mm"}} | TRUE | TRUE | TRUE | 12 | 12 | 24 | High | ISO 13485;21 CFR 820.75;ISO 14971 | molder;press;plastic molder;Engel;Arburg | {} | OQ establishes the "process window" (Melt Temp, Pack Pressure).
| Plastic Extruder | Production: Material Forming | Multi-Lumen Co-Extrusion | {"screw_speed":{"type":"number","units":"rpm"},"barrel_zones":{"type":"integer"},"vacuum_sizing":{"type":"boolean"}} | TRUE | TRUE | TRUE | 12 | 6 | 12 | High | ISO 13485;21 CFR 820.75;ISO 14971 | extrusion line;tube maker;catheter extruder | {} | Critical for catheter tubing dimensions (OD/ID). PQ uses actual resin lot.
| Pouch Sealer | Production: Packaging | Continuous Rotary Band | {"seal_width":{"type":"number","units":"mm"},"temp_setpoint":{"type":"number","units":"C"},"pressure":{"type":"number","units":"N"}} | TRUE | TRUE | TRUE | 12 | 12 | 12 | High | ISO 11607-2;ASTM F88;ASTM F1929 | heat sealer;band sealer;sterile barrier sealer | {} | PQ requires seal strength and dye penetration integrity testing.
| Lyophilizer | Sterilization & Decontamination | Pilot Scale Industrial | {"condenser_cap":{"type":"number","units":"kg"},"shelf_temp":{"type":"range","units":"C"},"vacuum":{"type":"number","units":"mTorr"}} | TRUE | TRUE | TRUE | 6 | 6 | 12 | High | ISO 13485;21 CFR 820.75;FDA QMSR | freeze dryer;sublimation unit;lyo | {} | Critical for stabilizing HCT/Ps or reagents. Cycle validation required.
| Digital Caliper | Measurement & Metrology | Wireless Data Logging | {"accuracy":{"type":"number","units":"mm"},"resolution":{"type":"number","units":"mm"},"range":{"type":"number","units":"mm"}} | TRUE | FALSE | FALSE | 12 |  |  | Medium | ISO 17025;FDA QMSR;ISO 13485 | micrometer;vernier;measuring tool | {} | Calibrated against NIST traceable standards. Data logging aids UDI.
| Particle Counter | Measurement & Metrology | Laser Scattering Air | {"flow_rate":{"type":"number","units":"LPM"},"channels":{"type":"array"},"sensitivity":{"type":"number","units":"um"}} | TRUE | TRUE | FALSE | 12 | 12 |  | High | ISO 14644-1;FDA QMSR;ISO 13485 | aerosol counter;cleanroom tester;particle meter | {} | Essential for cleanroom certification. Recalibration every 12 months.
| PH Meter | Analytical Laboratory | Benchtop Precision | {"ph_range":{"type":"range","units":"pH"},"slope_check":{"type":"boolean"},"temp_comp":{"type":"boolean"}} | TRUE | TRUE | FALSE | 6 | 12 |  | Medium | USP <791>;USP <1058>;ISO 13485 | acidity meter;electrode;pH probe | {} | Requires daily 2-point or 3-point calibration using certified buffers.
| Ultrasonic Welder | Production: Assembly & Joining | Press-Style Benchtop | {"frequency":{"type":"number","units":"kHz"},"trigger_force":{"type":"number","units":"N"},"amplitude":{"type":"number","units":"um"}} | TRUE | TRUE | TRUE | 12 | 6 | 12 | High | ISO 13485;21 CFR 820.75;ISO 14971 | welder;sonic bonder;plastic welder;Branson | {} | PQ evaluates weld strength and visual flash against specifications.
| Walk-In Cold Room | Cold Chain | Temperature Mapped | {"setpoint":{"type":"number","units":"C"},"redundant_cooling":{"type":"boolean"},"alarm_delays":{"type":"number","units":"min"}} | TRUE | TRUE | TRUE | 12 | 6 | 12 | High | FDA QMSR;ISO 13485;ISO 14644 | cold storage room;refrigerated warehouse;WICR | {} | PQ requires mapping with multiple sensors (e.g., 9-15) under empty/full load.
| Peristaltic Pump | Production: Assembly & Filling | Variable Speed Precision | {"flow_rate":{"type":"range","units":"mL/min"},"tubing_compatibility":{"type":"string"}} | TRUE | TRUE | TRUE | 12 | 12 | 12 | Medium | ISO 13485;21 CFR 820.75 | dosing pump;tubing pump;filler pump | {} | PQ confirms volume accuracy and repeatability for vial/pouch filling.
| CNC Vertical Machining Center | Precision Machining | 5-Axis Milling | {"spindle_speed":{"type":"number","units":"RPM"},"tool_positions":{"type":"integer"},"coolant_type":{"type":"enum","values":["Flood","Mist"]}} | TRUE | TRUE | TRUE | 12 | 6 | 24 | High | ISO 13485;21 CFR 820.70;ASTM F2026 | VMC;Haas;Mazak;Mill;Milling Machine | {} | OQ requires ball-bar testing and spindle runout verification.
| CNC Swiss Lathe | Precision Machining | Sliding Headstock | {"max_bar_diameter":{"type":"number","units":"mm"},"sub_spindle":{"type":"boolean"},"guide_bushing":{"type":"boolean"}} | TRUE | TRUE | TRUE | 12 | 6 | 24 | High | ISO 13485;21 CFR 820.70;ISO 14971 | Swiss Turn;Citizen;Star Lathe;Screw Machine | {} | Standard for orthopedic bone screws and dental implants.
| Wire EDM Machine | Precision Machining | Submerged Wire Cut | {"wire_diameter":{"type":"range","units":"mm"},"dielectric_fluid":{"type":"string"},"max_taper_angle":{"type":"number","units":"deg"}} | TRUE | TRUE | TRUE | 12 | 12 |  | High | ISO 13485;21 CFR 820.70 | EDM;Wire Cut;Electrical Discharge Machine | {} | Critical for ultra-fine geometries in surgical instruments.
| Laser Cutting System | Precision Machining | Fiber Laser Stent Cutter | {"laser_power":{"type":"number","units":"W"},"pulse_freq":{"type":"number","units":"kHz"},"gas_assist":{"type":"enum","values":["O2","N2","Ar"]}} | TRUE | TRUE | TRUE | 12 | 6 | 12 | High | ISO 13485;21 CFR 820.75 | Stent cutter;Laser mill;Tube cutter | {} | PQ requires validation of Heat Affected Zone (HAZ) and dimensional stability.
| Waterjet Cutting System | Precision Machining | Ultra-High Pressure CNC | {"pump_pressure":{"type":"number","units":"PSI"},"abrasive_flow":{"type":"number","units":"kg/min"},"nozzle_size":{"type":"number","units":"mm"}} | TRUE | TRUE | TRUE | 12 | 6 |  | Medium | ISO 13485;21 CFR 820.70 | Flow;Water cutter;Abrasive jet | {} | Used for rough shaping of orthopedic blanks.
| Coordinate Measuring Machine (CMM) | Metrology | Bridge-Style CNC | {"measuring_vol":{"type":"string"},"probe_type":{"type":"string"},"accuracy":{"type":"number","units":"um"}} | TRUE | TRUE | TRUE | 12 | 12 | 24 | High | ISO 10360;ISO 17025;FDA QMSR | CMM;Metrology station;Brown & Sharpe;Zeiss | {} | OQ involves verification using certified artifacts (Step Gauges/Spheres).
| Optical Comparator | Metrology | Digital Profile Projector | {"magnification":{"type":"array"},"screen_size":{"type":"number","units":"mm"},"data_processor":{"type":"boolean"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Medium | ISO 17025;FDA QMSR | Shadowgraph;Profile projector;Keyence IM | {} | Calibration verifies stage linear accuracy and magnification distortion.
| Clinical Chemistry Analyzer | IVD Laboratory | Discrete Automated | {"throughput":{"type":"number","units":"tests/hr"},"reagent_cooling":{"type":"boolean"},"sample_vol":{"type":"number","units":"uL"}} | TRUE | TRUE | TRUE | 6 | 6 | 12 | High | CLIA;USP <1058>;ISO 15189 | Chemistry system;Beckman;Siemens;IVD Analyzer | {} | USP <1058> Group C. Requires CSV and linearity validation.
| Hematology Analyzer | IVD Laboratory | 5-Part Differential | {"technology":{"type":"enum","values":["Flow Cytometry","Impedance"]},"parameters_counted":{"type":"integer"}} | TRUE | TRUE | TRUE | 6 | 6 | 12 | High | CLIA;USP <1058>;21 CFR 864 | CBC machine;Blood counter;Coulter counter | {} | Calibration often uses stabilized whole blood controls.
| Blood Gas Analyzer | IVD Laboratory | Cartridge-Based | {"measured_analytes":{"type":"array"},"warmup_time":{"type":"number","units":"sec"}} | TRUE | TRUE | TRUE | 6 | 12 |  | High | CLIA;USP <1058>;21 CFR 862 | ABG machine;POCT analyzer;Electrolyte meter | {} | Critical for surgical support. OQ includes verification of POC connectivity.
| ELISA Plate Reader | IVD Laboratory | Multimode (Abs/Fl/Lum) | {"wavelength_range":{"type":"range","units":"nm"},"shaking_modes":{"type":"array"},"temp_control":{"type":"boolean"}} | TRUE | TRUE | TRUE | 12 | 12 | 12 | Medium | USP <1058>;FDA QMSR | Microplate reader;Spectramax;Tecan;ELISA | {} | OQ verifies photometric accuracy, linearity, and plate alignment.
| ELISA Plate Washer | IVD Laboratory | 96-Well Automated | {"residual_vol":{"type":"number","units":"uL/well"},"wash_heads":{"type":"integer"},"vacuum_press":{"type":"number","units":"bar"}} | TRUE | TRUE | TRUE | 12 | 6 |  | Medium | USP <1058>;FDA QMSR | Plate washer;BioTek;Washer manifold | {} | PQ confirms efficient aspiration to prevent cross-contamination.
| Real-Time PCR Machine | IVD Laboratory | 96-Well qPCR | {"thermal_accuracy":{"type":"number","units":"C"},"optical_channels":{"type":"integer"},"lid_temp":{"type":"number","units":"C"}} | TRUE | TRUE | TRUE | 12 | 12 | 12 | High | USP <1058>;21 CFR Part 11;ISO 13485 | Thermocycler;qPCR;TaqMan;Bio-Rad | {} | Critical for viral/genetic testing. Requires strict temperature uniformity OQ.
| DNA Sequencer | IVD Laboratory | Next-Generation (NGS) | {"flowcell_capacity":{"type":"integer"},"run_time":{"type":"range","units":"hr"},"read_length":{"type":"number","units":"bp"}} | TRUE | TRUE | TRUE | 12 | 6 | 12 | High | USP <1058>;21 CFR Part 11;ISO 20916 | NGS;Illumina;Ion Torrent;Sequencing platform | {} | High software complexity; requires extensive CSV and data integrity validation.
| Flow Cytometer | IVD Laboratory | Multi-Laser Analyzer | {"lasers":{"type":"integer"},"detectors":{"type":"integer"},"max_event_rate":{"type":"number","units":"eps"}} | TRUE | TRUE | TRUE | 12 | 6 | 12 | High | USP <1058>;ISO 13485;21 CFR 1271 | FACS;Cell sorter;BD Facs;Flow machine | {} | Crucial for HCT/P cell viability and purity characterization.
| Automated Liquid Handler | IVD Laboratory | Robotic Pipetting Workcell | {"tip_types":{"type":"array"},"vol_range":{"type":"range","units":"uL"},"deck_capacity":{"type":"integer"}} | TRUE | TRUE | TRUE | 12 | 6 | 12 | High | USP <1058>;21 CFR Part 11 | Hamilton;Robotic pipettor;Liquid robot | {} | PQ must include gravimetric or photometric volume verification.
| Centerless Grinder | Precision Machining | Thru-Feed/In-Feed | {"grinding_wheel_dia":{"type":"number","units":"mm"},"spindle_power":{"type":"number","units":"kW"}} | TRUE | TRUE | TRUE | 12 | 6 |  | Medium | ISO 13485;21 CFR 820.70 | Grinder;Oshino;Precision grinder | {} | Used for guidewire and mandrel production. PQ verifies OD taper/roundness.
| Surface Profilometer | Metrology | Contact Stylus | {"vertical_range":{"type":"number","units":"um"},"stylus_force":{"type":"number","units":"mN"},"resolution":{"type":"number","units":"nm"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Medium | ISO 25178;ISO 17025;FDA QMSR | Ra tester;Roughness meter;Stylus profiler | {} | Critical for verifying implant surface finish (Ra/Rz).
| Universal Testing Machine (UTM) | Metrology | Dual Column Tensile | {"load_cell_cap":{"type":"number","units":"kN"},"speed_range":{"type":"range","units":"mm/min"},"extensometer":{"type":"boolean"}} | TRUE | TRUE | TRUE | 12 | 12 |  | High | ASTM E4;ISO 13485;ASTM F1717 | Instron;Tensile tester;Pull tester;Zwick | {} | Critical for destructive testing of implants/welds.
| Vision Inspection System | Metrology | 3D Profiling Sensor | {"fov_size":{"type":"string"},"depth_resolution":{"type":"number","units":"um"},"frame_rate":{"type":"number","units":"fps"}} | TRUE | TRUE | TRUE | 12 | 12 | 12 | High | ISO 13485;21 CFR Part 11;21 CFR 830 | Vision camera;Cognex;Keyence;Smart sensor | {} | Validation must include "known good" and "known bad" challenge runs.
| Pouch Sealer (Impulse) | Production: Packaging | Impulse Bar Sealer | {"seal_time":{"type":"number","units":"sec"},"seal_temp":{"type":"number","units":"C"},"cooling_time":{"type":"number","units":"sec"}} | TRUE | TRUE | TRUE | 12 | 12 | 12 | High | ISO 11607-2;ASTM F88 | Heat sealer;Bar sealer;Sterile barrier sealer | {} | Critical for sterile barrier integrity. PQ requires seal strength testing.
| Tray Sealer | Production: Packaging | Shuttle-Style Medical | {"platen_size":{"type":"string"},"seal_pressure":{"type":"number","units":"PSI"},"dwell_time":{"type":"number","units":"sec"}} | TRUE | TRUE | TRUE | 12 | 6 | 12 | High | ISO 11607-2;ISO 13485 | Blister sealer;Tray bonder;Alloyd | {} | OQ verifies temperature uniformity across the sealing platen.
| Form-Fill-Seal Machine | Production: Packaging | Horizontal Thermoform (HFFS) | {"forming_depth":{"type":"number","units":"mm"},"cycle_rate":{"type":"number","units":"CPM"},"vacuum_level":{"type":"number","units":"mbar"}} | TRUE | TRUE | TRUE | 12 | 6 | 12 | High | ISO 11607-1;21 CFR 820.75 | Thermoformer;Multivac;Packaging line | {} | PQ includes visual inspection for forming defects and seal integrity.
| Thermal Transfer Printer | Production: Labeling | High-Resolution UDI Printer | {"dpi":{"type":"number","units":"dots/in"},"ribbon_type":{"type":"string"},"print_width":{"type":"number","units":"mm"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Medium | 21 CFR 830;21 CFR Part 11 | Zebra printer;Label maker;Barcode printer | {} | Validation focuses on barcode scanability and UDI compliance.
| Inkjet Coder | Production: Labeling | Continuous Inkjet (CIJ) | {"nozzle_size":{"type":"number","units":"um"},"ink_composition":{"type":"string"},"throw_distance":{"type":"number","units":"mm"}} | TRUE | TRUE | FALSE | 12 | 6 |  | Medium | 21 CFR 830;ISO 13485 | Videojet;Batch coder;Lot printer | {} | OQ includes character clarity and solvent resistance testing.
| Label Applicator | Production: Labeling | Wrap-Around Automated | {"conveyor_speed":{"type":"number","units":"m/min"},"placement_accuracy":{"type":"number","units":"mm"}} | TRUE | TRUE | TRUE | 12 | 12 |  | Medium | 21 CFR 820;ISO 13485 | Labeler;Bottle labeler;Sticker machine | {} | PQ confirms placement consistency and absence of wrinkles/flagging.
| Checkweigher | Production: Inspection | High-Speed Dynamic | {"capacity":{"type":"number","units":"g"},"accuracy":{"type":"number","units":"mg"},"reject_mechanism":{"type":"string"}} | TRUE | TRUE | TRUE | 12 | 6 |  | Medium | NIST HB 44;21 CFR 820 | Weight checker;Inline scale;Rejecter | {} | OQ involves "known weight" challenge at various line speeds.
| HVAC Air Handling Unit | Facilities & Utilities | Cleanroom AHU | {"airflow_volume":{"type":"number","units":"CFM"},"filter_stages":{"type":"integer"},"static_pressure":{"type":"number","units":"in_wg"}} | TRUE | TRUE | TRUE | 12 | 6 | 12 | High | ISO 14644-1;FDA QMSR | Air handler;AC unit;Cleanroom supply | {} | Critical for environmental control (Class 5/7/8). PQ is Cleanroom Cert.
| HEPA Fan Filter Unit (FFU) | Facilities & Utilities | Ceiling Integrated | {"face_velocity":{"type":"number","units":"FPM"},"filter_efficiency":{"type":"number","units":"%"}} | TRUE | TRUE | FALSE | 12 | 12 | 12 | High | ISO 14644-3;NSF 49 | HEPA filter;Fan unit;Ceiling filter | {} | OQ involves filter integrity (PAO) leak testing.
| Compressed Air System | Facilities & Utilities | Oil-Free Scroll | {"max_pressure":{"type":"number","units":"PSI"},"dew_point":{"type":"number","units":"C"},"particulate_rating":{"type":"string"}} | TRUE | TRUE | TRUE | 12 | 6 | 12 | High | ISO 8573-1;FDA QMSR | Process air;CDA;Air compressor | {} | Validation requires air quality testing (Oil, Water, Particles).
| Pure Steam Generator | Facilities & Utilities | Distillation Column | {"steam_capacity":{"type":"number","units":"kg/hr"},"conductivity":{"type":"number","units":"uS/cm"}} | TRUE | TRUE | TRUE | 12 | 6 | 12 | High | ISO 17665;USP <1231> | Clean steam;PSG;Sterile steam | {} | Critical for autoclave supply. Quality must meet WFI standards.
| Pass-Through Chamber | Facilities & Utilities | Active HEPA Interlock | {"interlock_logic":{"type":"string"},"air_change_rate":{"type":"number","units":"ACH"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Medium | ISO 14644;FDA QMSR | Transfer hatch;Pass-box;Cleanroom hatch | {} | OQ verifies interlock timing and pressure differential recovery.
| Microbiological Air Sampler | Environment & Safety | Impacting Portable | {"sampling_rate":{"type":"number","units":"LPM"},"timer_accuracy":{"type":"number","units":"%"}} | TRUE | TRUE | FALSE | 12 | 12 |  | High | ISO 14698;FDA QMSR | Bio-sampler;Air monitor;Colony counter | {} | Essential for viable particle monitoring in aseptic areas.
| Water for Injection (WFI) System | Facilities & Utilities | Vapor Compression Still | {"output_flow":{"type":"number","units":"GPM"},"toc_level":{"type":"number","units":"ppb"},"endotoxin_spec":{"type":"number","units":"EU/mL"}} | TRUE | TRUE | TRUE | 6 | 6 | 12 | High | USP <1231>;ISO 13485 | WFI;Sterile water;VC Still | {} | Requires Phase I/II/III validation (up to 1 year for Phase III).
| Ethylene Oxide (EtO) Chamber | Sterilization & Decontamination | Industrial Contract | {"chamber_vol":{"type":"number","units":"m3"},"eto_concentration":{"type":"number","units":"mg/L"},"humidity_setpoint":{"type":"number","units":"%"}} | TRUE | TRUE | TRUE | 6 | 6 | 12 | High | ISO 11135;21 CFR 820.75 | Gas sterilizer;EtO room;Terminal sterilizer | {} | PQ requires 3 full cycles and 1-3 fractional/half-cycles.
| Gamma Irradiator | Sterilization & Decontamination | Continuous Tote Carrier | {"source_activity":{"type":"number","units":"kCi"},"dose_range":{"type":"range","units":"kGy"},"conveyor_speed":{"type":"number","units":"m/hr"}} | TRUE | TRUE | TRUE | 12 | 12 | 12 | High | ISO 11137-1;ISO 11137-2 | Radiation plant;Cobalt-60;Gamma plant | {} | PQ focuses on dose mapping and minimum/maximum dose ratios.
| VHP Generator | Sterilization & Decontamination | Mobile Decon Unit | {"h2o2_injection_rate":{"type":"number","units":"g/min"},"airflow_rate":{"type":"number","units":"m3/hr"}} | TRUE | TRUE | TRUE | 12 | 6 | 12 | High | ISO 14937;ISO 14644 | Hydrogen peroxide fogger;VHP unit;Steris | {} | Used for cleanroom and isolator decontamination. Cycle validation required.
| Micro-Molding Machine | Production: Material Forming | High-Resolution Desktop | {"clamp_force":{"type":"number","units":"kN"},"injection_velocity":{"type":"number","units":"mm/s"},"shot_resolution":{"type":"number","units":"mg"}} | TRUE | TRUE | TRUE | 12 | 12 | 24 | High | ISO 13485;21 CFR 820.75 | Mini-molder;Micro-injection;Small part press | {} | Used for tiny implant components. OQ verifies gate freeze-off timing.
| Electronic Pipette | Measurement & Metrology | Programmed Variable Vol | {"vol_range":{"type":"range","units":"uL"},"motor_accuracy":{"type":"number","units":"%"}} | TRUE | TRUE | FALSE | 6 | 12 |  | Low | ISO 8655;USP <1058> | Auto-pipette;Digital pipette;Eppendorf | {} | USP <1058> Group B. Calibrated gravimetrically.
| Ultrasonic Cleaner | Sterilization & Decontamination | Multi-Tank Aqueous | {"frequency":{"type":"number","units":"kHz"},"bath_temp":{"type":"number","units":"C"},"degas_time":{"type":"number","units":"min"}} | TRUE | TRUE | TRUE | 12 | 12 | 12 | Medium | ISO 15883;AAMI ST98 | Sonic bath;Parts washer;Branson bath | {} | OQ verifies ultrasonic activity uniformity using foil tests.
| Plasma Spray System | Production: Surface Treatment | Hydroxyapatite (HA) Coating | {"plate_power":{"type":"number","units":"kW"},"spray_distance":{"type":"number","units":"mm"},"powder_feed_rate":{"type":"number","units":"g/min"}} | TRUE | TRUE | TRUE | 12 | 6 | 12 | High | ISO 13779-2;ASTM F1147;FDA QMSR | HA sprayer;implant coater;thermal spray | {} | PQ must include HA crystallinity and adhesion strength testing.
| Hydrophilic Coating Line | Production: Surface Treatment | UV-Curable Catheter Coating | {"uv_dosage":{"type":"number","units":"J/cm2"},"dip_speed":{"type":"number","units":"mm/s"},"viscosity":{"type":"number","units":"cP"}} | TRUE | TRUE | TRUE | 6 | 6 | 12 | High | ISO 13485;ASTM F2397;USP 788 | lubritious coating;UV coater;dip tank | {} | Critical for catheter lubricity. OQ verifies UV intensity uniformity.
| Balloon Forming Machine | Production: Material Forming | Cardiovascular Blow-Molder | {"forming_pressure":{"type":"number","units":"psi"},"mold_temp":{"type":"number","units":"C"},"stretch_speed":{"type":"number","units":"mm/s"}} | TRUE | TRUE | TRUE | 6 | 6 | 12 | High | ISO 13485;FDA QMSR;ISO 10555 | balloon molder;PTCA former;stent balloon | {} | Critical for balloon wall thickness and burst pressure.
| Stent Crimping Machine | Production: Assembly | Automated Heart Valve/Stent | {"crimping_force":{"type":"number","units":"N"},"target_diameter":{"type":"number","units":"mm"},"dwell_time":{"type":"number","units":"sec"}} | TRUE | TRUE | TRUE | 6 | 12 | 12 | High | ISO 25539-2;FDA QMSR | stent loader;crimper;valve crimper | {} | OQ verifies diameter repeatability and radial force.
| Guidewire Coiling Machine | Production: Assembly | Precision Mandrel Winder | {"winding_tension":{"type":"number","units":"N"},"pitch_accuracy":{"type":"number","units":"mm"},"wire_gauge":{"type":"range","units":"AWG"}} | TRUE | TRUE | TRUE | 12 | 6 | 24 | High | ISO 13485;21 CFR 820.70 | coil winder;spring maker;guidewire winder | {} | PQ verifies coil integrity and flexibility (kink resistance).
| RF Induction Heater | Production: Assembly | Catheter Tipping & Bonding | {"rf_output":{"type":"number","units":"W"},"heat_time":{"type":"number","units":"sec"},"die_temp":{"type":"number","units":"C"}} | TRUE | TRUE | TRUE | 6 | 12 | 12 | High | ISO 13485;21 CFR 820.75 | RF welder;tip former;induction bonder | {} | Critical for atraumatic catheter tip formation.
| Centerless Grinder | Production: Precision Machining | Mandrel/Guidewire Finisher | {"wheel_speed":{"type":"number","units":"RPM"},"coolant_flow":{"type":"number","units":"L/min"},"taper_control":{"type":"boolean"}} | TRUE | TRUE | TRUE | 12 | 6 |  | Medium | ISO 13485;21 CFR 820.70 | guidewire grinder;mandrel finisher;OD grinder | {} | OQ verifies OD tolerance and surface finish (Ra).
| Bone Saw / Cutter | HCT/P Processing | Precision Allograft Saw | {"blade_speed":{"type":"number","units":"RPM"},"irrigation_flow":{"type":"number","units":"mL/min"},"depth_stop":{"type":"boolean"}} | TRUE | TRUE | TRUE | 12 | 12 | 12 | High | 21 CFR 1271.200;ISO 13485 | allograft saw;bone cutter;HCTP saw | {} | PQ verifies minimal manipulation and heat prevention in tissue.
| Tendon Stripper | HCT/P Processing | Mechanical Tissue Preparer | {"stripping_force":{"type":"number","units":"N"},"blade_gap":{"type":"number","units":"mm"},"sterile_housing":{"type":"boolean"}} | TRUE | TRUE | FALSE | 12 | 12 |  | High | 21 CFR 1271.200;ISO 13485 | ligament stripper;HCTP tool;tendon shaver | {} | Used for donor ligament preparation. Requires sterile contact check.
| Heart Valve Dissection Bench | HCT/P Processing | Aseptic Surgical Station | {"surface_material":{"type":"string"},"lighting_intensity":{"type":"number","units":"lux"},"ergonomics":{"type":"string"}} | TRUE | FALSE | FALSE |  | 12 |  | High | 21 CFR 1271.190;ISO 14644 | valve bench;HCTP dissection;tissue bench | {} | IQ verifies sterile surface finish and cleanroom integration.
| Tissue Dehydration System | HCT/P Processing | Automated Solvent/Temp | {"vacuum_level":{"type":"number","units":"mbar"},"heating_temp":{"type":"number","units":"C"},"solvent_conc":{"type":"number","units":"%"}} | TRUE | TRUE | TRUE | 12 | 6 | 12 | High | 21 CFR 1271.200;ISO 13485 | tissue processor;HCTP dryer;bone dehydrator | {} | PQ verifies residual moisture and tissue integrity.
| Cryogenic Vial Filling Line | Production: Packaging | Aseptic Biologics Filler | {"fill_accuracy":{"type":"number","units":"%"},"vial_size":{"type":"range","units":"mL"},"capping_torque":{"type":"number","units":"N-cm"}} | TRUE | TRUE | TRUE | 6 | 6 | 12 | High | ISO 13408-1;21 CFR 1271 | bio-filler;cryo filler;HCTP filling | {} | PQ includes media fills to verify aseptic capability.
| Automated Pipetting System | IVD Laboratory | Robotic Liquid Handler | {"vol_accuracy":{"type":"number","units":"%"},"cv_reproducibility":{"type":"number","units":"%"},"deck_positions":{"type":"integer"}} | TRUE | TRUE | TRUE | 6 | 6 | 12 | High | USP <1058>;21 CFR Part 11 | liquid handler;Hamilton;Tecan;pipettor | {} | PQ uses gravimetric or photometric verification of volumes.
| Microplate Washer | IVD Laboratory | 96/384-Well Automated | {"residual_vol":{"type":"number","units":"uL"},"aspirate_press":{"type":"number","units":"bar"},"manifold_type":{"type":"string"}} | TRUE | TRUE | TRUE | 12 | 12 |  | Medium | USP <1058>;FDA QMSR | plate washer;ELISA washer;BioTek | {} | OQ verifies aspiration efficiency to prevent cross-talk.
| UV-C Surface Sterilizer | Environment & Safety | Mobile Decon/Robotic | {"uv_intensity":{"type":"number","units":"uW/cm2"},"exposure_time":{"type":"number","units":"sec"},"safety_interlock":{"type":"boolean"}} | TRUE | TRUE | TRUE | 12 | 6 | 12 | High | ISO 14937;FDA QMSR | UV bot;UV-C light;room sterilizer | {} | PQ uses biological indicators (BIs) to verify log reduction.
| Contact Angle Goniometer | Metrology | Surface Hydrophilicity Tester | {"angle_range":{"type":"range","units":"deg"},"droplet_vol":{"type":"number","units":"uL"},"resolution":{"type":"number","units":"deg"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Medium | ISO 17025;FDA QMSR | contact angle meter;surface energy tester | {} | Essential for validating hydrophilic/plasma coatings.
| Catheter Lubricity Tester | Metrology | Pinch-Test Friction Meter | {"clamp_force":{"type":"number","units":"N"},"pull_speed":{"type":"number","units":"mm/s"},"force_res":{"type":"number","units":"mN"}} | TRUE | TRUE | TRUE | 12 | 12 |  | Medium | ISO 13485;ASTM F2397 | pinch tester;friction tester;catheter track | {} | PQ verifies "Mean Friction" reduction against specifications.
| Burst Pressure Tester | Metrology | Hydraulic Balloon Tester | {"max_pressure":{"type":"number","units":"psi"},"fluid_temp":{"type":"number","units":"C"},"ramp_rate":{"type":"number","units":"psi/s"}} | TRUE | TRUE | TRUE | 6 | 12 |  | High | ISO 10555-1;FDA QMSR | balloon burster;pressure tester;RBP tester | {} | Used to determine Rated Burst Pressure (RBP) for catheters.
| Torsional Yield Tester | Metrology | Bone Screw Torque Station | {"torque_cap":{"type":"number","units":"N-m"},"angular_res":{"type":"number","units":"deg"},"axial_load":{"type":"number","units":"N"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Medium | ASTM F543;ISO 17025 | screw tester;torque yield;ortho tester | {} | Critical for validating mechanical strength of orthopedic screws.
| Laser Micrometer | Metrology | Dual-Axis Dimensional Scan | {"od_range":{"type":"range","units":"mm"},"linearity":{"type":"number","units":"um"},"scan_rate":{"type":"number","units":"Hz"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Medium | ISO 17025;FDA QMSR | OD gauge;laser mic;Zumbach;Keyence | {} | Used for in-line extrusion or final catheter OD verification.
| Vision Inspection System | Metrology | High-Speed Batch Inspector | {"camera_res":{"type":"number","units":"MP"},"frame_rate":{"type":"number","units":"fps"},"lighting_type":{"type":"string"}} | TRUE | TRUE | TRUE | 12 | 12 | 12 | High | ISO 13485;21 CFR 820.70;21 CFR Part 11 | vision bot;Cognex;Keyence;optical inspector | {} | OQ requires "Red/Green" (Accept/Reject) challenge testing.
| Scanning Electron Microscope (SEM) | Metrology | Surface Morphology SEM | {"magnification":{"type":"range","units":"x"},"resolution":{"type":"number","units":"nm"},"eds_integrated":{"type":"boolean"}} | TRUE | TRUE | FALSE | 12 | 6 |  | Medium | ISO 13485;FDA QMSR | SEM;electron scope;surface analyzer | {} | Used for analyzing implant surface finish and failure modes.
| FTIR Spectrometer | Analytical Lab | Universal ATR FTIR | {"spectral_range":{"type":"range","units":"cm-1"},"resolution":{"type":"number","units":"cm-1"},"library_search":{"type":"boolean"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Medium | USP <854>;USP <1058>;ISO 13485 | FTIR;material ID;infrared scope | {} | USP <1058> Group C. Critical for raw material identification.
| GC/MS System | Analytical Lab | Triple Quadrupole GC/MS | {"mass_range":{"type":"range","units":"m/z"},"autosampler_cap":{"type":"integer"},"carrier_gas":{"type":"string"}} | TRUE | TRUE | TRUE | 12 | 6 | 12 | High | USP <1058>;21 CFR Part 11;ISO 10993-18 | GCMS;residual analyzer;EO residual tester | {} | Used for ethylene oxide (EO) residual and extractable testing.
| Jacketed Mixing Tank | Reagent Mfg | Reagent Buffer Prep | {"volume":{"type":"number","units":"L"},"temp_range":{"type":"range","units":"C"},"agitation_speed":{"type":"number","units":"RPM"}} | TRUE | TRUE | TRUE | 12 | 12 | 12 | Medium | 21 CFR 820.75;ASME BPVC | buffer tank;mixing vessel;stainless tank | {} | OQ verifies temperature control stability and mixing uniformity.
| High-Shear Mixer | Reagent Mfg | Emulsion Homogenizer | {"rotor_speed":{"type":"number","units":"RPM"},"shear_rate":{"type":"number","units":"s-1"},"sterile_design":{"type":"boolean"}} | TRUE | TRUE | TRUE | 6 | 6 | 12 | Medium | 21 CFR 820.75;FDA QMSR | homogenizer;Silverson;reagent mixer | {} | PQ verifies particle size distribution consistency across batches.
| Rocking Bioreactor | HCT/P Processing | Wave-Motion Single-Use | {"rock_angle":{"type":"number","units":"deg"},"rock_rate":{"type":"number","units":"rpm"},"do2_control":{"type":"boolean"}} | TRUE | TRUE | TRUE | 6 | 6 | 12 | High | 21 CFR 1271;FDA QMSR | wave bioreactor;cell culture bag;HCTP culture | {} | Used for expanding human cells. PQ focuses on cell viability.
| Fluorescence Cell Sorter | HCT/P Processing | Multi-Laser FACS | {"lasers":{"type":"integer"},"sort_purity":{"type":"number","units":"%"},"nozzle_size":{"type":"number","units":"um"}} | TRUE | TRUE | TRUE | 12 | 6 | 12 | High | 21 CFR 1271;USP <1058> | FACS;cell sorter;BD Melody | {} | OQ verifies laser alignment and sorting accuracy.
| Collaborative Robot (Cobot) | Assembly | 6-Axis Cleanroom Cobot | {"payload":{"type":"number","units":"kg"},"reach":{"type":"number","units":"mm"},"force_limiting":{"type":"boolean"}} | TRUE | TRUE | TRUE | 12 | 6 | 24 | Medium | ISO 10218;IEC 60601;FDA QMSR | UR5;Cobot;human-robot assembly | {} | OQ focuses on force-stop safety limits and E-stop logic.
| Modular Conveyor System | Assembly | Precision Pallet Transfer | {"belt_speed":{"type":"number","units":"m/min"},"index_precision":{"type":"number","units":"mm"},"cleanroom_class":{"type":"integer"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Low | ISO 14644;FDA QMSR | conveyor;pallet line;flex link | {} | IQ verifies cleanroom material compatibility (no outgassing).
| Particle Monitoring System | Facilities | Continuous Cleanroom Air | {"sensitivity":{"type":"number","units":"um"},"sample_freq":{"type":"number","units":"sec"},"alarm_logic":{"type":"string"}} | TRUE | TRUE | FALSE | 12 | 12 | 12 | High | ISO 14644-2;FDA QMSR | FMS;particle counter;cleanroom monitor | {} | Continuous monitoring for aseptic areas. OQ verifies alarm delays.
| Ducted Fume Hood | Facilities | Chemical Exhaust Station | {"face_velocity":{"type":"number","units":"fpm"},"sash_interlock":{"type":"boolean"},"low_flow_alarm":{"type":"boolean"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Medium | ANSI/ASHRAE 110;ISO 14644 | chemical hood;lab exhaust;fume station | {} | OQ involves smoke testing and face velocity mapping.
| Nitrogen Generator | Utilities | High-Purity PSA | {"output_flow":{"type":"number","units":"SCFH"},"purity_level":{"type":"number","units":"%"},"dew_point":{"type":"number","units":"C"}} | TRUE | TRUE | TRUE | 12 | 6 | 12 | Medium | ISO 8573-1;FDA QMSR | N2 gen;nitrogen maker;process gas | {} | PQ verifies gas purity over a 24-hour peak demand period.
| Total Organic Carbon (TOC) | Utilities | In-Line Water TOC Analyzer | {"lod":{"type":"number","units":"ppb"},"response_time":{"type":"number","units":"sec"},"oxidizer_type":{"type":"string"}} | TRUE | TRUE | TRUE | 6 | 12 | 12 | High | USP <643>;USP <1231> | TOC meter;water quality tester;Sievers | {} | Critical for WFI system monitoring. Calibrated with NIST standards.
| Vacuum Drying Oven | Production | Moisture Removal Unit | {"vacuum_level":{"type":"number","units":"mbar"},"temp_uniformity":{"type":"number","units":"C"},"shelf_count":{"type":"integer"}} | TRUE | TRUE | TRUE | 12 | 12 | 12 | Medium | 21 CFR 820.75;ISO 13485 | vacuum drier;parts oven;de-moisturizer | {} | Used for removing residuals from multi-lumen catheters.
| Automated Vial Washer | Cleaning | Rotary Reagent Bottle Washer | {"washes_per_hour":{"type":"integer"},"rinse_cycles":{"type":"integer"},"wfi_final_rinse":{"type":"boolean"}} | TRUE | TRUE | TRUE | 12 | 6 | 12 | Medium | ISO 15883;AAMI ST98 | bottle washer;vial cleaner;glass washer | {} | PQ requires riboflavin (UV dye) coverage/cleaning validation.
| LAL Kinetic Reader | Metrology | Endotoxin Plate Reader | {"wavelength":{"type":"number","units":"nm"},"cv_precision":{"type":"number","units":"%"},"software_v":{"type":"string"}} | TRUE | TRUE | TRUE | 12 | 12 |  | High | USP <85>;ISO 13485 | endotoxin reader;pyrogen tester;LAL station | {} | Critical for safety release of implantables and HCT/Ps.
| Vickers Hardness Tester | Metrology | Automated Micro-Hardness | {"load_range":{"type":"range","units":"gf"},"dwell_time":{"type":"number","units":"sec"},"accuracy":{"type":"number","units":"%"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Medium | ASTM E384;ISO 13485 | microhardness;metal tester;implant hardness | {} | Calibration uses certified hardness blocks (NIST traceable).
| Joint Wear Simulator | Metrology | Hip/Knee Durability Station | {"axial_load":{"type":"number","units":"N"},"cycle_count":{"type":"integer"},"stations":{"type":"integer"}} | TRUE | TRUE | TRUE | 12 | 6 | 24 | High | ISO 14242;ASTM F1717 | wear tester;hip simulator;joint machine | {} | PQ verifies long-term durability (>5 million cycles) of implants.
| In-Line Conductivity Meter | Utilities | Sanitary Water Meter | {"cell_constant":{"type":"number"},"temp_comp":{"type":"boolean"},"accuracy":{"type":"number","units":"uS/cm"}} | TRUE | TRUE | TRUE | 6 | 12 | 12 | High | USP <645>;USP <1231> | water cond;DI monitor;pure water meter | {} | Essential for PW/WFI loop monitoring. USP <645> Stage 1.
| Heat Staking Machine | Production: Assembly | Thermal Press Station | {"forming_temp":{"type":"number","units":"C"},"forming_force":{"type":"number","units":"N"},"dwell_time":{"type":"number","units":"sec"}} | TRUE | TRUE | TRUE | 12 | 12 | 12 | High | FDA QMSR;ISO 13485 | hot staking;plastic welder;thermal joiner | {} | PQ verifies joint pull-strength and material embrittlement.
| Plastic Resin Dryer | Production: Auxiliaries | Desiccant Hopper Dryer | {"dew_point":{"type":"number","units":"C"},"drying_temp":{"type":"number","units":"C"},"airflow":{"type":"number","units":"CFM"}} | TRUE | TRUE | TRUE | 12 | 6 | 24 | Medium | 21 CFR 820.70;ISO 13485 | hopper dryer;resin dryer;material prep | {} | Critical for moisture-sensitive resins (e.g., PEEK, Nylon).
| Mold Temp Controller | Production: Auxiliaries | Water/Oil Thermolator | {"max_temp":{"type":"number","units":"C"},"flow_rate":{"type":"number","units":"GPM"},"heater_cap":{"type":"number","units":"kW"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Medium | 21 CFR 820.70;ISO 13485 | thermolator;mold heater;chiller unit | {} | OQ verifies temperature stability at the mold inlet/outlet.
| Hot Runner Controller | Production: Auxiliaries | Multi-Zone Electrical | {"zones":{"type":"integer"},"power_output":{"type":"number","units":"W"},"alarm_logic":{"type":"string"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Medium | 21 CFR 820.70;ISO 13485 | manifold controller;mold runner;temp control | {} | OQ verifies individual zone response and thermocouple accuracy.
| ICP-MS Analyzer | Analytical Lab | Elemental Impurity System | {"mass_range":{"type":"range","units":"m/z"},"nebulizer_type":{"type":"string"},"lod":{"type":"number","units":"ppt"}} | TRUE | TRUE | TRUE | 12 | 6 | 12 | High | USP <232>;USP <233>;ISO 10993-18 | heavy metal tester;plasma mass spec;elemental analyzer | {} | Used for safety release (leachable metals) of implants.
| Laser Marker (Fiber) | Production: Labeling | Fiber Laser UDI System | {"wavelength":{"type":"number","units":"nm"},"pulse_freq":{"type":"number","units":"kHz"},"marking_speed":{"type":"number","units":"mm/s"}} | TRUE | TRUE | TRUE | 12 | 12 |  | Medium | 21 CFR 830;21 CFR Part 11 | UDI marker;fiber laser;metal engraver | {} | PQ verifies barcode grade (ISO/IEC 15415) on actual substrate.
| Laser Marker (CO2) | Production: Labeling | CO2 Non-Metal System | {"laser_power":{"type":"number","units":"W"},"spot_size":{"type":"number","units":"um"},"focal_length":{"type":"number","units":"mm"}} | TRUE | TRUE | TRUE | 12 | 12 |  | Medium | 21 CFR 830;21 CFR Part 11 | polymer marker;CO2 engraver;label laser | {} | Used for marking catheter shafts or plastic housings.
| DSC Analyzer | Analytical Lab | Differential Scanning Calorimeter | {"temp_range":{"type":"range","units":"C"},"heating_rate":{"type":"number","units":"C/min"},"atmosphere":{"type":"string"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Medium | USP <1058>;ISO 11357 | thermal analyzer;melt point;polymer ID | {} | Critical for verifying resin crystallinity and glass transition.
| TGA Analyzer | Analytical Lab | Thermogravimetric Analyzer | {"balance_res":{"type":"number","units":"ug"},"temp_max":{"type":"number","units":"C"},"sample_pan":{"type":"string"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Medium | USP <1058>;ISO 11358 | thermal stability;ash content;composition analyzer | {} | Used to detect filler content or thermal degradation in polymers.
| Automated Passivation Line | Production: Finishing | Nitric/Citric Acid Line | {"acid_conc":{"type":"number","units":"%"},"bath_temp":{"type":"number","units":"C"},"dwell_time":{"type":"number","units":"min"}} | TRUE | TRUE | TRUE | 6 | 6 | 12 | High | ASTM A967;ISO 13485 | acid bath;stainless finisher;passivator | {} | PQ must include corrosion resistance verification (e.g., salt spray).
| Electropolishing System | Production: Finishing | Anodic Chemical Bath | {"current_density":{"type":"number","units":"A/ft2"},"voltage":{"type":"number","units":"V"},"electrolyte_type":{"type":"string"}} | TRUE | TRUE | TRUE | 6 | 6 | 12 | High | ISO 13485;21 CFR 820.75 | EP system;anodic polisher;stent finisher | {} | Critical for smoothing surfaces of interventional metal devices.
| Cold Chain Monitor (IoT) | Distribution | Real-time Datalogger | {"temp_accuracy":{"type":"number","units":"C"},"ping_freq":{"type":"number","units":"min"},"battery_life":{"type":"integer","units":"months"}} | TRUE | TRUE | FALSE | 12 |  |  | Medium | 21 CFR Part 11;ISO 13485 | GPS logger;temp tracker;cellular monitor | {} | Used for validating HCT/P shipping lanes. Software must be CSV.
| Automated Bead Blaster | Production: Finishing | CNC Cabinet Blaster | {"nozzle_press":{"type":"number","units":"PSI"},"media_grit":{"type":"integer"},"cycle_time":{"type":"number","units":"sec"}} | TRUE | TRUE | TRUE | 12 | 6 |  | Medium | ISO 13485;21 CFR 820.70 | sand blaster;shot peener;surface prep | {} | OQ verifies surface roughness (Ra) uniformity on implants.
| Capillary Rheometer | Analytical Lab | Polymer Flow Tester | {"max_force":{"type":"number","units":"kN"},"shear_rate":{"type":"range","units":"s-1"},"die_geometry":{"type":"string"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Medium | ISO 11443;ASTM D3835 | flow tester;viscometer;extrusion rheometer | {} | Used to validate resin batch processability for extruders.
| Air Shower | Facilities | Cleanroom Entry Unit | {"air_velocity":{"type":"number","units":"m/s"},"cycle_time":{"type":"number","units":"sec"},"interlock_type":{"type":"string"}} | TRUE | TRUE | FALSE | 12 | 12 | 12 | Medium | ISO 14644-1;FDA QMSR | entry blower;de-dusting;cleanroom shower | {} | OQ verifies particulate removal efficiency and interlock logic.
| Solvent Vapor Degreaser | Cleaning | Multi-Stage Ultrasonic | {"vapor_temp":{"type":"number","units":"C"},"solvent_level":{"type":"number","units":"mm"},"boil_sump_temp":{"type":"number","units":"C"}} | TRUE | TRUE | TRUE | 12 | 6 | 12 | Medium | 21 CFR 820.70;ISO 13485 | vapor cleaner;parts degreaser;Branson cleaner | {} | PQ verifies residual solvent levels (NVR) after cleaning.
| Glove Integrity Tester | Environment & Safety | Isolator Pressure Decay | {"test_pressure":{"type":"number","units":"Pa"},"leak_rate":{"type":"number","units":"Pa/min"},"cycle_time":{"type":"number","units":"sec"}} | TRUE | TRUE | FALSE | 12 | 12 |  | High | ISO 14644-7;FDA QMSR | glove tester;barrier tester;isolator tool | {} | Critical for aseptic HCT/P processing. Requires daily or batch check.
| Adhesion Pull Tester | Metrology | Coating Bond Strength | {"max_load":{"type":"number","units":"N"},"pull_rate":{"type":"number","units":"mm/s"},"fixture_type":{"type":"string"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Medium | ASTM F1147;ISO 13779 | pull-off tester;coating tester;bond strength | {} | Used to validate HA or plasma-sprayed implant coatings.
| Micro-TIG Welder | Production: Assembly | Precision Arc Welder | {"current_range":{"type":"range","units":"A"},"pulse_duration":{"type":"number","units":"ms"},"shield_gas":{"type":"string"}} | TRUE | TRUE | TRUE | 12 | 12 | 12 | High | ISO 13485;21 CFR 820.75 | pulse arc;micro welder;surgical tool welder | {} | Used for surgical instruments. PQ verifies weld penetration/integrity.
| Nitrogen Flush System | Production: Packaging | Bag/Vial Gas Overlay | {"flow_rate":{"type":"number","units":"LPM"},"oxygen_residual":{"type":"number","units":"%"},"nozzle_vac":{"type":"boolean"}} | TRUE | TRUE | TRUE | 12 | 12 | 12 | Medium | 21 CFR 820.75;ISO 11607 | N2 flush;gas packer;O2 scavenger | {} | PQ verifies residual oxygen levels in the headspace of the package.
| SMT Reflow Oven | Production: Electronics | Multi-Zone Convection | {"heating_zones":{"type":"integer"},"belt_speed":{"type":"number","units":"mm/min"},"max_temp":{"type":"number","units":"C"}} | TRUE | TRUE | TRUE | 12 | 6 | 12 | High | IPC-7801;IPC-7530;J-STD-020 | solder oven;reflow station;SMT oven;PCB heater | {} | PQ requires thermal profiling using a data logger (e.g., KIC) on a test board.
| Solder Paste Printer | Production: Electronics | Automated Screen Printer | {"alignment_accuracy":{"type":"number","units":"um"},"squeegee_pressure":{"type":"number","units":"kg"},"cycle_time":{"type":"number","units":"sec"}} | TRUE | TRUE | TRUE | 12 | 12 |  | Medium | IPC-7525;ISO 13485 | stencil printer;paste applier;SMT printer | {} | OQ verifies aperture alignment and paste height consistency.
| Automated Optical Inspector (AOI) | Measurement & Metrology | Post-Reflow SMT Inspector | {"camera_res":{"type":"number","units":"um/pixel"},"inspection_speed":{"type":"number","units":"cm2/sec"},"fov":{"type":"string"}} | TRUE | TRUE | TRUE | 12 | 12 | 12 | High | IPC-A-610;21 CFR Part 11 | AOI;PCB inspector;solder joint checker | {} | OQ uses "Golden Boards" to verify false call rates and defect detection.
| Wave Soldering Machine | Production: Electronics | Lead-Free Dual Wave | {"solder_pot_temp":{"type":"number","units":"C"},"conveyor_angle":{"type":"number","units":"deg"},"fluxer_type":{"type":"string"}} | TRUE | TRUE | TRUE | 12 | 6 | 12 | High | J-STD-001;ISO 13485 | wave solder;thru-hole solder;selective wave | {} | Critical for PCBA integrity in life-support devices. PQ focuses on dross levels and wetting.
| Raman Spectrometer | Analytical Lab | Handheld RMID System | {"laser_wavelength":{"type":"number","units":"nm"},"spectral_range":{"type":"range","units":"cm-1"},"laser_safety_class":{"type":"string"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Medium | USP <858>;USP <1058>;21 CFR Part 11 | Raman;RMID;chemical ID;Vaya;Agilent | {} | USP <1058> Group C. Used for raw material ID through opaque containers.
| Lateral Flow Reagent Dispenser | Production: IVD | Nano-Dispense System | {"dispense_vol":{"type":"range","units":"nL"},"line_width":{"type":"number","units":"mm"},"position_accuracy":{"type":"number","units":"um"}} | TRUE | TRUE | TRUE | 6 | 6 | 12 | High | ISO 13485;FDA QMSR | BioDot;LFA striper;reagent dispenser | {} | PQ verifies protein binding consistency on nitrocellulose membranes.
| Guillotine Card Cutter | Production: IVD | Automated LFA Strip Cutter | {"cut_width":{"type":"number","units":"mm"},"strips_per_minute":{"type":"integer"},"blade_life":{"type":"integer","units":"cuts"}} | TRUE | TRUE | TRUE | 12 | 12 |  | Medium | ISO 13485;FDA QMSR | strip cutter;LFA slitter;card cutter | {} | OQ verifies strip width tolerance to ensure proper fit in plastic cassettes.
| Automated Lamination System | Production: IVD | Multi-Layer LFA Laminator | {"lamination_pressure":{"type":"number","units":"bar"},"alignment_sensor":{"type":"boolean"},"throughput":{"type":"number","units":"cards/hr"}} | TRUE | TRUE | TRUE | 12 | 12 |  | Medium | ISO 13485;FDA QMSR | LFA laminator;card assembler;membrane sticker | {} | Critical for consistent wicking between sample pad and membrane.
| Bio-Hazard Waste Autoclave | Sterilization & Decon | Industrial Decontamination | {"chamber_vol":{"type":"number","units":"L"},"max_temp":{"type":"number","units":"C"},"cycle_record":{"type":"boolean"}} | TRUE | TRUE | TRUE | 6 | 6 | 12 | High | 29 CFR 1910.1030;ISO 17665 | red bag sterilizer;waste autoclave;decon unit | {} | PQ must demonstrate complete inactivation of waste loads using BIs.
| Industrial Waste Shredder | Facilities & Waste | Sharps/Bio-Waste Shredder | {"motor_power":{"type":"number","units":"HP"},"shred_size":{"type":"number","units":"mm"},"noise_level":{"type":"number","units":"dB"}} | TRUE | TRUE | FALSE |  | 12 |  | Low | 21 CFR 820.70;OSHA | sharps shredder;waste grinder;medical shredder | {} | Used to render medical device prototypes or rejects non-functional and safe.
| Cryogenic Dry Shipper | Cold Chain | LN2 Absorbed Dewar | {"hold_time":{"type":"number","units":"days"},"inner_capacity":{"type":"number","units":"L"},"temp_logger":{"type":"boolean"}} | TRUE | FALSE | TRUE | 12 | 12 |  | High | 21 CFR 1271.265;ISO 13485 | LN2 shipper;HCTP transport;dry dewar | {} | PQ involves "Open-Door" and "Hold-Time" challenges with a dummy load.
| Automated Cell Counter | HCT/P Processing | Image-Based Viability | {"counting_range":{"type":"range","units":"cells/mL"},"viability_dye":{"type":"string"},"image_res":{"type":"number","units":"MP"}} | TRUE | TRUE | TRUE | 6 | 12 |  | High | 21 CFR 1271.160;USP <1058> | Vi-Cell;Countess;cell analyzer;HCTP counter | {} | OQ verifies count linearity against certified bead standards.
| Micro-Laser Welder | Production: Joining | Manual/Robotic Precision | {"spot_dia":{"type":"number","units":"um"},"pulse_energy":{"type":"number","units":"mJ"},"aiming_beam":{"type":"boolean"}} | TRUE | TRUE | TRUE | 12 | 6 | 12 | High | ISO 15609-4;ISO 13919-1 | jewelry welder;laser tacker;micro welder | {} | Used for heart valve frames and tiny electronics. PQ uses pull-testing.
| Benchtop Centrifuge | Laboratory | Multi-Rotor General | {"max_rpm":{"type":"number"},"max_g_force":{"type":"number"},"timer_range":{"type":"range","units":"min"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Medium | IEC 61010-2-020;FDA QMSR | spin-down;lab centrifuge;Eppendorf | {} | OQ verifies RPM accuracy with an external tachometer.
| Refrigerated Centrifuge | Laboratory | High-Speed Chilled | {"temp_range":{"type":"range","units":"C"},"chamber_cooling":{"type":"boolean"},"rotor_recognition":{"type":"boolean"}} | TRUE | TRUE | TRUE | 12 | 6 | 12 | Medium | IEC 61010-2-020;USP <1058> | chilled centrifuge;blood spinner;cold spin | {} | OQ includes verification of chamber temperature at max RPM.
| Biological Indicator Incubator | Testing | Rapid Readout Station | {"incubation_temp":{"type":"number","units":"C"},"well_count":{"type":"integer"},"auto_read":{"type":"boolean"}} | TRUE | TRUE | FALSE | 12 | 12 |  | High | ISO 11138;ANSI/AAMI ST79 | BI incubator;Attest reader;sterility tester | {} | Critical for autoclave release. OQ verifies temp stability of individual wells.
| Handheld Barcode Verifier | Labeling | 1D/2D ANSI Grade | {"grading_std":{"type":"enum","values":},"aperture":{"type":"number","units":"mil"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Medium | 21 CFR 830;ISO 15426 | UDI verifier;label grader;barcode scanner | {} | Not just a scanner; this validates the "Print Quality" of UDI labels.
| Environmental Test Chamber | Testing | Large Scale Reach-In | {"temp_range":{"type":"range","units":"C"},"humidity_range":{"type":"range","units":"%RH"},"interior_volume":{"type":"number","units":"m3"}} | TRUE | TRUE | TRUE | 6 | 6 | 12 | High | ISO 13485;ICH Q1A | stability chamber;humidity box;shelf-life oven | {} | PQ requires mapping the chamber for uniformity under loaded conditions.
| Motorized Force Gauge | Measurement & Metrology | Tensile/Compression Stand | {"force_capacity":{"type":"number","units":"N"},"speed_range":{"type":"range","units":"mm/min"},"travel_limit":{"type":"number","units":"mm"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Medium | ISO 17025;FDA QMSR | pull tester;force stand;Mark-10;Chatillon | {} | Used for simple tensile tests (e.g., Luer lock strength).
| Digital Pressure Gauge | Measurement & Metrology | High-Precision Master | {"range":{"type":"range","units":"PSI"},"accuracy_class":{"type":"number","units":"%"},"wetted_parts":{"type":"string"}} | TRUE | FALSE | FALSE | 6 |  |  | Medium | ISO 17025;FDA QMSR | master gauge;reference pressure;pneumatic meter | {} | Typically used as a reference standard to calibrate process sensors.
| Cleanroom Washer-Extractor | Facilities: Cleaning | Barrier Style Industrial | {"load_capacity":{"type":"number","units":"kg"},"water_quality":{"type":"enum","values":},"g_force":{"type":"number"}} | TRUE | TRUE | TRUE | 12 | 6 | 12 | Medium | ISO 14644;ISO 13485;21 CFR 820.70 | garment washer;barrier washer;laundry extractor | {} | PQ verifies particulate reduction and bioburden levels on cleanroom suits.
| Cleanroom Tumble Dryer | Facilities: Cleaning | HEPA-Filtered Exhaust | {"temp_range":{"type":"range","units":"C"},"airflow_volume":{"type":"number","units":"CFM"},"filter_efficiency":{"type":"number","units":"%"}} | TRUE | TRUE | TRUE | 12 | 6 | 12 | Medium | ISO 14644-14;ISO 13485 | suit dryer;cleanroom dryer;HEPA dryer | {} | PQ involves particle count verification within the drum during operation.
| Specialized Gas Manifold | Utilities: Gas | Automatic Changeover System | {"gas_purity":{"type":"number","units":"%"},"max_pressure":{"type":"number","units":"PSI"},"leak_rate":{"type":"number","units":"sccm"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Medium | ISO 7396;FDA QMSR;ISO 13485 | gas switcher;bulk gas manifold;O2 manifold | {} | OQ verifies pressure switchover logic and alarm triggers.
| NAT Screening System | HCT/P: Screening | High-Throughput PCR | {"sensitivity":{"type":"number","units":"IU/mL"},"lod":{"type":"number","units":"copies"},"throughput":{"type":"number","units":"tests/run"}} | TRUE | TRUE | TRUE | 6 | 6 | 12 | High | 21 CFR 1271;21 CFR 610;USP <1058> | NAT analyzer;HIV/HCV screener;nucleic acid tester | {} | Mandatory for HCT/P donor eligibility. Requires extensive CSV.
| Automated Immunoassay Analyzer | HCT/P: Screening | Chemiluminescent (CLIA) | {"analytical_sensitivity":{"type":"number"},"reagent_cooling":{"type":"boolean"},"sample_probe_cv":{"type":"number","units":"%"}} | TRUE | TRUE | TRUE | 6 | 6 | 12 | High | 21 CFR 1271;21 CFR 866;CLIA | CLIA machine;infectious disease screener;donor tester | {} | Used for HBsAg, anti-HIV, and syphilis screening of HCT/P donors.
| Bone Mill (Automated) | HCT/P: Processing | Precision Allograft Grinder | {"blade_speed":{"type":"number","units":"RPM"},"output_size":{"type":"range","units":"um"},"material":{"type":"string"}} | TRUE | TRUE | TRUE | 12 | 12 | 12 | High | 21 CFR 1271.200;ISO 13485 | bone grinder;allograft mill;morselizer | {} | PQ verifies consistent particle size distribution for bone grafts.
| Automatic Colony Picker | Laboratory: Micro | Imaging-Based Robotic | {"picking_speed":{"type":"number","units":"colonies/hr"},"camera_res":{"type":"number","units":"MP"},"plate_capacity":{"type":"integer"}} | TRUE | TRUE | TRUE | 12 | 6 | 12 | Medium | GAMP 5;ISO 13485;21 CFR Part 11 | colony bot;microbiology robot;QPix | {} | OQ verifies picking accuracy (targeting) and cross-contamination prevention.
| V-Blender (Twin Shell) | Production: Mixing | Sanitary Powder Mixer | {"total_volume":{"type":"number","units":"L"},"max_fill_level":{"type":"number","units":"%"},"mixing_speed":{"type":"number","units":"RPM"}} | TRUE | TRUE | TRUE | 12 | 12 | 12 | Medium | 21 CFR 820.75;ISO 13485 | twin shell;V-mixer;powder blender | {} | PQ verifies blend uniformity (RSD < 5%) using thief sampling.
| Decellularization System | HCT/P: Processing | Automated Perfusion Unit | {"pump_flow_rate":{"type":"range","units":"mL/min"},"pressure_limit":{"type":"number","units":"mmHg"},"temp_control":{"type":"boolean"}} | TRUE | TRUE | TRUE | 6 | 12 | 12 | High | 21 CFR 1271;ISO 13485 | tissue stripper;de-cell unit;HCTP processor | {} | PQ must demonstrate complete removal of cellular DNA from tissue scaffold.
| GC-TD System | Analytical Lab | Thermal Desorption GC | {"desorption_temp":{"type":"number","units":"C"},"trap_cooling":{"type":"number","units":"C"},"oven_max":{"type":"number","units":"C"}} | TRUE | TRUE | TRUE | 12 | 6 |  | High | ISO 10993-7;ISO 10993-18;USP <1058> | residual GC;outgassing tester;TD-GCMS | {} | Used for analyzing volatile residuals in orthopedic implants.
| Continuous Air Particle Manifold | Facilities: Monitoring | Isokinetic Sampling System | {"port_count":{"type":"integer"},"sample_flow":{"type":"number","units":"LPM"},"vacuum_source":{"type":"string"}} | TRUE | TRUE | FALSE | 12 | 12 | 12 | High | ISO 14644-2;FDA QMSR | FMS manifold;cleanroom air monitor | {} | OQ verifies vacuum stability at all sampling points simultaneously.
| Vacuum Bubble Leak Tester | Production: Inspection | ASTM D3078 System | {"vacuum_level":{"type":"number","units":"inHg"},"chamber_vol":{"type":"number","units":"L"},"timer_accuracy":{"type":"number","units":"sec"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Medium | ASTM D3078;ISO 11607 | bubble tester;pouch leak tester;immersion leak | {} | Used for sterile barrier integrity testing (package validation).
| Sintering Furnace | Production: Additive | High-Temp Vacuum/Inert | {"max_temp":{"type":"number","units":"C"},"ramp_rate":{"type":"number","units":"C/min"},"atmosphere":{"type":"string"}} | TRUE | TRUE | TRUE | 6 | 6 | 12 | High | ISO 13485;ASTM F2924;ASTM F3001 | metal sinter;additive oven;3D print furnace | {} | PQ involves temp mapping and density/hardness verification of parts.
| Micro-Milling Machine | Production: Machining | 5-Axis Dental/Ortho | {"spindle_rpm":{"type":"number"},"tool_dia_min":{"type":"number","units":"mm"},"position_accuracy":{"type":"number","units":"um"}} | TRUE | TRUE | TRUE | 12 | 6 | 24 | High | ISO 13485;21 CFR 820.70 | dental mill;micro-CNC;small part mill | {} | Used for custom bone plates and dental abutments.
| MALDI-TOF MS | Analytical Lab | Microbial Identification | {"mass_range":{"type":"range","units":"Da"},"laser_type":{"type":"string"},"library_size":{"type":"integer"}} | TRUE | TRUE | TRUE | 12 | 6 |  | High | USP <1058>;21 CFR Part 11;ISO 13485 | MALDI;microbial ID;Biotyper | {} | Used for rapid environmental isolate identification in cleanrooms.
| Cleanroom Garment Cabinet | Facilities | HEPA-Filtered Laminar | {"hepa_efficiency":{"type":"number","units":"%"},"air_velocity":{"type":"number","units":"fpm"},"uv_light":{"type":"boolean"}} | TRUE | FALSE | FALSE | 12 | 12 |  | Low | ISO 14644;FDA QMSR | gowning rack;HEPA cabinet;suit storage | {} | IQ verifies laminar flow and surface cleanliness levels.
| Rotary Evaporator | Laboratory: Reagent | Reagent Solvent Recovery | {"bath_temp":{"type":"number","units":"C"},"rotation_speed":{"type":"number","units":"RPM"},"vacuum_cap":{"type":"number","units":"mbar"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Medium | 21 CFR 820.70;ISO 13485 | rotovap;solvent stripper;reagent concentrator | {} | Used in IVD reagent preparation to concentrate analytes.
| Bone Saw (Manual/Pneumatic) | HCT/P: Processing | Reusable Surgical Tool | {"air_pressure_req":{"type":"number","units":"psi"},"stroke_speed":{"type":"number","units":"spm"},"blade_compatibility":{"type":"string"}} | TRUE | TRUE | TRUE | 12 | 12 |  | High | ISO 13485;ISO 17664;21 CFR 1271 | surgical saw;ortho saw;allograft saw | {} | PQ focuses on heat prevention to maintain tissue osteoinductivity.
| Nitrogen/Oxygen Gas Mixer | Utilities | High-Accuracy Blender | {"mix_accuracy":{"type":"number","units":"%"},"flow_range":{"type":"range","units":"LPM"},"fail_safe":{"type":"boolean"}} | TRUE | TRUE | TRUE | 6 | 12 | 12 | High | 21 CFR 820.70;ISO 13485 | gas blender;proportional mixer;hypoxic mixer | {} | Used for hypoxic cell culture (HCT/P). OQ verifies mix ratio stability.
| Automated Tape Applicator | Production: Packaging | Medical Grade Taper | {"tape_width":{"type":"number","units":"mm"},"application_force":{"type":"number","units":"N"},"speed":{"type":"number","units":"m/min"}} | TRUE | TRUE | TRUE | 12 | 12 |  | Low | ISO 13485;FDA QMSR | box taper;packaging taper;case sealer | {} | PQ verifies seal continuity and adhesive bonding to the substrate.
| Cryogenic Bulk Storage Tank | Cold Chain | Vacuum-Insulated LN2 | {"capacity":{"type":"number","units":"L"},"static_evap_rate":{"type":"number","units":"%/day"},"pressure_rating":{"type":"number","units":"psi"}} | TRUE | TRUE | TRUE | 12 | 12 | 24 | High | 21 CFR 1271.260;ISO 13485 | LN2 tank;bulk cryo;liquid nitrogen reservoir | {} | PQ involves "Hold Time" study and vacuum integrity verification.
| Bulk CO2 Storage Tank | Facilities: Gas | Low-Pressure Refrigerated | {"storage_temp":{"type":"number","units":"C"},"purity_spec":{"type":"number","units":"%"},"fill_level_alarm":{"type":"boolean"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Medium | ISO 7396;FDA QMSR | CO2 bulk;incubator gas supply;gas silo | {} | Essential for high-volume HCT/P incubator farms. OQ verifies alarm thresholds.
| Inline UDI Barcode Scanner | Production: Labeling | 360-Degree Code Reader | {"scan_rate":{"type":"number","units":"scans/sec"},"symbology_support":{"type":"array"},"decoding_grade":{"type":"string"}} | TRUE | TRUE | TRUE | 12 | 12 |  | High | 21 CFR 830;21 CFR Part 11 | UDI scanner;barcode bot;Cognex inline;Keyence SR | {} | PQ verifies read rates at maximum conveyor speed with degraded codes.
| Spinal Construct Fatigue Tester | Metrology: Mechanical | Static/Dynamic Axial-Torsion | {"max_axial_load":{"type":"number","units":"kN"},"freq_range":{"type":"range","units":"Hz"},"actuator_stroke":{"type":"number","units":"mm"}} | TRUE | TRUE | TRUE | 6 | 6 | 24 | High | ASTM F1717;ASTM F2706;ISO 12189 | spine tester;implant fatigue;Zwick spine | {} | Critical for validating vertebral body fusion devices and rod systems.
| Torsional Bone Screw Tester | Metrology: Mechanical | Automated Torque-Angle | {"torque_cap":{"type":"number","units":"N-m"},"angular_res":{"type":"number","units":"deg"},"axial_preload":{"type":"number","units":"N"}} | TRUE | TRUE | TRUE | 12 | 12 |  | High | ASTM F543;ISO 17025 | screw torque;breaking angle;bone screw validation | {} | PQ verifies torsional yield strength and maximum torque to failure.
| Dynamic Hip Simulator | Metrology: Mechanical | Multi-Station Wear Tester | {"stations":{"type":"integer"},"load_profile":{"type":"string"},"fluid_temp_control":{"type":"boolean"}} | TRUE | TRUE | TRUE | 12 | 6 | 24 | High | ISO 7206-4;ISO 14242;ASTM F3047M | hip simulator;wear tester;prosthetic durability | {} | PQ typically runs for 5 million+ cycles with bovine serum lubricant.
| Knee Simulator | Metrology: Mechanical | Multi-Axis Displacement-Control | {"flexion_range":{"type":"range","units":"deg"},"axial_force":{"type":"number","units":"kN"},"sampling_rate":{"type":"number","units":"Hz"}} | TRUE | TRUE | TRUE | 12 | 6 | 24 | High | ISO 14243;ASTM F2777;ISO 21536 | knee simulator;TKR tester;joint wear | {} | PQ involves gravimetric wear measurement of the tibial insert.
| Micro-Hardness Tester | Metrology: Material | Vickers/Knoop Semi-Auto | {"test_load":{"type":"range","units":"gf"},"dwell_time":{"type":"number","units":"sec"},"objective_lens":{"type":"array"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Medium | ASTM E384;ISO 17025 | Vickers tester;implant hardness;Knoop scope | {} | OQ verifies indent measurement accuracy using NIST-traceable blocks.
| DC Plasma Spray System | Production: Surface Treatment | Hydroxyapatite Coater | {"plate_power":{"type":"number","units":"kW"},"spray_dist":{"type":"number","units":"mm"},"feed_gas_ratio":{"type":"string"}} | TRUE | TRUE | TRUE | 12 | 6 | 12 | High | ISO 13779-2;ASTM F1147 | HA sprayer;DC plasma;implant coater | {} | PQ must include HA crystallinity (XRD) and coating adhesion testing.
| Vacuum Sintering Furnace | Production: Additive | High-Temp Inert Gas | {"max_temp":{"type":"number","units":"C"},"vacuum_level":{"type":"number","units":"mTorr"},"temp_uniformity":{"type":"number","units":"C"}} | TRUE | TRUE | TRUE | 6 | 6 | 12 | High | ISO 13485;ASTM F2924 | sinter oven;metal 3D print oven;Ti furnace | {} | PQ involves 9-point thermal mapping and part density analysis.
| Automated Grit Blaster | Production: Finishing | CNC Surface Texturizer | {"nozzle_pressure":{"type":"number","units":"psi"},"media_grit_size":{"type":"integer"},"cycle_time":{"type":"number","units":"sec"}} | TRUE | TRUE | TRUE | 12 | 12 |  | Medium | ISO 13485;FDA QMSR | sandblaster;implant finisher;bead blast | {} | OQ verifies Ra (surface roughness) consistency across the implant surface.
| DSC Analyzer | Analytical Lab: Materials | High-Sensitivity Calorimeter | {"temp_range":{"type":"range","units":"C"},"heating_rate":{"type":"number","units":"C/min"},"baseline_noise":{"type":"number","units":"uW"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Medium | USP <1058>;ISO 11357 | DSC;thermal analyzer;polymer crystallinity | {} | USP <1058> Group C. Used to validate PEEK/Polymer raw materials.
| TGA Analyzer | Analytical Lab: Materials | Thermogravimetric Balance | {"balance_resolution":{"type":"number","units":"ug"},"max_temp":{"type":"number","units":"C"},"atmosphere_control":{"type":"array"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Medium | USP <1058>;ISO 11358 | thermal weight;TGA;ash content analyzer | {} | Critical for detecting moisture content and filler ratios in medical plastics.
| LC-MS/MS System | Analytical Lab: Trace | Triple Quadrupole E&L | {"mass_range":{"type":"range","units":"m/z"},"sensitivity":{"type":"number","units":"pg"},"source_type":{"type":"string"}} | TRUE | TRUE | TRUE | 12 | 6 | 12 | High | ISO 10993-18;21 CFR Part 11 | LCMS;extractables tester;leachables analyzer | {} | Critical for biocompatibility/E&L profiles of implants and catheters.
| ICP-OES Analyzer | Analytical Lab: Trace | Elemental Inductive Plasma | {"wavelength_range":{"type":"range","units":"nm"},"lod":{"type":"number","units":"ppb"},"nebulizer_type":{"type":"string"}} | TRUE | TRUE | TRUE | 12 | 6 | 12 | Medium | USP <232>;USP <233> | metal analyzer;heavy metal tester;ICP | {} | Replacement for traditional USP <231> heavy metal tests.
| Digital Profile Projector | Metrology: Optical | 2D Vision Measurement | {"fov_size":{"type":"string"},"magnification":{"type":"number"},"accuracy":{"type":"number","units":"um"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Medium | ISO 17025;FDA QMSR | optical comparator;Keyence IM;shadowgraph | {} | OQ verifies pixel-to-dimension scaling using a certified glass slide.
| Ultrasonic Flow Meter | Metrology: Fluid | Non-Invasive Clamp-On | {"flow_range":{"type":"range","units":"LPM"},"pipe_dia":{"type":"range","units":"mm"},"accuracy":{"type":"number","units":"%"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Medium | ISO 17025;FDA QMSR | flow meter;clamp flow;catheter flow tester | {} | Used to validate WFI/PW loop flow velocities (non-contact).
| Peel Strength Tester | Metrology: Packaging | 180-Degree T-Peel Stand | {"force_cap":{"type":"number","units":"N"},"crosshead_speed":{"type":"number","units":"mm/min"},"stroke":{"type":"number","units":"mm"}} | TRUE | TRUE | TRUE | 12 | 12 |  | High | ASTM F88;ISO 11607 | peel tester;seal strength;pouch tester | {} | PQ verifies sterile barrier seal strength consistency.
| Bubble Emission Leak Tester | Metrology: Packaging | ASTM D3078 Vacuum Tank | {"vacuum_max":{"type":"number","units":"inHg"},"tank_vol":{"type":"number","units":"L"},"timer_acc":{"type":"number","units":"sec"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Medium | ASTM D3078;ISO 11607 | bubble tester;pouch leak bath;immersion tester | {} | Essential for determining gross leaks in flexible packaging.
| Dye Penetration Integrity Bath | Metrology: Packaging | ASTM F1929 System | {"dye_composition":{"type":"string"},"rinse_station":{"type":"boolean"},"lighting_std":{"type":"string"}} | TRUE | FALSE | FALSE |  | 12 |  | Medium | ASTM F1929;ISO 11607 | dye leak;wicking tester;package integrity | {} | IQ verifies dye formulation and visual detection area lighting.
| Heart-Lung Machine Console | Cardiovascular Support | Cardiopulmonary Bypass Console | {"pump_heads":{"type":"integer"},"battery_runtime":{"type":"number","units":"min"},"integrated_monitors":{"type":"integer"}} | TRUE | TRUE | TRUE | 12 | 6 | 12 | High | 21 CFR 870.4230;IEC 60601-1 | CPB console;perfusion pump;bypass machine | {} | OQ verifies all alarm logic (bubble, level, pressure) and battery switchover.
| Roller-Type Blood Pump | Cardiovascular Support | CPB Arterial/Suction Pump | {"flow_range":{"type":"range","units":"L/min"},"occlusion_setting":{"type":"string"},"rpm_accuracy":{"type":"number","units":"%"}} | TRUE | TRUE | TRUE | 6 | 12 | 12 | High | 21 CFR 870.4380;ISO 13485 | arterial pump;roller pump;blood mover | {} | PQ verifies flow rate consistency against calibrated flow meters.
| Cardiopulmonary Temp Controller | Cardiovascular Support | CPB Heater-Cooler Unit | {"temp_range":{"type":"range","units":"C"},"heating_cap":{"type":"number","units":"W"},"cooling_cap":{"type":"number","units":"W"}} | TRUE | TRUE | TRUE | 12 | 6 | 12 | High | 21 CFR 870.4250;IEC 60601-1 | CPB heater;cooler unit;thermal exchanger | {} | OQ verifies temp stability and microbial control procedures for water tanks.
| High-Power X-Ray Generator | Diagnostic Imaging | Cardiac Cath Lab Generator | {"power_output":{"type":"number","units":"kW"},"kvp_range":{"type":"range","units":"kV"},"pulse_freq":{"type":"number","units":"fps"}} | TRUE | TRUE | TRUE | 12 | 12 | 24 | High | 21 CFR 1020.30;IEC 60601-2-43 | X-ray source;imaging generator;high-voltage tank | {} | OQ verifies pulse width (3-8 ms) and kVp accuracy for radiation safety.
| Digital Flat-Panel Detector | Diagnostic Imaging | Dynamic Cardiac X-Ray | {"field_size":{"type":"number","units":"in"},"resolution":{"type":"number","units":"lp/mm"},"frame_rate":{"type":"number","units":"fps"}} | TRUE | TRUE | FALSE | 12 | 12 |  | High | 21 CFR 1020;ISO 13485 | FPD;dynamic detector;DR panel | {} | OQ verifies DQE (detective quantum efficiency) and signal-to-noise ratio.
| Automated Contrast Injector | Diagnostic Imaging | Angiographic Power Injector | {"injection_rate":{"type":"range","units":"mL/s"},"pressure_limit":{"type":"number","units":"psi"},"syringe_vol":{"type":"number","units":"mL"}} | TRUE | TRUE | TRUE | 12 | 12 |  | Medium | 21 CFR 870.1650;FDA QMSR | power injector;dye injector;Medrad | {} | OQ verifies pressure-cutoff and volume accuracy to prevent vessel damage.
| Mobile VHP Decon Robot | Environment & Safety | Autonomous Room Fogger | {"h2o2_conc":{"type":"number","units":"%"},"cycle_time":{"type":"number","units":"min"},"room_mapping":{"type":"boolean"}} | TRUE | TRUE | TRUE | 12 | 6 | 12 | High | ISO 22441;ISO 14644 | VHP bot;cleanroom robot;Steris bot | {} | PQ requires biological indicator (BI) 6-log reduction mapping of rooms.
| Active Bio-Air Sampler | Environment & Safety | Isokinetic Cleanroom Sampler | {"flow_rate":{"type":"number","units":"LPM"},"sampling_vol":{"type":"number","units":"L"},"timer_acc":{"type":"number","units":"sec"}} | TRUE | TRUE | FALSE | 12 | 12 |  | High | ISO 14698;FDA QMSR | bio-sampler;active air monitor;micro sampler | {} | OQ verifies flow rate stability (e.g., 100 LPM) across different media types.
| Endoscope Washer-Disinfector | Cleaning | Automated (AER) | {"chamber_count":{"type":"integer"},"leak_test_integrated":{"type":"boolean"},"disinfectant_temp":{"type":"number","units":"C"}} | TRUE | TRUE | TRUE | 12 | 6 | 12 | High | ISO 15883-4;AAMI ST98 | AER;scope washer;reprocessor | {} | PQ verifies high-level disinfection (HLD) using simulated use soil tests.
| H2O2 Gas Plasma Sterilizer | Sterilization | Low-Temp Vapor Unit | {"chamber_vol":{"type":"number","units":"L"},"cycle_temp":{"type":"number","units":"C"},"plasma_gen":{"type":"string"}} | TRUE | TRUE | TRUE | 6 | 6 | 12 | High | ISO 14937;FDA QMSR | STERRAD;plasma sterilizer;VHP unit | {} | Ideal for electronics/fiberoptics. PQ requires 3 consecutive successful BI runs.
| Micro-Resistance Tester | Metrology | 4-Wire Kelvin Meter | {"range":{"type":"range","units":"mOhm"},"test_current":{"type":"number","units":"mA"},"accuracy":{"type":"number","units":"%"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Medium | IPC-6012 Class 3;ISO 17025 | Kelvin bridge;ohm meter;solder joint tester | {} | Essential for validating active implantable PCBA grounding paths.
| Defibrillator Analyzer | Metrology: Safety | Discharge Energy Tester | {"energy_range":{"type":"range","units":"J"},"waveform_analysis":{"type":"boolean"},"internal_load":{"type":"number","units":"Ohm"}} | TRUE | TRUE | FALSE | 12 | 12 |  | High | IEC 60601-2-4;ISO 17025 | defib tester;pacer analyzer;safety tester | {} | Used to validate energy delivery of external and internal defibrillators.
| Electrical Safety Analyzer | Metrology: Safety | Automated Leakage Tester | {"leakage_range":{"type":"range","units":"uA"},"ground_bond_test":{"type":"boolean"},"standards_support":{"type":"array"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Medium | IEC 60601-1;AAMI ES60601-1 | safety bot;leakage tester;Fluke ESA | {} | Mandatory for all active medical devices before release.
| Infusion Pump Calibration Standard | Metrology: Fluid | Precision Mass Flow Analyzer | {"flow_range":{"type":"range","units":"mL/hr"},"volume_acc":{"type":"number","units":"%"},"occlusion_press":{"type":"number","units":"psi"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Medium | IEC 60601-2-24;AAMI TIR101 | pump tester;flow analyzer;IV tester | {} | Used for validating the accuracy of drug delivery systems.
| Ventilator Gas Flow Analyzer | Metrology: Gas | High-Speed Respiratory Tester | {"flow_range":{"type":"range","units":"LPM"},"pressure_range":{"type":"range","units":"cmH2O"},"o2_conc":{"type":"boolean"}} | TRUE | TRUE | FALSE | 6 | 12 |  | High | ISO 80601-2-12;ISO 17025 | vent tester;respiratory analyzer;gas flow | {} | Critical for validating airway pressure and volume safety in ventilators.
| Hemodynamic Monitoring Console | Cardiovascular Monitoring | Multi-Channel Polygraph | {"waveforms_count":{"type":"integer"},"ecg_sample_rate":{"type":"number","units":"Hz"},"software_v":{"type":"string"}} | TRUE | TRUE | TRUE | 12 | 12 | 12 | High | IEC 60601-2-27;21 CFR Part 11 | polygraph;cath monitor;hemodynamic station | {} | OQ verifies real-time radiological image and ECG synchronization.
| Liquid Particle Counter | Metrology: IVD | Light Obscuration Analyzer | {"particle_size_range":{"type":"array"},"sample_vol":{"type":"number","units":"mL"},"viscosity_max":{"type":"number","units":"cP"}} | TRUE | TRUE | TRUE | 12 | 12 |  | High | USP <788>;ISO 13485 | sub-visible counter;liquid LPC;particle analyzer | {} | Critical for validating cleanliness of sterile injectables and tubing.
| Biological Indicator Auto-Reader | Testing: Sterility | Fluorescence Incubator | {"well_count":{"type":"integer"},"incubation_temp":{"type":"number","units":"C"},"detection_time":{"type":"number","units":"min"}} | TRUE | TRUE | FALSE | 12 | 12 |  | High | ISO 11138;FDA QMSR | Attest reader;BI incubator;release station | {} | OQ verifies temperature stability of each independent well.
| Laser Power / Energy Meter | Metrology: Production | Welding Validation Probe | {"wavelength_range":{"type":"range","units":"nm"},"power_limit":{"type":"number","units":"W"},"nist_traceable":{"type":"boolean"}} | TRUE | FALSE | FALSE | 12 |  |  | Medium | ISO 15609-4;ISO 17025 | laser probe;power puck;welding meter | {} | Reference standard to qualify laser welder power settings during OQ.
| Intra-Aortic Balloon Pump Console | Cardiovascular Support | Counterpulsation Console | {"inflation_speed":{"type":"number","units":"ms"},"sync_modes":{"type":"array"},"internal_battery":{"type":"boolean"}} | TRUE | TRUE | TRUE | 12 | 6 | 12 | High | 21 CFR 870.1025;IEC 60601-1 | IABP;balloon console;counter-pulsator | {} | PQ verifies trigger timing accuracy against ECG and pressure waveforms.
| Nitinol Shape Setting Furnace | Material Forming | Vacuum Heat Treat | {"max_temp":{"type":"number","units":"C"},"vacuum_level":{"type":"number","units":"mbar"},"quench_time":{"type":"number","units":"sec"}} | TRUE | TRUE | TRUE | 6 | 6 | 12 | High | ASTM F2082;ASTM F2516;ASTM F2063 | nitinol oven;shape set;vacuum furnace;Af tester | {} | PQ involves Active Austenite Finish (Af) temperature verification.
| Ultrasonic Spray Coater | Surface Treatment | Drug Eluting Stent (DES) | {"ultrasonic_freq":{"type":"number","units":"kHz"},"flow_rate":{"type":"number","units":"uL/min"},"nozzle_speed":{"type":"number","units":"mm/s"}} | TRUE | TRUE | TRUE | 6 | 6 | 12 | High | ISO 13485;FDA QMSR;ISO 10993 | DES coater;ultrasonic sprayer;Sono-Tek;stent coater | {} | PQ verifies drug load uniformity and presence of webbing.
| Catheter Balloon Folder | Assembly | Multi-Wing Pleater | {"fold_pressure":{"type":"number","units":"psi"},"heat_set_temp":{"type":"number","units":"C"},"dwell_time":{"type":"number","units":"sec"}} | TRUE | TRUE | TRUE | 12 | 12 |  | High | ISO 10555-4;FDA QMSR | balloon pleater;folder;winging machine | {} | Critical for profile and withdrawal testing per ISO 10555-4.
| Industrial Micro-CT | Metrology | 3D Non-Destructive | {"kvp":{"type":"number","units":"kV"},"voxel_size":{"type":"number","units":"um"},"max_part_size":{"type":"string"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Medium | ISO 15708;ISO 17025 | Micro-CT;Zeiss Metrotom;NDT scanner | {} | OQ verifies resolution against certified phantom materials.
| Automated Reagent Vial Labeler | Labeling | High-Speed Wraparound | {"label_acc":{"type":"number","units":"mm"},"speed":{"type":"number","units":"VPM"},"udi_verified":{"type":"boolean"}} | TRUE | TRUE | TRUE | 12 | 12 |  | Medium | 21 CFR 830;21 CFR Part 11 | vial labeler;IVD labeler;Zebra wrap | {} | PQ verifies label adhesion and scannability at peak rates.
| Liquid Scintillation Counter | Analytical Lab | Beta-Emitter Counter | {"efficiency":{"type":"number","units":"%"},"background":{"type":"number","units":"CPM"},"isotopes":{"type":"array"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Medium | USP ;ISO 13485 | scintillation;Beckman counter;beta tester | {} | Used in HCT/P laboratories for radioactive isotope counting.
| Laser Diffraction Particle Sizer | Analytical Lab | Sub-Micron Sizer | {"range":{"type":"range","units":"um"},"obscuration":{"type":"number","units":"%"},"lasers":{"type":"integer"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Medium | USP ;USP | Malvern sizer;Horiba;particle distribution | {} | USP Group B. Calibrated with NIST latex spheres.
| Salt Bath Furnace | Material Forming | Nitinol Fluidized Bed | {"bath_temp":{"type":"number","units":"C"},"cooling_rate":{"type":"number","units":"C/s"},"atmosphere":{"type":"string"}} | TRUE | TRUE | TRUE | 6 | 6 | 12 | High | ASTM F2063;ISO 13485 | nitinol salt bath;shape setting bath;fluidized bed | {} | OQ verifies temperature uniformity of the salt medium.
| Orthopedic Surgical Drill | HCT/P Processing | High-Torque Pneumatic | {"rpm_max":{"type":"number"},"torque_limit":{"type":"number","units":"N-m"},"air_press":{"type":"number","units":"psi"}} | TRUE | TRUE | TRUE | 12 | 12 |  | High | 21 CFR 1271;ISO 17664 | surgical drill;ortho power;allograft drill | {} | PQ focuses on preventing thermal necrosis in tissue grafts.
| Orthopedic Reamer | HCT/P Processing | Flexible Shaft Reamer | {"cutting_dia":{"type":"number","units":"mm"},"material_comp":{"type":"string"},"rpm_range":{"type":"range"}} | TRUE | TRUE | TRUE | 12 | 12 |  | High | 21 CFR 1271;ISO 17664 | bone reamer;medullary tool;flexible reamer | {} | Qualified as reusable processing equipment for donor tissue.
| Nano-scale Reagent Dispenser | IVD Production | Piezo-Electric Jetting | {"drop_vol":{"type":"number","units":"nL"},"pitch_acc":{"type":"number","units":"um"},"cv_pct":{"type":"number","units":"%"}} | TRUE | TRUE | TRUE | 6 | 6 | 12 | High | ISO 13485;FDA QMSR | nano-dispenser;BioDot;micro-jet | {} | PQ verifies coefficient of variation (CV) for dispense accuracy.
| Dynamic Mechanical Analyzer | Analytical Lab | Viscoelastic Tester | {"freq_range":{"type":"range","units":"Hz"},"max_force":{"type":"number","units":"N"},"temp_acc":{"type":"number","units":"C"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Medium | USP ;ISO 6721 | DMA;polymer tester;stiffness analyzer | {} | Critical for characterization of bioabsorbable polymers.
| X-Ray Photoelectron Spectrometer | Analytical Lab | Surface Chemistry XPS | {"beam_dia":{"type":"number","units":"um"},"sensitivity":{"type":"number","units":"cps"},"pass_energy":{"type":"number","units":"eV"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Medium | ISO 13485;FDA QMSR | XPS;ESCA;surface analyzer;implant surface | {} | High-complexity Group C instrument. Used for passivated surfaces.
| Automated Pouch Loader | Packaging | Robotic Inserter | {"cycle_time":{"type":"number","units":"sec"},"placement_acc":{"type":"number","units":"mm"},"vacuum_pick":{"type":"boolean"}} | TRUE | TRUE | TRUE | 12 | 12 |  | Low | ISO 11607;ISO 13485 | pouch loader;product inserter;packaging robot | {} | PQ verifies zero-defect orientation to prevent seal compromise.
| Suture Attachment Machine | Joining | Needle-Suture Swager | {"swage_force":{"type":"number","units":"N"},"pull_force":{"type":"number","units":"N"},"die_type":{"type":"string"}} | TRUE | TRUE | TRUE | 12 | 12 | 12 | High | 21 CFR 820.75;ISO 13485 | swager;suture joiner;needle swage | {} | PQ requires USP pull-test verification for 3 consecutive batches.
| Automated Guidewire Grinder | Machining | Precision Profile Grinder | {"taper_angle":{"type":"number","units":"deg"},"wheel_rpm":{"type":"number"},"coolant_flow":{"type":"number","units":"L/min"}} | TRUE | TRUE | TRUE | 12 | 6 |  | Medium | ISO 13485;21 CFR 820 | guidewire grinder;mandrel grinder;profile grind | {} | OQ verifies repeatability of multi-segment tapers.
| Central Vacuum System | Utilities | Medical Grade Oil-Free | {"flow_cap":{"type":"number","units":"SCFM"},"vacuum_level":{"type":"number","units":"inHg"},"redundancy":{"type":"boolean"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Medium | ISO 7396;FDA QMSR | plant vac;process vacuum;house vac | {} | OQ verifies vacuum stability under maximum line demand.
| RO Skid (Reverse Osmosis) | Utilities | Pharmaceutical Grade | {"flow_perm":{"type":"number","units":"GPM"},"perm_cond":{"type":"number","units":"uS/cm"},"rejection":{"type":"number","units":"%"}} | TRUE | TRUE | TRUE | 6 | 6 | 12 | High | USP ;ISO 13485 | RO system;water maker;USP water | {} | Phase I validation requires daily monitoring for 2-4 weeks.
| Deionization (DI) Skid | Utilities | Electrodeionization (EDI) | {"conductivity":{"type":"number","units":"uS/cm"},"resin_type":{"type":"string"},"flow_rate":{"type":"number","units":"LPM"}} | TRUE | TRUE | TRUE | 6 | 6 | 12 | High | USP ;ISO 13485 | DI water;EDI skid;demineralizer | {} | Critical for removing final ions from process rinse water.
| Multi-Parameter Patient Simulator | Metrology | Vital Sign Standard | {"heart_rate":{"type":"range","units":"BPM"},"waveforms":{"type":"array"},"nibp_acc":{"type":"number"}} | TRUE | FALSE | FALSE | 12 |  |  | Medium | IEC 60601;ISO 17025 | patient sim;NIBP tester;vital sign tester | {} | Standard to qualify monitors and therapeutic heart devices.
| Active Lead Continuity Tester | Active Implantable | 4-Wire Kelvin Bench | {"test_current":{"type":"number","units":"mA"},"resistance_range":{"type":"range","units":"mOhm"},"dwell_time":{"type":"number","units":"ms"}} | TRUE | TRUE | TRUE | 12 | 12 |  | High | ISO 27186;IEC 60601-1;IPC-6012 | lead tester;Kelvin bridge;pacer lead tester | {} | Critical for verifying grounding paths in active implantable leads.
| Helium Mass Spectrometer | Metrology | Vacuum Leak Detector | {"sensitivity":{"type":"number","units":"mbar.L/s"},"vacuum_source":{"type":"string"},"sniff_mode":{"type":"boolean"}} | TRUE | TRUE | FALSE | 12 | 12 |  | High | ISO 13485;21 CFR 820 | helium leak;He tester;mass spec leak | {} | Gold standard for hermeticity testing of pacemaker cans and valves.
| Automated Flow Bench | Metrology | Catheter Patency Tester | {"pressure":{"type":"number","units":"psi"},"flow_range":{"type":"range","units":"mL/min"},"fluid_temp":{"type":"number","units":"C"}} | TRUE | TRUE | TRUE | 12 | 12 |  | High | ISO 10555;FDA Guidance 824 | patency tester;catheter flow;flow bench | {} | PQ verifies flow rate consistency across 3 consecutive batches.
| Large-Scale Bioreactor | Reagent Mfg | Stainless Steel Stirred | {"volume":{"type":"number","units":"L"},"do2_range":{"type":"range","units":"%"},"temp_uniformity":{"type":"number","units":"C"}} | TRUE | TRUE | TRUE | 12 | 6 | 12 | High | 21 CFR 820.75;ISO 13485 | fermenter;aseptic tank;cell culture tank | {} | PQ includes media fills and metabolic profile stability.
| Tangential Flow Filtration (TFF) | Reagent Mfg | Ultrafiltration Skid | {"tmp_pressure":{"type":"number","units":"bar"},"crossflow_rate":{"type":"number","units":"LPM"},"membrane_area":{"type":"number","units":"m2"}} | TRUE | TRUE | TRUE | 12 | 6 | 12 | Medium | USP ;ISO 13485 | TFF skid;dialysis skid;concentration unit | {} | OQ verifies flux recovery and clean-water permeate stability.
| Industrial Reagent Lyophilizer | Sterilization | Aseptic Shelf Dryer | {"condenser_temp":{"type":"number","units":"C"},"shelf_mapping":{"type":"boolean"},"vacuum_acc":{"type":"number","units":"mTorr"}} | TRUE | TRUE | TRUE | 6 | 6 | 12 | High | ISO 13485;21 CFR 820.75 | industrial freeze dryer;bulk lyo;aseptic dryer | {} | PQ involves a 9-point thermal mapping of shelf uniformity.
| Syringe Visual Inspection System | Inspection | High-Speed Automated | {"particle_lod":{"type":"number","units":"um"},"inspection_speed":{"type":"number","units":"syringes/min"},"camera_count":{"type":"integer"}} | TRUE | TRUE | TRUE | 12 | 12 | 12 | High | ISO 13485;21 CFR Part 11 | AVI;Knapp test;particle inspector | {} | OQ uses Knapp-test vials (known particulate) to verify detection LOD.
| Micro-Torque Lead Attacher | Assembly | Automated Pacer Lead | {"torque_limit":{"type":"number","units":"N-cm"},"angular_feedback":{"type":"boolean"},"software_v":{"type":"string"}} | TRUE | TRUE | FALSE | 12 | 12 |  | High | ISO 13485;21 CFR 820 | torque bot;lead attacher;pacer assembler | {} | Critical for attaching leads to pacemaker headers without over-stressing.
| WFI Circulating Loop Skid | Utilities | Hot Storage & Dist | {"loop_temp":{"type":"number","units":"C"},"velocity_min":{"type":"number","units":"m/s"},"turbulent_flow":{"type":"boolean"}} | TRUE | TRUE | TRUE | 6 | 6 | 12 | High | USP ;ISO 13485 | water loop;WFI skid;hot water tank | {} | PQ Phase III requires 365 days of monitoring to capture seasonal variation.
| Vacuum Skin Packaging System | Packaging | Sterile Barrier Skin-Pack | {"vacuum_level":{"type":"number","units":"mbar"},"seal_temp":{"type":"number","units":"C"},"dwell_time":{"type":"number","units":"sec"}} | TRUE | TRUE | TRUE | 12 | 12 | 12 | High | ISO 11607;21 CFR 820.75 | skin packer;tray sealer;barrier sealer | {} | PQ verifies seal integrity and absence of creases in the skin layer.
| Electrosurgical RF Generator | Therapeutic | High-Frequency Console | {"output_power":{"type":"number","units":"W"},"load_impedance":{"type":"range","units":"Ohm"},"leakage_max":{"type":"number","units":"mA"}} | TRUE | TRUE | FALSE | 12 | 12 |  | High | IEC 60601-2-2;FDA 820 | ESU;cautery unit;RF console | {} | OQ verifies output accuracy across the full impedance range (10-1000 Ohms).
| Harmonic Scalpel Test Rig | Therapeutic | Ultrasonic Safety Bench | {"frequency":{"type":"number","units":"kHz"},"tip_amplitude":{"type":"number","units":"um"},"clamping_force":{"type":"number","units":"N"}} | TRUE | TRUE | FALSE | 12 | 12 |  | High | IEC 60601-1;ISO 13485 | ultrasonic scalpel;Harmonic tester;blade bench | {} | Used to calibrate and validate vibration performance of surgical blades.
| Micro-Wire Laser Stripper | Assembly | Non-Contact Insulation | {"laser_wavelength":{"type":"number","units":"nm"},"focal_acc":{"type":"number","units":"um"},"strip_length":{"type":"number","units":"mm"}} | TRUE | TRUE | TRUE | 12 | 12 |  | Medium | ISO 13485;ANSI Z136.1 | laser stripper;wire stripper;pacer lead prep | {} | Critical for active leads; ensures no mechanical nicks to conductor wires.
| Crimping Force Monitor | Metrology | Inline Solderless Join | {"force_range":{"type":"range","units":"N"},"sampling_rate":{"type":"number","units":"Hz"},"reject_limits":{"type":"string"}} | TRUE | TRUE | FALSE | 12 | 12 |  | High | ISO 13485;21 CFR 820 | crimp checker;force analyzer;lead crimp | {} | In-line monitoring for active implantable lead crimping.
| Reagent Solvent Still | Utilities | Solvent Purification | {"distillation_rate":{"type":"number","units":"L/hr"},"purity_cutoff":{"type":"number","units":"%"},"exhaust_safeties":{"type":"boolean"}} | TRUE | TRUE | TRUE | 12 | 12 |  | Medium | ISO 13485;21 CFR 820 | solvent reclaimer;reagent still;distiller | {} | OQ verifies vapor temperature and reflux ratio control.
| Clean-In-Place (CIP) Skid | Cleaning | Automated Loop System | {"flow_rate":{"type":"number","units":"LPM"},"conductivity_cutoff":{"type":"number","units":"uS/cm"},"cycle_time":{"type":"number","units":"min"}} | TRUE | TRUE | TRUE | 12 | 12 | 12 | High | ISO 13485;21 CFR 820 | CIP unit;reagent washer;loop cleaner | {} | PQ involves riboflavin (UV dye) coverage testing in production vessels.
| Sterilize-In-Place (SIP) System | Sterilization | Pure Steam Control | {"steam_pressure":{"type":"number","units":"psi"},"temp_mapping":{"type":"boolean"},"drain_temp":{"type":"number","units":"C"}} | TRUE | TRUE | TRUE | 12 | 12 | 12 | High | ISO 17665;21 CFR 820 | SIP control;steam-in-place;aseptic system | {} | PQ verifies lethality (Fo) at the most difficult-to-sterilize ports.
| Multi-Stage Ultrasonic Passivation | Finishing | High-Throughput Line | {"acid_bath_count":{"type":"integer"},"rinse_resistivity":{"type":"number","units":"MOhm"},"sonic_freq":{"type":"number","units":"kHz"}} | TRUE | TRUE | TRUE | 12 | 12 | 12 | High | ASTM A967;ISO 13485 | acid passivation;sonic finisher;implant line | {} | PQ verifies corrosion resistance (ASTM B117) of finished parts.
| Nitrogen Blanket Controller | Control | Tank Headspace Unit | {"pressure_setpoint":{"type":"number","units":"mbar"},"oxygen_lod":{"type":"number","units":"ppm"},"alarm_logic":{"type":"string"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Medium | ISO 13485;21 CFR 820 | inerting unit;N2 blanket;headspace controller | {} | OQ verifies fast-response oxygen purging during tank draw-down.
| Process Gas Mass Flow Controller | Utilities | High-Purity Gas MFC | {"gas_compatibility":{"type":"string"},"range":{"type":"range","units":"SCCM"},"linearity":{"type":"number","units":"%"}} | TRUE | TRUE | FALSE | 6 | 12 |  | Medium | ISO 17025;FDA 820 | MFC;gas meter;N2 controller | {} | Reference standard for controlling gas mix in laser welding or furnaces.
| High-Force Orthopedic Press | Production: Assembly | Hydraulic/Pneumatic Press | {"max_force":{"type":"number","units":"kN"},"stroke_limit":{"type":"number","units":"mm"},"platen_size":{"type":"string"}} | TRUE | TRUE | TRUE | 12 | 12 |  | High | ISO 13485;21 CFR 820.75;ASTM F1147 | ortho press;press-fit;HA press;implant joiner | {} | PQ verifies pull-out force and interference fit stability.
| Automated Potting Line | Production: Electronics | Active Implantable Potting | {"dispense_accuracy":{"type":"number","units":"%"},"cure_temp":{"type":"number","units":"C"},"degas_cycle":{"type":"boolean"}} | TRUE | TRUE | TRUE | 6 | 6 | 12 | High | ISO 13485;21 CFR 820.75;ISO 10993 | electronics potting;epoxy filler;pacer sealer | {} | Critical for hermeticity in active implants. OQ verifies zero-bubble fill.
| Biodegradable Extruder | Material Forming | PCL/PLGA Specialty Setup | {"melt_temp_acc":{"type":"number","units":"C"},"screw_torque":{"type":"number","units":"Nm"},"vacuum_venting":{"type":"boolean"}} | TRUE | TRUE | TRUE | 12 | 6 | 12 | High | ISO 13485;ASTM F1983;ASTM F1088 | PCL extruder;bio-resorbable line;PLGA maker | {} | Critical moisture/temp control to prevent polymer degradation.
| Vacuum Potting Chamber | Production: Assembly | De-aerated Resin Filler | {"vacuum_level":{"type":"number","units":"mbar"},"chamber_temp":{"type":"number","units":"C"},"pot_life_alarm":{"type":"boolean"}} | TRUE | TRUE | FALSE | 12 | 12 |  | High | ISO 13485;21 CFR 820 | vacuum potting;resin degasser;encapsulator | {} | OQ verifies vacuum pull-down time and leak rate consistency.
| Solder Joint Strength Tester | Metrology | Automated PCBA Pull-Tester | {"force_range":{"type":"range","units":"N"},"speed_range":{"type":"range","units":"mm/s"},"fixture_type":{"type":"string"}} | TRUE | TRUE | TRUE | 12 | 12 |  | Medium | IPC-A-610;ISO 17025 | solder puller;PCBA tester;bond tester | {} | Used to validate SMT reflow profiles per IPC Class 3.
| Precision Bone Plate Bender | Production | CNC Contour Bender | {"bend_angle_acc":{"type":"number","units":"deg"},"max_thickness":{"type":"number","units":"mm"},"material":{"type":"string"}} | TRUE | TRUE | TRUE | 12 | 12 | 24 | Medium | ISO 13485;21 CFR 820.70 | plate bender;ortho bender;trauma tool | {} | PQ verifies absence of surface cracks/nicks during bending.
| Bio-Resorbable Sintering Oven | Material Forming | Low-Temp Controlled Atmos | {"max_temp":{"type":"number","units":"C"},"temp_uniformity":{"type":"number","units":"C"},"gas_type":{"type":"string"}} | TRUE | TRUE | TRUE | 6 | 6 | 12 | High | ISO 13485;ASTM F1983 | PCL sinter;scaffold oven;bio-oven | {} | PQ involves molecular weight (MW) retention testing post-process.
| Atmospheric Plasma Treater | Surface Treatment | In-Line Jet Plasma | {"plasma_power":{"type":"number","units":"W"},"gas_flow":{"type":"number","units":"LPM"},"speed":{"type":"number","units":"m/s"}} | TRUE | TRUE | TRUE | 12 | 6 | 12 | Medium | ISO 13485;21 CFR 820.75 | plasma wand;surface activator;jet plasma | {} | PQ verifies surface energy (dynes) using Dyne pens or goniometer.
| Suture Anchor Induction Heater | Production: Assembly | Precision Insert Bonder | {"frequency":{"type":"number","units":"kHz"},"heat_time":{"type":"number","units":"sec"},"power_output":{"type":"number","units":"W"}} | TRUE | TRUE | TRUE | 6 | 12 | 12 | High | ISO 13485;21 CFR 820.75 | anchor heater;insert bonder;suture welder | {} | PQ verifies pull-out force of the suture from the anchor.
| Active Battery Cycle Tester | Metrology | Implantable Cell Analyzer | {"volt_acc":{"type":"number","units":"mV"},"curr_acc":{"type":"number","units":"uA"},"channels":{"type":"integer"}} | TRUE | TRUE | TRUE | 12 | 12 |  | High | IEC 60601-1;ISO 13485;21 CFR Part 11 | battery cycler;cell tester;pacer battery | {} | Critical for validating long-term reliability of lithium primary cells.
| RF Shielded Test Box | Environment | Faraday Active Device Box | {"shielding_db":{"type":"number","units":"dB"},"freq_range":{"type":"range","units":"GHz"},"interlocks":{"type":"boolean"}} | TRUE | TRUE | FALSE |  | 12 |  | High | IEC 60601-1-2;FDA QMSR | RF box;Faraday cage;pacer test box | {} | OQ verifies attenuation levels at targeted frequencies.
| Pouch Rotary Loader | Production: Packaging | High-Speed Aseptic Loader | {"cycle_speed":{"type":"number","units":"PPM"},"vacuum_acc":{"type":"number","units":"mbar"},"material":{"type":"string"}} | TRUE | TRUE | TRUE | 12 | 12 |  | Medium | ISO 11607;21 CFR 820.75 | pouch loader;product inserter;rotary loader | {} | PQ verifies zero-pinch or fold errors in sterile barrier.
| Catheter Tip Grinder | Production | Precision Mandrel Grinder | {"wheel_rpm":{"type":"number"},"feed_acc":{"type":"number","units":"um"},"coolant":{"type":"boolean"}} | TRUE | TRUE | TRUE | 12 | 6 |  | Medium | ISO 13485;21 CFR 820 | tip grinder;mandrel finisher;catheter bench | {} | PQ verifies atraumatic tip geometry and OD consistency.
| Automated Mandrel Sizer | Production: Auxiliaries | Laser Straightener/Sizer | {"straightness_tol":{"type":"number","units":"um"},"speed":{"type":"number","units":"m/min"},"laser_v":{"type":"string"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Low | ISO 13485;FDA QMSR | core straightener;mandrel sizer;wire sizer | {} | OQ verifies TIR (Total Indicator Reading) of straightened wire.
| Implantable Sensor Cal Bath | Metrology | Precision Thermal Source | {"temp_stability":{"type":"number","units":"C"},"uniformity":{"type":"number","units":"C"},"vol":{"type":"number","units":"L"}} | TRUE | TRUE | TRUE | 6 | 12 |  | High | ISO 17025;FDA QMSR | cal bath;sensor bath;thermal master | {} | Reference standard for qualifying therapeutic heart sensors.
| Laser Suture Driller | Production: Machining | Micro-Hole CO2/UV Laser | {"hole_dia":{"type":"number","units":"um"},"pulse_freq":{"type":"number","units":"kHz"},"depth_ctrl":{"type":"boolean"}} | TRUE | TRUE | TRUE | 12 | 6 | 12 | High | ISO 13485;21 CFR 820.75 | needle driller;suture laser;micro-drill | {} | Critical for swaged needle/suture attachment quality.
| Micro-Injection Insert Molder | Material Forming | Vertical Cleanroom Press | {"clamp_force":{"type":"number","units":"kN"},"shot_size":{"type":"number","units":"g"},"insert_sensor":{"type":"boolean"}} | TRUE | TRUE | TRUE | 12 | 12 | 24 | High | ISO 13485;21 CFR 820.75 | insert molder;vertical press;needle molder | {} | OQ establishes "Insert Detect" safety logic and process window.
| Tissue Valve Sewing Stand | Production | Aseptic Surgical Bench | {"lighting_lux":{"type":"number"},"ergonomics":{"type":"string"},"sterile_surface":{"type":"boolean"}} | TRUE | FALSE | FALSE |  | 12 |  | High | 21 CFR 1271.190;ISO 14644 | heart valve bench;sewing station;HCTP bench | {} | IQ verifies laminar airflow integration and light intensity.
| High-Intensity UV Cure Chamber | Production: Assembly | Batch UV Irradiation | {"uv_intensity":{"type":"number","units":"mW/cm2"},"cooling_fan":{"type":"boolean"},"timer_acc":{"type":"number","units":"sec"}} | TRUE | TRUE | TRUE | 6 | 6 | 12 | High | ISO 13485;21 CFR 820.75 | UV box;adhesive oven;light cure chamber | {} | PQ verifies "Degree of Cure" (e.g., Shore D hardness) across full load.
| Surgical Tool Torque Tester | Metrology | Haptic Feedback Calibrator | {"torque_range":{"type":"range","units":"N-cm"},"angular_acc":{"type":"number","units":"deg"},"peak_hold":{"type":"boolean"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Medium | ISO 17025;FDA QMSR | torque analyzer;tool cal;haptic tester | {} | Used to validate "Click" accuracy of torque-limiting drivers.
| Robotic Aseptic Filling Line | Production: Assembly & Filling | Isolator-Integrated Vial/Syringe | {"fill_range":{"type":"range","units":"mL"},"accuracy":{"type":"number","units":"%"},"cleanroom_class":{"type":"enum","values":}} | TRUE | TRUE | TRUE | 6 | 6 | 12 | High | ISO 13408-1;21 CFR 820.75;21 CFR 1271 | vial filler;syringe robot;aseptic filler | {} | PQ includes media fills to verify 100% aseptic integrity.
| Automated Surgical Instrument Finisher | Production: Finishing | Hemostat/Forceps Polisher | {"belt_speed":{"type":"number","units":"RPM"},"media_type":{"type":"string"},"stations":{"type":"integer"}} | TRUE | TRUE | TRUE | 12 | 12 | 24 | Medium | ISO 13485;FDA QMSR;ISO 17664 | instrument polisher;forceps finisher;hemostat polisher | {} | PQ verifies consistent surface roughness (Ra) and absence of burrs.
| HVAC "Active Smoke" Fogger | Facilities: Validation | High-Output Visualization Fogger | {"smoke_duration":{"type":"number","units":"min"},"particle_safety":{"type":"boolean"},"fog_density":{"type":"string"}} | TRUE | TRUE | FALSE | 12 | 12 | 12 | Low | ISO 14644-3;FDA Guidance (Cleanroom) | smoke generator;cleanroom fogger;HVAC visualizer | {} | Used for "Smoke Studies" to visualize laminar airflow patterns.
| Pharmaceutical Spray Dryer | Production: Material Forming | Drug-Device Combination System | {"inlet_temp":{"type":"range","units":"C"},"feed_rate":{"type":"number","units":"mL/min"},"gas_flow":{"type":"number","units":"LPM"}} | TRUE | TRUE | TRUE | 6 | 6 | 12 | High | ISO 13485;21 CFR Part 11;cGMP | drug-device dryer;micro-particle dryer;GEA dryer | {} | PQ involves particle size distribution and drug-loading stability.
| Ultra-Precision IOL Lathe | Production: Material Forming | Ophthalmic Diamond Turner | {"spindle_speed":{"type":"number","units":"RPM"},"surface_finish":{"type":"number","units":"nm"},"positional_acc":{"type":"number","units":"um"}} | TRUE | TRUE | TRUE | 12 | 6 | 24 | High | ISO 11979-3;21 CFR 886.3600 | intraocular lathe;contact lens lathe;diamond turner | {} | PQ must include diopter accuracy and modulation transfer function (MTF).
| RFID Tag Writer/Encoder | Production: Labeling | UHF/ISO 18000 Encoder | {"freq_range":{"type":"range","units":"MHz"},"protocol":{"type":"enum","values":},"write_speed":{"type":"number","units":"tags/min"}} | TRUE | TRUE | TRUE | 12 | 12 | 12 | Medium | ISO 18000-63;21 CFR 830 | chip writer;smart labeler;UHF encoder | {} | OQ verifies "Write Success" rate and data lock integrity.
| Suture Needle Punch / Borer | Production: Machining | Laser/Mechanical Hole Borer | {"hole_dia":{"type":"range","units":"um"},"stroke_depth":{"type":"number","units":"mm"},"laser_type":{"type":"string"}} | TRUE | TRUE | TRUE | 12 | 12 | 24 | High | ISO 13485;21 CFR 820.75 | borer;suture driller;needle hole maker | {} | PQ verifies hole depth consistency for needle swage pull-out strength.
| Robotic Surgical Instrument Tester | Metrology: Surgical | Haptic Feedback Calibrator | {"force_feedback_res":{"type":"number","units":"mN"},"axial_acc":{"type":"number","units":"um"},"degree_of_freedom":{"type":"integer"}} | TRUE | TRUE | TRUE | 6 | 6 | 12 | High | IEC 60601-1;ASTM F3107 | haptic calibrator;robotic surgical arm tester;da vinci tester | {} | Critical for validating robotic effector accuracy and safety stop.
| Automated Micro-Screwdriver | Production: Assembly | Torque/Angle Control Station | {"torque_range":{"type":"range","units":"Ncm"},"rpm":{"type":"number"},"torque_accuracy":{"type":"number","units":"%"}} | TRUE | TRUE | TRUE | 12 | 12 | 12 | Medium | ISO 13485;21 CFR 820.70 | torque driver;automated screw;assembly tool | {} | OQ verifies "Low-Torque" limits to prevent housing stripping.
| Digital Pathology Slide Scanner | Metrology | High-Res Tissue Digitizer | {"magnification":{"type":"enum","values":["20x","40x"]},"resolution":{"type":"number","units":"um/pixel"},"slide_cap":{"type":"integer"}} | TRUE | TRUE | FALSE | 12 | 12 | 12 | Medium | ISO 13485;21 CFR Part 11 | histology scanner;slide digitizer;pathology bot | {} | Critical for IVD/HCT/P quality checks. OQ verifies color fidelity.
| UV-Vis Spectrophotometer | Analytical Lab | Dual-Beam Scanning Unit | {"wavelength_range":{"type":"range","units":"nm"},"stray_light":{"type":"number","units":"%"},"photometric_acc":{"type":"number","units":"Abs"}} | TRUE | TRUE | FALSE | 12 | 12 | 12 | Medium | USP <1058>;USP <857> | spectrometer;lab UV;light meter | {} | USP <1058> Group B. Calibrated for wavelength accuracy and linearity.
| Karl Fischer Titrator | Analytical Lab | Moisture Analyzer | {"moisture_range":{"type":"range","units":"ppm"},"resolution":{"type":"number","units":"ug"},"method":{"type":"enum","values":["Volumetric","Coulometric"]}} | TRUE | TRUE | FALSE | 12 | 12 | 12 | Medium | USP <1058>;USP <921> | moisture analyzer;titrator;KF water tester | {} | Used to validate residual moisture in lyophilized HCT/P products.
| Dissolution Test System | Analytical Lab | USP Apparatus 1/2 Station | {"bath_temp_acc":{"type":"number","units":"C"},"paddle_rpm":{"type":"number"},"vessel_vol":{"type":"number","units":"mL"}} | TRUE | TRUE | TRUE | 6 | 12 | 12 | Medium | USP <1058>;USP <711> | tablet tester;dissolution bath;paddle tester | {} | PQ verifies release rate of drug-eluting catheters or beads.
| TOC (Total Organic Carbon) Analyzer | Utilities | Online High-Temp Oxidation | {"lod":{"type":"number","units":"ppb"},"response_time":{"type":"number","units":"sec"},"oxidizer_type":{"type":"string"}} | TRUE | TRUE | TRUE | 6 | 12 | 12 | High | USP <643>;USP <1231> | water toc;carbon analyzer;Sievers | {} | Essential utility control for WFI and Purified Water systems.
| Automatic Colony Counter | Laboratory: Micro | High-Res Imaging Counter | {"detection_limit":{"type":"number","units":"um"},"image_res":{"type":"number","units":"MP"},"plate_dia":{"type":"number","units":"mm"}} | TRUE | TRUE | FALSE | 12 | 12 | 12 | Medium | USP <1058>;FDA QMSR | micro counter;bacterial counter;plate counter | {} | OQ verifies separation logic for overlapping bacterial colonies.
| Burst Strength Tester | Metrology: Packaging | Mullen-Style Hydraulic | {"burst_range":{"type":"range","units":"psi"},"diaphragm_dia":{"type":"number","units":"mm"},"data_acc":{"type":"number","units":"%"}} | TRUE | TRUE | FALSE | 12 | 12 | 12 | Medium | ISO 11607-1;ASTM D3786 | Mullen tester;packaging burst;pouch strength | {} | Used to qualify secondary shipping cartons for medical devices.
| Handheld Raman Spectrometer | Metrology | Handheld RMID Unit | {"laser_wavelength":{"type":"number","units":"nm"},"spectral_range":{"type":"range","units":"cm-1"},"laser_safety_class":{"type":"string"}} | TRUE | TRUE | FALSE | 12 | 12 | 12 | Medium | USP <858>;21 CFR Part 11 | handheld raman;RMID;Vaya;Agilent | {} | OQ verifies chemical ID accuracy through clear/amber containers.
| Non-Contact Optical Micrometer | Metrology | High-Speed Laser Scan | {"od_range":{"type":"range","units":"mm"},"scan_rate":{"type":"number","units":"Hz"},"resolution":{"type":"number","units":"um"}} | TRUE | TRUE | FALSE | 12 | 12 | 12 | Medium | ISO 17025;FDA 820 | laser mic;optical gauge;non-contact sizer | {} | Used for real-time diameter monitoring during catheter extrusion.
| Muffle Furnace | Laboratory | High-Temp Ashing Unit | {"temp_max":{"type":"number","units":"C"},"chamber_vol":{"type":"number","units":"L"},"venting":{"type":"boolean"}} | TRUE | TRUE | FALSE | 12 | 12 | 12 | Medium | ISO 13485;21 CFR 820.70 | lab oven;ashing furnace;high temp lab | {} | Used to determine residue on ignition (ROI) for device materials.
| Automated Tissue Microarrayer | HCT/P Processing | Precision Tissue Puncher | {"punch_dia":{"type":"range","units":"mm"},"tray_capacity":{"type":"integer"},"pos_acc":{"type":"number","units":"um"}} | TRUE | TRUE | TRUE | 12 | 12 | 12 | Medium | 21 CFR 1271;ISO 13485 | tissue array;TMA bot;HCTP punch | {} | Used for high-throughput screening of donor tissue biopsy cores.
| Selective Wave Soldering | Production: Electronics | Lead-Free Point-to-Point | {"solder_pot_temp":{"type":"number","units":"C"},"nozzle_dia":{"type":"number","units":"mm"},"drag_speed":{"type":"number","units":"mm/s"}} | TRUE | TRUE | TRUE | 12 | 6 | 12 | High | IPC-7801;J-STD-001;ISO 13485 | selective solder;point solder;mini-wave | {} | PQ requires micro-sectioning of solder joints to verify intermetallic thickness.
| Orthopedic Screw Coater | Production: Surface Treatment | Automated Spin/Spray | {"spin_speed":{"type":"number","units":"RPM"},"spray_pressure":{"type":"number","units":"psi"},"cure_time":{"type":"number","units":"min"}} | TRUE | TRUE | TRUE | 6 | 12 | 12 | High | ISO 13485;ASTM F1147;FDA 820.75 | screw coater;HA coater;bone screw finisher | {} | PQ verifies coating mass and thickness uniformity across 3 batches.
| Cleanroom AGV/AMR | Facilities: Transport | Autonomous Mobile Robot | {"path_accuracy":{"type":"number","units":"mm"},"cleanroom_class":{"type":"integer"},"payload":{"type":"number","units":"kg"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Low | ISO 14644-14;ANSI/RIA R15.08 | cleanroom bot;AMR;material robot;mobile platform | {} | OQ verifies obstacle avoidance logic and "E-Stop" response within cleanroom.
| Precision Needle Grinder | Production: Machining | CNC Multi-Bevel Grinder | {"grinding_angle":{"type":"number","units":"deg"},"wheel_rpm":{"type":"number"},"coolant_flow":{"type":"number","units":"L/min"}} | TRUE | TRUE | TRUE | 12 | 6 | 24 | Medium | ISO 13485;21 CFR 820.70 | cannula grinder;needle maker;point grinder | {} | OQ verifies "Point Geometry" consistency using 3D vision measurement.
| Vial Capping / Crimping | Production: Packaging | Aseptic Aluminum Crimper | {"crimp_force":{"type":"number","units":"N"},"vial_height_acc":{"type":"number","units":"mm"},"reject_logic":{"type":"boolean"}} | TRUE | TRUE | TRUE | 12 | 12 | 12 | High | ISO 13408-1;21 CFR 1271 | vial capper;crimper;seal crimper | {} | PQ involves container closure integrity (CCI) testing after crimping.
| High-Shear Granulator | Reagent Mfg | Wet Reagent Granulator | {"impeller_speed":{"type":"number","units":"RPM"},"chopper_speed":{"type":"number","units":"RPM"},"liquid_feed_rate":{"type":"number","units":"g/min"}} | TRUE | TRUE | TRUE | 12 | 6 | 12 | Medium | 21 CFR 820.75;ISO 13485 | wet granulator;mixer;IVD granulator | {} | PQ verifies final particle size distribution and moisture content.
| 3D Laser Profile Scanner | Metrology | High-Res Surface Map | {"z_resolution":{"type":"number","units":"um"},"xy_resolution":{"type":"number","units":"um"},"scan_speed":{"type":"number","units":"mm/s"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Medium | ISO 10360-8;ISO 17025 | profile scanner;laser map;surface scanner | {} | Used to qualify orthopedic implant surface roughness and texture.
| Automated Suture Braider | Production | Multi-Carrier Fiber Braider | {"mandrel_speed":{"type":"number","units":"RPM"},"carrier_tension":{"type":"number","units":"N"},"picks_per_inch":{"type":"integer"}} | TRUE | TRUE | TRUE | 12 | 6 | 24 | High | ISO 13485;21 CFR 820.75 | fiber braider;suture maker;wire braider | {} | PQ verifies tensile strength and diameter consistency of the braid.
| Bio-Hazard Pass-through | Sterilization | Autoclave/VHP Interlock | {"steam_pressure":{"type":"number","units":"bar"},"cycle_time":{"type":"number","units":"min"},"door_interlock":{"type":"boolean"}} | TRUE | TRUE | TRUE | 6 | 6 | 12 | High | ISO 17665;29 CFR 1910.1030 | waste pass;decon hatch;sterilizer hatch | {} | OQ verifies that only one door can open and seal integrity remains high.
| Vapor Phase Reflow Oven | Production: Electronics | High-Density PCB Reflow | {"fluid_boiling_point":{"type":"number","units":"C"},"preheat_time":{"type":"number","units":"sec"},"pressure":{"type":"number","units":"mbar"}} | TRUE | TRUE | TRUE | 12 | 6 | 12 | High | IPC Class 3;J-STD-020 | vapor reflow;condensation solder;PCB oven | {} | Used for complex medical sensor PCBs. PQ verifies zero-voiding joints.
| Titanium Anodizing Line | Production: Finishing | Type II/III Color Line | {"rectifier_voltage":{"type":"number","units":"V"},"acid_temp":{"type":"number","units":"C"},"dwell_time":{"type":"number","units":"sec"}} | TRUE | TRUE | TRUE | 6 | 12 | 12 | Medium | ISO 13485;21 CFR 820.75 | color anodizer;Ti finisher;ortho line | {} | PQ verifies color uniformity and absence of acid residue (NVR).
| Optical Fusion Splicer | Production: Assembly | Fiber-Optic Catheter Splice | {"splice_loss":{"type":"number","units":"dB"},"tensile_strength":{"type":"number","units":"N"},"alignment_mode":{"type":"string"}} | TRUE | TRUE | TRUE | 12 | 12 |  | High | ISO 13485;FDA 820 | laser catheter splice;fiber splicer;fusion unit | {} | OQ verifies arc power stability and core alignment accuracy.
| Cleanroom Air Velocity Meter | Metrology | Hot-Wire Anemometer | {"range":{"type":"range","units":"fpm"},"accuracy":{"type":"number","units":"%"},"datalogging":{"type":"boolean"}} | TRUE | FALSE | FALSE | 12 |  |  | Low | ISO 14644-3;ISO 17025 | anemometer;airflow probe;hood tester | {} | Calibration must be performed against a NIST-traceable wind tunnel.
| Aseptic Liquid Dosing Robot | Production: Assembly | Robotic Syringe/Vial Doser | {"vol_accuracy":{"type":"number","units":"%"},"cycle_speed":{"type":"number","units":"doses/hr"},"sterile_path":{"type":"boolean"}} | TRUE | TRUE | TRUE | 6 | 6 | 12 | High | ISO 13408-1;21 CFR Part 11 | dosing bot;liquid robot;filling arm | {} | PQ requires "Precision Fill" validation across the full volume range.
| Cleanroom Garment Scanner | Metrology | Particulate Release Tester | {"sensitivity":{"type":"number","units":"um"},"test_duration":{"type":"number","units":"sec"},"sample_flow":{"type":"number","units":"LPM"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Low | IEST-RP-CC003;ISO 14644 | suit scanner;garment tester;particle release | {} | Used to qualify laundry cycles for reusable cleanroom gowns.
| Guidewire Coating Dip Tank | Production: Coating | Automated Hydrophilic Dip | {"withdraw_speed":{"type":"number","units":"mm/s"},"tank_viscosity":{"type":"number","units":"cP"},"dwell_time":{"type":"number","units":"sec"}} | TRUE | TRUE | TRUE | 12 | 12 |  | High | ISO 13485;ASTM F2397 | dip coater;lubricity tank;guidewire line | {} | PQ verifies final lubricity (friction reduction) of the wire.
| Digital Pacer Lead Gauge | Production: Inspection | X-Ray Dimensional Gauge | {"resolution":{"type":"number","units":"um"},"kVp_output":{"type":"number","units":"kV"},"software_v":{"type":"string"}} | TRUE | TRUE | TRUE | 12 | 12 |  | High | 21 CFR 1020;21 CFR Part 11 | lead gauge;X-ray inspector;lead inspector | {} | OQ verifies image sharpness and pixel-to-micron calibration.
| Smart Conveyor Redirecter | Integration | Vision-Triggered Diverter | {"redirect_time":{"type":"number","units":"ms"},"trigger_logic":{"type":"string"},"speed_max":{"type":"number","units":"m/min"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Medium | 21 CFR 820.70;ISO 13485 | diverter;sortation arm;reject gate | {} | OQ verifies "Fail-Safe" reject logic (reject if air pressure drops).
| High-Resolution IOL Sorter | Production: Assembly | Ophthalmic Lens Inserter | {"position_acc":{"type":"number","units":"um"},"force_limit":{"type":"number","units":"N"},"cycle_time":{"type":"number","units":"sec"}} | TRUE | TRUE | TRUE | 12 | 12 |  | High | ISO 11979-3;ISO 13485 | lens sorter;IOL loader;optical assembly | {} | Critical for handling intraocular lenses without edge damage.
| Plasma Surface Activator | Production | Low-Pressure O2 Plasma | {"rf_power":{"type":"number","units":"W"},"gas_pressure":{"type":"number","units":"mbar"},"cycle_temp":{"type":"number","units":"C"}} | TRUE | TRUE | TRUE | 12 | 6 | 12 | Medium | ISO 10993-1;ISO 13485 | plasma cleaner;surface prep;O2 plasma | {} | PQ verifies "Wettability" (Dyne level) on plastic catheter manifolds.
| Laser Plastic Welder | Production: Assembly | Quasi-Simultaneous / Contour | {"laser_power":{"type":"number","units":"W"},"clamp_force":{"type":"number","units":"N"},"scan_speed":{"type":"number","units":"mm/s"}} | TRUE | TRUE | TRUE | 12 | 12 | 12 | High | ISO 13485;21 CFR 820.75;ISO 10993 | plastic laser;ClearWeld;LPKF welder | {} | PQ verifies joint pull-strength and hermeticity on transparent plastics.
| Guidewire Tip Shaper | Production: Finishing | Automated Thermal Former | {"forming_temp":{"type":"number","units":"C"},"mandrel_dia":{"type":"number","units":"mm"},"dwell_time":{"type":"number","units":"sec"}} | TRUE | TRUE | TRUE | 12 | 6 | 24 | High | ISO 13485;21 CFR 820.70;ISO 10555 | tip shaper;J-tip maker;catheter wire shaper | {} | PQ verifies tip geometry (e.g., J-curve) and absence of heat-damage.
| Sterilization Transfer Hatch | Facilities: Decon | VHP/UV-C Pass-through | {"cycle_time":{"type":"number","units":"min"},"h2o2_conc":{"type":"number","units":"ppm"},"door_interlock":{"type":"boolean"}} | TRUE | TRUE | TRUE | 6 | 6 | 12 | High | ISO 14644;ISO 14937;21 CFR 820 | pass-box;clean-side hatch;sterile transfer | {} | PQ involves BI log-reduction testing to verify "Clean-Side" entry safety.
| Induction Cap Sealer | Production: Packaging | High-Frequency Foil Sealer | {"line_speed":{"type":"number","units":"m/min"},"induction_power":{"type":"number","units":"W"},"height_adjustment":{"type":"number","units":"mm"}} | TRUE | TRUE | TRUE | 12 | 12 |  | Medium | ISO 11607;FDA QMSR | cap sealer;foil sealer;Enercon;induction bonder | {} | OQ verifies seal integrity across the full induction field width.
| Rotational Bond Strength Tester | Metrology | Catheter Torque-to-Failure | {"torque_cap":{"type":"number","units":"N-cm"},"rotation_speed":{"type":"number","units":"RPM"},"jaw_alignment":{"type":"number","units":"um"}} | TRUE | TRUE | TRUE | 12 | 12 |  | High | ISO 10555-1;ASTM F2394 | torque tester;bond strength;catheter torque | {} | PQ verifies proximal-to-distal torque transmission reliability.
| Dynamic Mechanical Analyzer (DMA) | Analytical Lab | Viscoelastic Characterization | {"force_range":{"type":"range","units":"N"},"freq_range":{"type":"range","units":"Hz"},"temp_acc":{"type":"number","units":"C"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Medium | USP <1058>;ISO 6721 | DMA;polymer stiffness;viscoelastic tester | {} | Critical for bio-absorbable polymer degradation studies.
| Torsional Fatigue Tester | Metrology | Dynamic Orthopedic Bench | {"torque_load":{"type":"number","units":"N-m"},"cycles":{"type":"integer"},"env_chamber":{"type":"boolean"}} | TRUE | TRUE | TRUE | 6 | 6 | 24 | High | ASTM F1717;ASTM F543 | screw fatigue;spinal fatigue;dynamic torsion | {} | PQ verifies implant survival through >1 million dynamic cycles.
| Hydrostatic Pressure Tester | Metrology | Catheter Burst Stand | {"max_pressure":{"type":"number","units":"psi"},"fluid_temp":{"type":"number","units":"C"},"timer_acc":{"type":"number","units":"sec"}} | TRUE | TRUE | TRUE | 6 | 12 |  | High | ISO 10555;FDA Guidance 824 | burst tester;RBP bench;catheter pressure | {} | Used to validate "Rated Burst Pressure" (RBP) specifications.
| Robotic Pipetting Workstation | Laboratory | Automated IVD Preparation | {"vol_accuracy":{"type":"number","units":"%"},"cv_reproducibility":{"type":"number","units":"%"},"deck_cap":{"type":"integer"}} | TRUE | TRUE | TRUE | 6 | 6 | 12 | High | GAMP 5;21 CFR Part 11;USP <1058> | Hamilton;pipetting robot;liquid handler | {} | PQ requires gravimetric volume verification of robotic heads.
| High-Resolution Digital Microscope | Metrology | 4K Automated Inspection | {"magnification":{"type":"range","units":"x"},"measuring_acc":{"type":"number","units":"um"},"z_stacking":{"type":"boolean"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Medium | ISO 13485;FDA 820 | Keyence microscope;digital scope;metrology scope | {} | OQ verifies calibration against a certified glass scale (NIST).
| Micro-CNC Drill | Production: Machining | High-RPM Suture Needle Borer | {"drill_speed":{"type":"number","units":"RPM"},"feed_acc":{"type":"number","units":"um"},"bit_sensor":{"type":"boolean"}} | TRUE | TRUE | TRUE | 12 | 6 | 24 | High | ISO 13485;21 CFR 820.75 | needle drill;micro-borer;suture punch | {} | PQ verifies hole concentricity for swaged suture reliability.
| Automated Label Rewinder | Production: Labeling | Vision-Integrated Counter | {"rewind_speed":{"type":"number","units":"m/s"},"label_count_acc":{"type":"number","units":"%"},"tension_ctrl":{"type":"boolean"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Low | ISO 13485;FDA 820 | label counter;web rewinder;roll checker | {} | OQ verifies "Missing Label" detection and recount accuracy.
| Nitrogen Air-Curtain Cabinet | Facilities | Aseptic Component Storage | {"n2_purity":{"type":"number","units":"%"},"airflow_vel":{"type":"number","units":"fpm"},"hepa_eff":{"type":"number","units":"%"}} | TRUE | TRUE | FALSE | 12 | 12 | 12 | Medium | ISO 14644-1;FDA QMSR | sterile rack;N2 cabinet;component storage | {} | OQ verifies positive pressure and inert atmosphere stability.
| Ultrasonic Stent Coater | Surface Treatment | DES Polymer/Drug Sprayer | {"nozzle_freq":{"type":"number","units":"kHz"},"droplet_size":{"type":"number","units":"um"},"spray_dist":{"type":"number","units":"mm"}} | TRUE | TRUE | TRUE | 6 | 6 | 12 | High | ISO 13485;FDA 820.75;ISO 10993 | stent coater;ultrasonic sprayer;DES line | {} | PQ verifies drug distribution (mapping) and absence of webbing.
| Glove Integrity Tester | Environment & Safety | Wireless Pressure Decay | {"test_pressure":{"type":"number","units":"Pa"},"lod_sensitivity":{"type":"number","units":"um"},"vessel_cap":{"type":"integer"}} | TRUE | TRUE | FALSE | 12 | 12 |  | High | ISO 14644-7;FDA QMSR | glove tester;isolator checker;leak tester | {} | Mandatory for daily pre-use checks in aseptic HCT/P processing.
| Vacuum Decay Leak Tester | Production: Packaging | Non-Destructive Package Check | {"vacuum_level":{"type":"number","units":"mbar"},"sensitivity":{"type":"number","units":"um"},"chamber_vol":{"type":"number","units":"L"}} | TRUE | TRUE | TRUE | 12 | 12 |  | Medium | ASTM F2338;ISO 11607 | CCI tester;pouch leak;non-destruct leak | {} | OQ verifies "Reject" logic using a master leak standard.
| ICP-MS Elemental Analyzer | Analytical Lab | Heavy Metal Impurity System | {"lod":{"type":"number","units":"ppt"},"mass_resolution":{"type":"number"},"autosampler":{"type":"boolean"}} | TRUE | TRUE | TRUE | 12 | 6 | 12 | High | USP <232>;USP <233>;ICH Q3D | ICPMS;metal tester;impurity analyzer | {} | Used for safety release of implantable devices (elemental risk).
| High-Shear Emulsifier | Reagent Mfg | IVD Reagent Homogenizer | {"shear_rate":{"type":"number","units":"s-1"},"vessel_vol":{"type":"number","units":"L"},"sanitary_seal":{"type":"boolean"}} | TRUE | TRUE | TRUE | 12 | 12 | 12 | Medium | 21 CFR 820.75;ISO 13485 | reagent mixer;Silverson;homogenizer | {} | PQ verifies consistent particle size reduction across batches.
| Sieve Shaker (Analytical) | Analytical Lab | Digital Particle Distribution | {"amplitude":{"type":"number","units":"mm"},"freq_acc":{"type":"number","units":"%"},"timer_res":{"type":"number","units":"sec"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Medium | ISO 13485;FDA 820 | particle sizer;vibratory sieve;mesh shaker | {} | Used to qualify incoming raw material (metal/plastic powders).
| Automated Reagent Bottle Capper | Production: Packaging | Torque-Controlled Rotary | {"torque_range":{"type":"range","units":"lb-in"},"capping_speed":{"type":"number","units":"BPM"},"reject_count":{"type":"boolean"}} | TRUE | TRUE | TRUE | 12 | 12 |  | Medium | ISO 13485;21 CFR 820 | bottle capper;capping bot;IVD capper | {} | OQ verifies torque "Upper/Lower" limits and reject accuracy.
| Multi-Carrier Catheter Braider | Production: Assembly | 16/32/48 Carrier System | {"wire_tension":{"type":"number","units":"N"},"picks_per_inch":{"type":"integer"},"braid_angle":{"type":"number","units":"deg"}} | TRUE | TRUE | TRUE | 12 | 6 | 24 | High | ISO 13485;21 CFR 820.75 | Steeger;braiding line;catheter reinforcement | {} | PQ verifies torque transmission and kink-resistance of the braided shaft.
| High-Speed Component Vision Sorter | Production: Inspection | Micro-Part Automated Sorter | {"detection_speed":{"type":"number","units":"parts/min"},"camera_res":{"type":"number","units":"MP"},"rejection_logic":{"type":"string"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Medium | ISO 13485;21 CFR 820.70;21 CFR Part 11 | part sorter;screw sorter;vision bowl feeder | {} | OQ verifies "Known Bad" part rejection at maximum throughput rates.
| Cleanroom Air Handling Unit (AHU) | Utilities: HVAC | Recirculating Cleanroom Unit | {"airflow_volume":{"type":"number","units":"CFM"},"external_static_press":{"type":"number","units":"in_wg"},"filter_stages":{"type":"integer"}} | TRUE | TRUE | TRUE | 12 | 6 | 12 | High | ISO 14644-1;FDA Guidance (Cleanroom) | cleanroom AHU;HVAC unit;air handler | {} | PQ is equivalent to annual Cleanroom Certification (Particles/Viables).
| Precision Die-Cutting Press | Production: Forming | Medical Adhesive/Bandage Press | {"cutting_force":{"type":"number","units":"kN"},"web_tension":{"type":"number","units":"N"},"registration_acc":{"type":"number","units":"mm"}} | TRUE | TRUE | TRUE | 12 | 12 |  | Medium | ISO 13485;21 CFR 820.70 | die cutter;bandage maker;rotary press | {} | PQ verifies dimensional accuracy of cut shapes and edge quality.
| Metal 3D Printer (L-PBF) | Production: Additive | Laser Powder Bed Fusion | {"laser_power":{"type":"number","units":"W"},"oxygen_level":{"type":"number","units":"ppm"},"layer_thickness":{"type":"number","units":"um"}} | TRUE | TRUE | TRUE | 6 | 6 | 12 | High | ISO 13485;ASTM F2924;ASTM F3001 | metal 3D;SLM printer;titanium printer | {} | PQ involves density testing and coupon tensile bars for every build.
| Tissue Mincer / Grinder | HCT/P Processing | Automated Aseptic Mincer | {"blade_speed":{"type":"number","units":"RPM"},"screen_size":{"type":"number","units":"um"},"sterile_path":{"type":"boolean"}} | TRUE | TRUE | TRUE | 12 | 12 | 12 | High | 21 CFR 1271.200;ISO 13485 | HCTP mincer;bone grinder;tissue processor | {} | PQ verifies consistent particle size and maintenance of tissue viability.
| Optical Fiber Polisher | Production: Finishing | Automated Ferrule Polisher | {"polishing_force":{"type":"number","units":"N"},"cycle_time":{"type":"number","units":"sec"},"pad_grit":{"type":"integer"}} | TRUE | TRUE | TRUE | 12 | 12 |  | Medium | ISO 13485;21 CFR 820 | fiber polisher;laser tip finisher;endoscope polisher | {} | Critical for surgical laser fiber efficiency. OQ verifies tip geometry.
| Hot Isostatic Press (HIP) | Production: Heat Treat | Densification Vessel | {"max_pressure":{"type":"number","units":"bar"},"max_temp":{"type":"number","units":"C"},"soak_time":{"type":"number","units":"min"}} | TRUE | TRUE | TRUE | 6 | 12 | 12 | High | ISO 13485;ASTM F2924 | HIP vessel;densifier;3D print finisher | {} | Used to eliminate internal porosity in 3D-printed metal implants.
| Automated Hemostat Grinder | Production: Machining | Surgical Tool Profile Grinder | {"wheel_speed":{"type":"number","units":"RPM"},"feed_acc":{"type":"number","units":"um"},"fixture_type":{"type":"string"}} | TRUE | TRUE | TRUE | 12 | 12 | 24 | Medium | ISO 13485;ISO 17664 | hemostat grinder;tool sharpener;forceps maker | {} | PQ verifies "Locking Strength" and surface finish of surgical tools.
| Powder Sieve System (Additive) | Production: Auxiliaries | Ultrasonic Raw Material Prep | {"mesh_size":{"type":"number","units":"um"},"ultrasonic_power":{"type":"number","units":"W"},"inert_blanket":{"type":"boolean"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Medium | ISO 13485;ASTM F2924 | powder sifter;additive sifter;metal powder prep | {} | OQ verifies particle size distribution (PSD) of sieved metal powders.
| 3D Digital Surfometer | Metrology | Non-Contact Interferometer | {"z_resolution":{"type":"number","units":"nm"},"xy_pixel_size":{"type":"number","units":"um"},"tilt_acc":{"type":"number","units":"deg"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Medium | ISO 25178;ISO 17025 | white light scope;surface mapper;Ra tester | {} | Used to validate orthopedic implant surface topography (S-parameters).
| Dynamic Light Scattering (DLS) | Analytical Lab | Nano-Particle Sizing System | {"sizing_range":{"type":"range","units":"nm"},"temp_ctrl":{"type":"number","units":"C"},"sensitivity":{"type":"number"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Medium | USP <429>;USP <1058> | DLS;Malvern Zetasizer;nano-sizer | {} | USP <1058> Group B. Used for sizing liposomal drug-combination carriers.
| High-Torque Surgical Reamer | HCT/P Processing | Medullary Canal Prep Tool | {"max_torque":{"type":"number","units":"N-m"},"rpm_accuracy":{"type":"number","units":"%"},"cannulation_dia":{"type":"number","units":"mm"}} | TRUE | TRUE | TRUE | 12 | 12 |  | High | 21 CFR 1271;ISO 17664 | bone reamer;femoral reamer;allograft reamer | {} | PQ focuses on preventing thermal necrosis in allograft tissue.
| Cryogenic Floor Centrifuge | HCT/P Processing | High-Capacity Chilled Separator | {"max_g_force":{"type":"number"},"temp_range":{"type":"range","units":"C"},"imbalance_cutoff":{"type":"boolean"}} | TRUE | TRUE | TRUE | 6 | 12 | 12 | High | 21 CFR 1271;IEC 61010-2-020 | blood centrifuge;component separator;cold spin | {} | Critical for large-scale cord blood or cell therapy processing.
| Flow-through Dissolution | Analytical Lab | USP Apparatus 4 System | {"cell_type":{"type":"string"},"pump_precision":{"type":"number","units":"%"},"temp_uniformity":{"type":"number","units":"C"}} | TRUE | TRUE | TRUE | 6 | 6 | 12 | Medium | USP <711>;USP <1058> | USP 4;elution tester;stent dissolution | {} | Ideal for validating drug release from stent-grafts or drug-eluting beads.
| Active Device Burn-In Station | Production: Electronics | Multi-Slot Thermal Stress | {"temp_setpoint":{"type":"number","units":"C"},"current_monitoring":{"type":"boolean"},"slot_capacity":{"type":"integer"}} | TRUE | TRUE | TRUE | 12 | 12 | 12 | High | ISO 13485;FDA 820.70 | stress rack;aging station;PCB test rack | {} | PQ verifies "Early Life Failure" (infant mortality) reduction in electronics.
| Hemodynamic Simulator | Metrology | Patient Vital Sign Standard | {"pressure_waveforms":{"type":"array"},"ecg_accuracy":{"type":"number","units":"%"},"co2_sim":{"type":"boolean"}} | TRUE | FALSE | FALSE | 12 |  |  | Medium | IEC 60601-2-27;ISO 17025 | patient sim;vital sign master;cal simulator | {} | Reference standard to qualify surgical monitors and therapeutic heart devices.
| Automated Large-Bag Sealer | Production: Packaging | IV / Fluid Bag Sealer | {"seal_width":{"type":"number","units":"mm"},"temp_acc":{"type":"number","units":"C"},"dwell_time":{"type":"number","units":"sec"}} | TRUE | TRUE | TRUE | 12 | 6 | 12 | High | ISO 11607-2;21 CFR 820.75 | IV bag sealer;fluid sealer;large format sealer | {} | PQ involves burst testing and seal integrity on high-volume liquid bags.
| Ultrasonic Mandrel Cleaner | Cleaning | Aseptic Core Wire Washer | {"bath_temp":{"type":"number","units":"C"},"ultrasonic_freq":{"type":"number","units":"kHz"},"wash_steps":{"type":"integer"}} | TRUE | TRUE | TRUE | 12 | 12 | 12 | Medium | ISO 15883;ISO 13485 | mandrel washer;wire cleaner;ultrasonic tank | {} | PQ verifies residual solvent and particulate levels on catheter mandrels.
| Micro-Arc (TIG) Suture Welder | Production: Joining | Needle-Suture Joining Station | {"arc_current":{"type":"number","units":"A"},"gas_shield":{"type":"string"},"electrode_acc":{"type":"number","units":"um"}} | TRUE | TRUE | TRUE | 12 | 12 | 12 | High | ISO 13485;21 CFR 820.75 | suture welder;needle joiner;micro-TIG | {} | PQ verifies tensile strength (USP pull test) for 3 consecutive batches.
| Automated Tablet Press | Production: Drug-Device | Rotary Multi-Station | {"stations":{"type":"integer"},"compression_force":{"type":"number","units":"kN"},"ejection_force":{"type":"number","units":"N"}} | TRUE | TRUE | TRUE | 6 | 6 | 12 | High | ISO 13485;21 CFR 820.75;cGMP | pill maker;punching machine;tablet stamper | {} | PQ verifies tablet weight uniformity, hardness, and friability.
| X-Ray Photoelectron Spectrometer (XPS) | Metrology: Surface Science | High-Vacuum Surface Analyzer | {"energy_res":{"type":"number","units":"eV"},"vacuum_base":{"type":"number","units":"mbar"},"source_type":{"type":"string"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Medium | ISO 13485;ISO 17025;FDA QMSR | XPS;ESCA;surface chemist;implant analyzer | {} | Critical for qualifying passivation layers and HA coating chemistry.
| Cleanroom Modular Wall System | Facilities: Infrastructure | Hard-Wall Modular Partition | {"surface_material":{"type":"string"},"outgassing_class":{"type":"string"},"fire_rating":{"type":"string"}} | TRUE | TRUE | FALSE |  | 12 |  | Low | ISO 14644-1;FDA QMSR;NFPA 75 | cleanroom wall;modular partition;ISO wall | {} | OQ verifies room airtightness and differential pressure stability.
| Aseptic IV Bag Filling Machine | Production: Filling | Isolator-Integrated Automated | {"fill_acc":{"type":"number","units":"%"},"bag_range":{"type":"range","units":"mL"},"vhp_integrated":{"type":"boolean"}} | TRUE | TRUE | TRUE | 6 | 6 | 12 | High | ISO 13408-1;21 CFR 820.75;21 CFR 1271 | IV filler;bag filler;vial filler;saline machine | {} | PQ requires "Media Fills" to verify 100% aseptic integrity.
| IV Bag Automated Visual Inspector | Production: Inspection | Multi-Camera Particle Detection | {"lod_sensitivity":{"type":"number","units":"um"},"camera_count":{"type":"integer"},"speed":{"type":"number","units":"bags/min"}} | TRUE | TRUE | TRUE | 12 | 12 | 12 | High | ISO 13485;21 CFR Part 11;USP 790 | AVI;bag inspector;particle checker | {} | PQ uses "Knapp-test" bags with known sub-visible contaminants.
| Scanning Acoustic Microscope (SAM) | Metrology: NDT | High-Freq Ultrasonic Imaging | {"acoustic_freq":{"type":"number","units":"MHz"},"spatial_res":{"type":"number","units":"um"},"immersion_fluid":{"type":"string"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Medium | ISO 13485;ASTM F601 | SAM;acoustic scope;ultrasonic image;delamination scope | {} | Used to detect internal voids in active implantable hermetic seals.
| Microneedle Coating System | Production: Coating | Micron-Scale Dip/Spray | {"dipping_speed":{"type":"number","units":"mm/s"},"droplet_vol":{"type":"number","units":"uL"},"temp_acc":{"type":"number","units":"C"}} | TRUE | TRUE | TRUE | 6 | 12 | 12 | High | ISO 13485;FDA 820.75;ASTM F2100 | microneedle coater;array coater;micro-dip | {} | PQ verifies drug load uniformity per needle across the patch array.
| Drug-Eluting Microsphere Coater | Production: Coating | Microsphere Encapsulator | {"polymer_conc":{"type":"number","units":"%"},"batch_vol":{"type":"number","units":"mL"},"solvent_type":{"type":"string"}} | TRUE | TRUE | TRUE | 12 | 6 | 12 | High | ISO 13485;21 CFR 820.75;USP 1112 | bead coater;DEB coater;sphere maker | {} | PQ verifies release rate and particle size distribution consistency.
| Atomic Force Microscope (AFM) | Metrology: Surface Science | High-Res Topography Probe | {"xy_resolution":{"type":"number","units":"nm"},"z_resolution":{"type":"number","units":"nm"},"probe_mode":{"type":"string"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Medium | ISO 13485;ISO 17025;ASTM E2859 | AFM;nano-scope;surface probe;nano-topography | {} | Used to qualify nano-textured implant surfaces for bone growth.
| In-Circuit Test (ICT) Bench | Production: Electronics | Bed-of-Nails Automated | {"test_points":{"type":"integer"},"voltage_max":{"type":"number","units":"V"},"accuracy":{"type":"number","units":"%"}} | TRUE | TRUE | TRUE | 12 | 12 |  | High | IPC-A-610;ISO 13485;21 CFR 820.70 | bed of nails;ICT;PCBA circuit tester | {} | Critical for active implants; OQ verifies all test vector coverage.
| Functional Test (FCT) Bench | Production: Electronics | Device Parameter Simulator | {"test_params":{"type":"array"},"comm_bus":{"type":"enum","values":},"cycle_time":{"type":"number","units":"sec"}} | TRUE | TRUE | TRUE | 12 | 12 | 12 | High | IEC 60601-1;21 CFR 820.70;ISO 13485 | FCT;device tester;implant simulator | {} | PQ validates real-world firmware response to sensor stimulus.
| Flying Probe PCBA Tester | Production: Electronics | Fixtureless Automated Tester | {"probe_acc":{"type":"number","units":"um"},"speed":{"type":"number","units":"probes/sec"},"test_coverage":{"type":"number","units":"%"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Medium | IPC-A-610;ISO 13485 | flying probe;Takaya;PCBA tester | {} | Used for high-mix/low-volume medical boards where ICT is not feasible.
| Solderability Tester | Metrology: Electronics | Wetting Balance Dip-Look | {"pot_temp":{"type":"number","units":"C"},"force_res":{"type":"number","units":"mN"},"dwell_acc":{"type":"number","units":"sec"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Medium | J-STD-002;ISO 13485;MIL-STD-883 | wetting balance;solder tester;component checker | {} | Used to qualify incoming leads/components for life-support PCBs.
| X-Ray Fluorescence (XRF) | Metrology: Material | Benchtop RoHS/Metal ID | {"elements_detected":{"type":"array"},"lod_limit":{"type":"number","units":"ppm"},"xray_acc":{"type":"number","units":"%"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Medium | ISO 13485;ISO 17025;RoHS | XRF;RoHS tester;metal ID;elemental analyzer | {} | Critical for identifying material mix-ups in implantable metals.
| PIND Testing System | Metrology: NDT | Particle Impact Noise Detector | {"acceleration":{"type":"number","units":"g"},"frequency":{"type":"number","units":"Hz"},"db_threshold":{"type":"number"}} | TRUE | TRUE | FALSE | 12 | 12 |  | High | MIL-STD-883 Method 2020;ISO 13485 | PIND;particle noise;hermetic tester;loose part detector | {} | Used to detect loose conductive particles inside active implantable cans.
| High-Pot / Dielectric Tester | Metrology: Safety | Dielectric Strength Analyzer | {"test_voltage":{"type":"range","units":"kV"},"leakage_acc":{"type":"number","units":"uA"},"ramp_rate":{"type":"number","units":"V/s"}} | TRUE | TRUE | FALSE | 12 | 12 |  | High | IEC 60601-1;UL 60601-1;ISO 17025 | hipot;safety tester;insulation tester | {} | Mandatory for 100% release testing of patient-contact active devices.
| Nitrogen Glovebox | Environment | Aseptic Reagent Formulation | {"o2_ppm":{"type":"number","units":"ppm"},"h2o_ppm":{"type":"number","units":"ppm"},"chamber_pressure":{"type":"number","units":"Pa"}} | TRUE | TRUE | TRUE | 12 | 12 | 12 | Medium | ISO 14644-7;FDA QMSR;21 CFR 1271 | formulation box;inert glovebox;dry box | {} | Critical for anaerobic HCT/P or moisture-sensitive IVD formulation.
| SMT Stencil Cleaning System | Production: Electronics | Automated Aqueous/Solvent | {"wash_cycle":{"type":"number","units":"min"},"cleaning_agent":{"type":"string"},"stencil_size":{"type":"string"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Low | IPC-7526;ISO 13485 | stencil washer;paste cleaner;PCB cleaner | {} | OQ verifies removal of solder paste from fine-pitch apertures.
| Component Laser Trimmer | Production: Machining | Active Resistor/Cap Adjuster | {"trim_acc":{"type":"number","units":"%"},"spot_size":{"type":"number","units":"um"},"laser_type":{"type":"string"}} | TRUE | TRUE | TRUE | 12 | 6 | 24 | High | ISO 13485;21 CFR 820.70 | laser trimmer;resistor trim;thin film laser | {} | Used to tune circuit parameters in life-critical sensor assemblies.
| Walk-In Stability Chamber | Testing: Stability | ICH Temperature/Humidity | {"temp_range":{"type":"range","units":"C"},"humidity_range":{"type":"range","units":"%RH"},"int_vol":{"type":"number","units":"m3"}} | TRUE | TRUE | TRUE | 6 | 6 | 12 | High | ICH Q1A;FDA QMSR;ISO 13485 | stability room;aging chamber;ICH room | {} | PQ involves comprehensive 3D mapping of temp/humidity uniformity.
| Automated Suture Spooler | Assembly | Precision Spooling System | {"winding_tension":{"type":"number","units":"N"},"spool_width":{"type":"number","units":"mm"},"pitch_acc":{"type":"number","units":"um"}} | TRUE | TRUE | TRUE | 12 | 12 | 24 | High | ISO 13485;21 CFR 820.75 | suture winder;spool bot;thread winder | {} | OQ verifies intelligent motor conversion of electronic commands to motion. [1]
| Heart Valve Leaflet Flex Tester | Metrology | Dynamic Durability Stand | {"flex_frequency":{"type":"number","units":"Hz"},"fluid_temp":{"type":"number","units":"C"},"target_cycles":{"type":"integer"}} | TRUE | TRUE | TRUE | 6 | 6 | 12 | High | ISO 5840-1;ISO 13485 | leaflet tester;valve flexer;tissue fatigue | {} | Durability testing often required beyond 200 million cycles.
| Mobile UV-C Disinfection Tower | Sterilization | Mobile Germicidal Unit | {"uv_dosage":{"type":"number","units":"mJ/cm2"},"room_sensors":{"type":"integer"},"cycle_time":{"type":"number","units":"min"}} | TRUE | TRUE | TRUE | 12 | 6 | 12 | High | ISO 14937;FDA QMSR | UV-C bot;room decon;disinfection tower | {} | Used for terminal cleanroom disinfection; requires biological indicator kill validation. [2, 3]
| Cannula Electrochemical Polisher | Finishing | Tip Sharpening Bath | {"acid_temp":{"type":"number","units":"C"},"current_density":{"type":"number","units":"A/dm2"},"dwell_time":{"type":"number","units":"sec"}} | TRUE | TRUE | TRUE | 6 | 12 | 12 | Medium | ISO 13485;ASTM B912 | cannula polisher;tip finisher;needle polisher | {} | PQ verifies atraumatic tip geometry and surface Ra < 0.1 µm.
| Guidewire Weld Strength Tester | Metrology | Tensile Breaking Bench | {"force_range":{"type":"range","units":"N"},"test_speed":{"type":"number","units":"mm/min"},"fixture_type":{"type":"string"}} | TRUE | TRUE | TRUE | 12 | 12 |  | High | ASTM F2516;ISO 10555 | weld puller;breaking bench;wire tester | {} | Used for destructive testing of core-to-coil welds in catheters.
| IOL Centration Metrology Gauge | Metrology | Interferometric Lens Mapper | {"diopter_acc":{"type":"number","units":"D"},"axis_acc":{"type":"number","units":"deg"},"resolution":{"type":"number","units":"um"}} | TRUE | TRUE | FALSE | 12 | 12 |  | High | ISO 11979-2;ISO 17025 | IOL centration;lens mapper;diopter gauge | {} | Critical for verifying optical center offset in toric/multifocal lenses.
| HCT/P Cryogenic Shaker | HCT/P Processing | Variable Speed Bioreactor | {"temp_range":{"type":"range","units":"C"},"shaking_rate":{"type":"number","units":"RPM"},"cycle_timer":{"type":"number","units":"sec"}} | TRUE | TRUE | TRUE | 6 | 12 | 12 | High | 21 CFR 1271;FDA QMSR | cryo mixer;bio-shaker;HCTP culture | {} | PQ verifies homogenization of additives without damaging cell membranes.
| Laminar Airflow Workstation | Facilities | ISO Class 5 Clean Bench | {"face_velocity":{"type":"number","units":"fpm"},"filter_eff":{"type":"number","units":"%"},"uv_timer":{"type":"boolean"}} | TRUE | TRUE | TRUE | 12 | 12 | 12 | High | ISO 14644-1;ISO 14644-3 | clean bench;laminar hood;aseptic bench | {} | PQ involves annual filter integrity and airflow visualization (smoke) tests. [4]
| Syringe Plunger Force Tester | Metrology | Piston Friction Analyzer | {"breakloose_force":{"type":"number","units":"N"},"glide_force":{"type":"number","units":"N"},"stroke":{"type":"number","units":"mm"}} | TRUE | TRUE | TRUE | 12 | 12 |  | Medium | ISO 7886-1;ISO 13485 | glide force;syringe tester;plunger bench | {} | Evaluates the performance of silicone lubricants on syringe barrels.
| Luer Connector Verification Gauge | Metrology | ISO 80369-7 Master Gauge | {"taper_acc":{"type":"number","units":"um"},"surface_finish":{"type":"string"},"thread_form":{"type":"string"}} | TRUE | FALSE | FALSE | 12 |  |  | Medium | ISO 80369-7;ISO 17025 | luer gauge;connector master;ISO 80369 | {} | Used as the reference standard to qualify injection molded connectors.
| PEEK Implant Vacuum Degasser | Production | Resin Moisture Removal | {"vacuum_level":{"type":"number","units":"mbar"},"tray_temp":{"type":"number","units":"C"},"chamber_vol":{"type":"number","units":"L"}} | TRUE | TRUE | TRUE | 12 | 12 |  | Medium | ISO 13485;21 CFR 820.70 | peek dryer;resin degasser;material oven | {} | Critical to prevent hydrolytic degradation during molding.
| Orthopedic Saw Blade Grinder | Machining | CNC Tooth Sharpener | {"pitch_acc":{"type":"number","units":"um"},"wheel_rpm":{"type":"number"},"coolant":{"type":"boolean"}} | TRUE | TRUE | TRUE | 12 | 12 | 24 | Medium | ISO 13485;21 CFR 820 | blade grinder;ortho sharpener;saw grinder | {} | OQ verifies tooth profile consistency against the master drawing. [5]
| Micro-Electrode Pinner | Assembly | Robotic Active Lead Joiner | {"pin_placement_acc":{"type":"number","units":"um"},"force_feedback":{"type":"boolean"},"vision_align":{"type":"boolean"}} | TRUE | TRUE | TRUE | 12 | 6 | 12 | High | ISO 13485;21 CFR 820 | lead pinner;electrode robot;active joiner | {} | High-complexity system for attaching active implantable electrodes.
| Aseptic Component Buffer Tank | HCT/P Processing | Stainless Jacketed Storage | {"vol_total":{"type":"number","units":"L"},"temp_uniformity":{"type":"number","units":"C"},"cip_ready":{"type":"boolean"}} | TRUE | TRUE | TRUE | 12 | 12 | 12 | High | 21 CFR 1271;FDA QMSR | buffer tank;HCTP storage;aseptic vessel | {} | PQ involves media hold-time validation for 7-14 days. [4]
| HEPA Filter Photometer | Metrology | Aerosol Leak Tester | {"particle_sensitivity":{"type":"number","units":"um"},"sample_flow":{"type":"number","units":"LPM"},"datalogging":{"type":"boolean"}} | TRUE | FALSE | FALSE | 12 |  |  | Medium | ISO 14644-3;FDA 820 | photometer;filter tester;PAO tester | {} | Calibration verifies linear response to aerosol concentrations.
| Multi-Lumen Flow Verifier | Metrology | Simultaneous Channel Tester | {"pressure_per_lumen":{"type":"array"},"target_flow":{"type":"number","units":"mL/min"},"cycle_time":{"type":"number","units":"sec"}} | TRUE | TRUE | TRUE | 12 | 12 |  | High | ISO 10555;ISO 13485 | flow bench;lumen tester;patency checker | {} | OQ verifies channel-to-channel cross-talk prevention logic.
| Drug-Resorbable Scaffold Sinterer | Material Forming | Controlled Atmosphere Furnace | {"max_temp":{"type":"number","units":"C"},"temp_acc":{"type":"number","units":"C"},"gas_mix":{"type":"string"}} | TRUE | TRUE | TRUE | 6 | 6 | 12 | High | ISO 13485;ASTM F1983 | PCL sinter;bio-oven;combo furnace | {} | PQ involves molecular weight (IV) and drug-loading stability tests.
| PIND Testing Stand | Metrology | Particle Impact Noise Detector | {"acceleration_range":{"type":"range","units":"g"},"freq_acc":{"type":"number","units":"%"},"software_v":{"type":"string"}} | TRUE | TRUE | FALSE | 12 | 12 |  | High | MIL-STD-883 M2020;ISO 13485 | PIND;particle noise;hermetic tester | {} | OQ verifies detection of 100 µm gold wire challenge standard.
| IVD Lateral Flow Slitter Bench | Production | Card Slitting System | {"cut_width_tol":{"type":"number","units":"mm"},"throughput":{"type":"number","units":"cards/hr"},"blade_life":{"type":"integer"}} | TRUE | TRUE | TRUE | 12 | 12 |  | Medium | ISO 13485;FDA QMSR | card slitter;LFA slitter;strip cutter | {} | PQ verifies strip width consistency to ensure fit in plastic housings.
| Cleanroom Pressure Gauge (Master) | Metrology | Precision Manometer | {"accuracy_class":{"type":"number","units":"%"},"range":{"type":"range","units":"in_wg"},"traceability":{"type":"string"}} | TRUE | FALSE | FALSE | 12 |  |  | Medium | ISO 14644;ISO 17025 | master manometer;pressure standard;DP gauge | {} | Standard to qualify cleanroom differential pressure monitoring systems. [6]
| X-Ray Powder Diffractometer (XRPD) | Analytical Lab | Crystallinity Analyzer | {"2theta_range":{"type":"range","units":"deg"},"step_size":{"type":"number","units":"deg"},"phase_id_accuracy":{"type":"number","units":"%"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Medium | BS ISO 13779-2;USP <1058> | XRD;crystallinity tester;HA phase analyzer | {} | Used to qualify hydroxyapatite (HA) coating phase purity.
| Energy-Dispersive Spectroscope (EDS) | Analytical Lab | Elemental Mapping Unit | {"detection_limit":{"type":"number","units":"wt%"},"elements_range":{"type":"array"},"integration_v":{"type":"string"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Medium | ISO 13485;ASTM E1508 | EDS;EDX;elemental mapper;composition analyzer | {} | Integrated with SEM for contaminant identification and alloy verification.
| MES (Mfg Execution System) | Control | Paperless Lot Tracking | {"user_access_levels":{"type":"integer"},"database_redundancy":{"type":"boolean"},"software_version":{"type":"string"}} | TRUE | TRUE | TRUE |  |  | 12 | High | 21 CFR Part 11;21 CFR 820.70(i);GAMP 5 | eDHR;lot tracker;digital batch record | {} | Critical for electronic Device History Records (eDHR) and traceability.
| LIMS (Lab Information Mgmt) | Control | Quality Control Database | {"sample_tracking_acc":{"type":"number","units":"%"},"audit_trail_log":{"type":"boolean"},"instrument_interface":{"type":"array"}} | TRUE | TRUE | TRUE |  |  | 12 | High | 21 CFR Part 11;ISO 17025;GAMP 5 | lab software;QC database;sample manager | {} | Used for electronic data capture from qualified lab instruments.
| Orthopedic Screw Washer-Dryer | Cleaning | Automated Basket System | {"cycle_temp":{"type":"number","units":"C"},"detergent_conc":{"type":"number","units":"%"},"ultrasonic_integrated":{"type":"boolean"}} | TRUE | TRUE | TRUE | 12 | 6 | 12 | Medium | ISO 15883;AAMI ST98 | screw washer;implant cleaner;passivation washer | {} | PQ verifies residual oil removal (TOC/NVR) using representative loads.
| Wireless Thermal Mapping Kit | Metrology | Multi-Channel Logger System | {"temp_accuracy":{"type":"number","units":"C"},"logger_count":{"type":"integer"},"radio_freq":{"type":"number","units":"GHz"}} | TRUE | FALSE | FALSE | 6 |  |  | Medium | ISO 17025;ISO 17665;USP <1058> | Kaye kit;thermal mapping;validation loggers | {} | Reference standard to qualify autoclaves, freezers, and ovens.
| Automated Microbiological Spreader | Lab: Micro | Spiral Plater / Spreader | {"spread_pattern":{"type":"enum","values":},"dilution_ratio":{"type":"number"},"plate_cap":{"type":"integer"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Medium | ISO 13485;FDA QMSR | agar spreader;micro plater;colony spreader | {} | OQ verifies uniform distribution of environmental samples on agar.
| Laser Anemometer | Metrology | Non-Invasive Air Velocity | {"velocity_range":{"type":"range","units":"m/s"},"response_time":{"type":"number","units":"ms"},"nist_traceable":{"type":"boolean"}} | TRUE | FALSE | FALSE | 12 |  |  | Low | ISO 14644-3;ISO 17025 | air flow probe;cleanroom anemometer;laminar meter | {} | Standard to qualify cleanroom laminar flow and BSC face velocities.
| Active Smoke Generator | Facilities | Cleanroom Air Visualizer | {"smoke_purity":{"type":"string"},"particle_safety":{"type":"boolean"},"tank_vol":{"type":"number","units":"L"}} | TRUE | FALSE | FALSE |  | 12 |  | Low | ISO 14644-3;FDA Guidance (Cleanroom) | smoke wand;HVAC visualizer;airflow fogger | {} | Used for mandatory "Smoke Studies" to prove laminar protection.
| Automated J-Tip Former | Production | Guidewire Thermal Shaper | {"mold_temp":{"type":"number","units":"C"},"wire_tension":{"type":"number","units":"N"},"cycle_speed":{"type":"number","units":"parts/min"}} | TRUE | TRUE | TRUE | 12 | 6 | 24 | High | ISO 13485;21 CFR 820.75;ISO 10555 | tip shaper;J-tip bot;guidewire maker | {} | PQ verifies curvature radius and tensile integrity of the distal tip.
| IPG Programmer Bench | Active Implantable | Pacemaker Testing Station | {"telemetry_range":{"type":"number","units":"m"},"pulse_freq_acc":{"type":"number","units":"%"},"battery_sim":{"type":"boolean"}} | TRUE | TRUE | TRUE | 6 | 12 |  | High | IEC 60601-1;21 CFR 870.1750 | pacer programmer;IPG tester;heart device bench | {} | Critical for qualifying final firmware load on active heart devices.
| Hemodialysis Tube Extruder | Material Forming | High-Speed PVC/TPU Line | {"output_rate":{"type":"number","units":"m/min"},"wall_thick_acc":{"type":"number","units":"um"},"sizing_vacuum":{"type":"boolean"}} | TRUE | TRUE | TRUE | 12 | 6 | 12 | High | ISO 13485;21 CFR 820.75;ISO 10993 | dialysis line;blood tubing maker;extruder | {} | PQ verifies surface smoothness and kink resistance for blood circuits.
| Endoscope Leak Tester | Inspection | Automated Pressure Decay | {"test_pressure":{"type":"number","units":"kPa"},"leak_lod":{"type":"number","units":"mL/min"},"internal_battery":{"type":"boolean"}} | TRUE | TRUE | TRUE | 12 | 12 |  | High | ISO 15883-4;ISO 13485 | scope leak;AER tester;internal seal checker | {} | Mandatory to ensure distal seal integrity before high-level disinfection.
| BP Monitor Calibrator | Metrology | Sphygmomanometer Master | {"pressure_acc":{"type":"number","units":"mmHg"},"sim_waveforms":{"type":"array"},"hose_conn":{"type":"integer"}} | TRUE | FALSE | FALSE | 12 |  |  | Medium | ISO 81060-2;ISO 17025 | NIBP master;BP tester;cuff calibrator | {} | Standard to qualify digital and manual blood pressure monitors.
| Multi-Axis Tissue Bioreactor | HCT/P Processing | Dynamic Scaffold Stimulator | {"tensile_force":{"type":"number","units":"N"},"flex_freq":{"type":"number","units":"Hz"},"medium_temp_acc":{"type":"number","units":"C"}} | TRUE | TRUE | TRUE | 6 | 6 | 12 | High | 21 CFR 1271.160;ISO 13485 | tissue bot;tendon grower;scaffold bioreactor | {} | PQ involves biochemical analysis of cell/scaffold integration levels.
| Automated Slide Scanner | Metrology | Digital Pathology Station | {"pixel_resolution":{"type":"number","units":"um"},"focus_acc":{"type":"number","units":"%"},"batch_cap":{"type":"integer"}} | TRUE | TRUE | FALSE | 12 | 12 | 12 | Medium | ISO 13485;21 CFR Part 11 | digital slide;pathology bot;histology scanner | {} | OQ verifies color fidelity and "Out-of-Focus" error detection logic.
| Cleanroom Garment HEPA Vac | Cleaning | Class 10/ISO 4 Approved | {"hepa_eff":{"type":"number","units":"%"},"airflow":{"type":"number","units":"CFM"},"exhaust_particle_acc":{"type":"number"}} | TRUE | FALSE | FALSE |  | 12 |  | Low | ISO 14644-1;NSF 49 | HEPA vacuum;cleanroom vac;gowning vac | {} | IQ verifies exhaust filtration does not compromise cleanroom particles.
| HA Coating Thickness Gauge | Metrology | Eddy Current / Ultrasonic | {"range":{"type":"range","units":"um"},"accuracy":{"type":"number","units":"um"},"probe_dia":{"type":"number","units":"mm"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Medium | ISO 13779-2;ISO 17025 | thickness gauge;HA meter;coating meter | {} | Used to qualify hydroxyapatite layer depth on orthopedic implants.
| Gowning Door Interlock | Control | Aseptic Buffer Logic | {"dwell_time":{"type":"number","units":"sec"},"pressure_trigger":{"type":"number","units":"Pa"},"fail_safe":{"type":"boolean"}} | TRUE | TRUE | FALSE |  | 12 |  | High | ISO 14644-1;FDA QMSR | airlock logic;cleanroom interlock;door controller | {} | OQ verifies "Both Doors Closed" logic to prevent cleanroom collapse.
| Nitrogen Dew-Point Meter | Metrology | Utilities Purity Probe | {"dp_range":{"type":"range","units":"C"},"pressure_max":{"type":"number","units":"bar"},"accuracy":{"type":"number","units":"C"}} | TRUE | FALSE | FALSE | 12 |  |  | Medium | ISO 8573-1;ISO 17025 | moisture probe;N2 purity;dry gas meter | {} | Reference standard to qualify nitrogen/air dryers for manufacturing.
| Orthopedic Laser Imprinter | Finishing | Fiber Laser Marker | {"wavelength":{"type":"number","units":"nm"},"pulse_width":{"type":"number","units":"ns"},"marking_area":{"type":"string"}} | TRUE | TRUE | TRUE | 12 | 12 |  | Medium | 21 CFR 830;ISO 13485 | UDI imprinter;implant marker;laser etcher | {} | PQ verifies UDI scannability and contrast on Ti/SS alloys.
| Catheter Luer-Leak Tester | Metrology | ISO 80369-7 Bench | {"test_pressure":{"type":"number","units":"kPa"},"leak_sensitivity":{"type":"number","units":"mL/min"},"cycle_timer":{"type":"number","units":"sec"}} | TRUE | TRUE | TRUE | 12 | 12 |  | High | ISO 80369-7;ISO 10555 | hub tester;luer leak;connector bench | {} | OQ verifies pressure decay accuracy against calibrated orifices.
| WFI Vapor Compression Still | Utilities | Large-Scale Distiller | {"output_capacity":{"type":"number","units":"L/hr"},"feed_water_cond":{"type":"number","units":"uS/cm"},"temp_acc":{"type":"number","units":"C"}} | TRUE | TRUE | TRUE | 6 | 6 | 12 | High | USP ;ISO 13485 | VC still;WFI maker;vapor compressor | {} | Phase 1 validation (2-4 weeks) monitors distillate daily.
| Multiple-Effect WFI Still | Utilities | Multi-Column Distiller | {"column_count":{"type":"integer"},"steam_consumption":{"type":"number","units":"kg/hr"},"conductivity_cutoff":{"type":"number","units":"uS/cm"}} | TRUE | TRUE | TRUE | 6 | 6 | 12 | High | USP ;ISO 13485 | multi-effect;WFI still;sterile water plant | {} | PQ focuses on pyrogen removal efficiency and steam quality.
| Orthopedic Vibratory Finisher | Finishing | Centrifugal Disc Polisher | {"motor_rpm":{"type":"number"},"bowl_volume":{"type":"number","units":"L"},"amplitude":{"type":"number","units":"mm"}} | TRUE | TRUE | TRUE | 12 | 12 | 24 | Medium | ISO 13485;21 CFR 820.70 | vibratory bowl;implant tumbler;disc finisher | {} | PQ verifies consistent mass finish (Ra) across bulk batches.
| Reagent Vessel CIP Skid | Cleaning | Automated Loop Station | {"rinse_volume":{"type":"number","units":"L"},"cond_control":{"type":"number","units":"uS/cm"},"pump_flow":{"type":"number","units":"LPM"}} | TRUE | TRUE | TRUE | 12 | 12 | 12 | High | ISO 15883;AAMI ST98 | loop washer;CIP unit;vessel cleaner | {} | PQ verifies 100% removal of active reagents via riboflavin test.
| LFA Reagent Nano-Dispenser | IVD Production | Piezo-Electric Striper | {"drop_vol":{"type":"number","units":"nL"},"line_width_acc":{"type":"number","units":"um"},"alignment_acc":{"type":"number","units":"um"}} | TRUE | TRUE | TRUE | 6 | 12 | 12 | High | ISO 13485;FDA QMSR | reagent striper;LFA bot;BioDot bench | {} | PQ verifies coefficient of variation (CV) for protein load.
| Process Air Heatless Dryer | Utilities | Twin-Tower Desiccant | {"outlet_dew_point":{"type":"number","units":"C"},"cycle_time":{"type":"number","units":"min"},"press_drop":{"type":"number","units":"psi"}} | TRUE | TRUE | FALSE | 12 | 12 | 12 | Medium | ISO 8573-1;FDA 820 | air dryer;desiccant tower;CDA dryer | {} | OQ verifies dew point stability during tower switchover.
| WFI Sanitization UV Unit | Utilities | High-Intensity 254nm | {"uv_dosage":{"type":"number","units":"mW/cm2"},"lamp_life_alarm":{"type":"boolean"},"flow_limit":{"type":"number","units":"LPM"}} | TRUE | TRUE | TRUE | 6 | 12 | 12 | High | USP ;ISO 13485 | water UV;sanitization unit;UV lamp | {} | OQ verifies irradiance mapping within the reactor chamber.
| High-Force Implant Pull-Tester | Metrology | Ortho Screw Bench | {"axial_force_cap":{"type":"number","units":"kN"},"crosshead_acc":{"type":"number","units":"um"},"sampling_rate":{"type":"number","units":"Hz"}} | TRUE | TRUE | TRUE | 12 | 12 |  | High | ASTM F543;ISO 17025 | pull tester;screw bench;breaking station | {} | Used for destructive testing of orthopedic screw insertion torque.
| Bone Plate Contour Scanner | Metrology | 3D Laser Profilometer | {"xy_resolution":{"type":"number","units":"um"},"z_accuracy":{"type":"number","units":"um"},"travel_x_y":{"type":"string"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Medium | ISO 10360-8;ISO 17025 | contour tracer;3D scanner;plate mapper | {} | OQ verifies scan accuracy against certified step-artifacts.
| Heart Valve Leaflet Gauge | Metrology | Non-Contact Optical | {"resolution":{"type":"number","units":"um"},"measuring_vol":{"type":"string"},"software_v":{"type":"string"}} | TRUE | TRUE | FALSE | 12 | 12 |  | High | ISO 5840-1;ISO 17025 | leaflet gauge;tissue gauge;valve bench | {} | Critical for verifying tissue thickness uniformity in bioprotheses.
| Automated Syringe Silconizer | Assembly | Spray/Wipe Deposition | {"silicone_mass":{"type":"number","units":"mg"},"nozzle_pos_acc":{"type":"number","units":"um"},"viscosity_range":{"type":"range"}} | TRUE | TRUE | TRUE | 6 | 12 | 12 | High | ISO 7886-1;ISO 13485 | syringe coater;lubricant spray;piston coater | {} | PQ verifies glide force and break-loose force consistency.
| Cleanroom Active Pass-box | Facilities | Fan-Filtered Interlock | {"hepa_eff":{"type":"number","units":"%"},"air_velocity":{"type":"number","units":"fpm"},"interlock_logic":{"type":"string"}} | TRUE | TRUE | FALSE | 12 | 12 | 12 | Medium | ISO 14644-1;FDA QMSR | active pass;HEPA hatch;cleanroom pass | {} | OQ verifies pressure recovery time and door logic safety.
| HCT/P Tissue Morselizer | HCT/P Processing | Automated Bone Grinder | {"morsel_size":{"type":"range","units":"mm"},"motor_torque":{"type":"number","units":"Nm"},"sterile_housing":{"type":"boolean"}} | TRUE | TRUE | TRUE | 12 | 12 | 12 | High | 21 CFR 1271.200;ISO 13485 | bone grinder;morselizer;HCTP mill | {} | PQ verifies consistent granule size for bone void fillers.
| Mandrel Straightening Bench | Production | Guidewire Core Bench | {"straightness_acc":{"type":"number","units":"um/m"},"tension_cap":{"type":"number","units":"N"},"mandrel_dia":{"type":"range"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Low | ISO 13485;FDA QMSR | wire straightener;core bench;mandrel rig | {} | OQ verifies TIR (Total Indicator Reading) reduction on core wires.
| Marker Band Swaging Tool | Assembly | Catheter Band Joiner | {"swage_dia_acc":{"type":"number","units":"um"},"force_limit":{"type":"number","units":"N"},"die_type":{"type":"string"}} | TRUE | TRUE | TRUE | 12 | 12 |  | High | ISO 13485;21 CFR 820.75 | band swager;catheter joiner;swage tool | {} | PQ verifies band retention force and catheter lumen patency.
| Central Gas Distribution Skid | Utilities | Bulk Aseptic Supply | {"gas_purity":{"type":"number","units":"%"},"static_pressure":{"type":"number","units":"psi"},"manifold_v":{"type":"string"}} | TRUE | TRUE | FALSE | 12 | 12 | 12 | Medium | ISO 7396;FDA 820 | bulk gas;gas farm;distribution skid | {} | OQ verifies leak rate and purity at terminal use-points.
| Robotic Implant Deburring | Machining | CNC Surface Polisher | {"tool_force":{"type":"number","units":"N"},"path_repeatability":{"type":"number","units":"um"},"cycle_time":{"type":"number","units":"sec"}} | TRUE | TRUE | TRUE | 12 | 12 |  | Medium | ISO 13485;21 CFR 820.70 | deburr bot;implant polisher;robotic finishing | {} | PQ verifies removal of machining burrs without part distortion.
| Active Lead Fatigue Simulator | Metrology | Accelerated Stress Rig | {"flex_angle":{"type":"number","units":"deg"},"cycle_count":{"type":"integer"},"saline_temp":{"type":"number","units":"C"}} | TRUE | TRUE | TRUE | 6 | 6 | 24 | High | ISO 14708-2;ISO 13485 | lead flexer;pacer stress;fatigue rig | {} | PQ involves continuous electrical monitoring during stress cycling.
| Catheter Trackability Rig | Metrology | Tortuous Path Simulator | {"insertion_speed":{"type":"number","units":"mm/s"},"path_geometry":{"type":"string"},"force_sensitivity":{"type":"number","units":"mN"}} | TRUE | TRUE | TRUE | 12 | 12 |  | High | ISO 10555-1;FDA Guidance 824 | IDTE;track tester;pushability rig;tortuous path | {} | PQ verifies "Mean Insertion Force" across 3 consecutive batches.
| Orthopedic Tray Inkjet Coder | Finishing | Large-Character Industrial | {"ink_durability":{"type":"string"},"print_res":{"type":"number","units":"dpi"},"barcode_acc":{"type":"number","units":"%"}} | TRUE | TRUE | TRUE | 12 | 12 |  | Medium | 21 CFR 830;ISO 13485 | tray imprinter;inkjet coder;sterilizable marker | {} | PQ verifies ink adhesion after repeated autoclave cycles (ISO 17664).
| Automated Pacer Header Joiner | Assembly | Precision Laser/Adhesive | {"placement_acc":{"type":"number","units":"um"},"cure_temp":{"type":"number","units":"C"},"uv_power":{"type":"number","units":"W"}} | TRUE | TRUE | TRUE | 6 | 12 |  | High | ISO 13485;IEC 60601-1 | header bonder;pacer assembly;active joiner | {} | Critical for hermeticity of active implant connectors.
| Cryogenic Bag Labeler | Labeling | Low-Temp Adhesive Station | {"label_adhesion":{"type":"number","units":"N/cm"},"operating_temp":{"type":"number","units":"C"},"barcode_grade":{"type":"string"}} | TRUE | TRUE | TRUE | 12 | 12 |  | High | 21 CFR 1271.290;ISO 13485 | cryo labeler;bag marker;HCTP labeler | {} | PQ verifies label integrity at -196°C (liquid nitrogen phase).
| Automated Slide De-stainer | Lab: Histology | Precision Solvent Station | {"bath_timer_acc":{"type":"number","units":"sec"},"solvent_recycle":{"type":"boolean"},"slides_per_batch":{"type":"integer"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Medium | ISO 13485;FDA QMSR | slide washer;histology cleaner;HCTP prep | {} | OQ verifies removal of excess background stain for IHC clarity.
| Cleanroom DP Server | Control | Multi-Room Pressure Monitor | {"sensor_count":{"type":"integer"},"alarm_log_redundancy":{"type":"boolean"},"scan_rate":{"type":"number","units":"Hz"}} | TRUE | TRUE | FALSE | 12 |  | 12 | High | ISO 14644-2;21 CFR Part 11 | FMS server;pressure database;cleanroom monitor | {} | Software must be validated to ensure audit trail for pressure excursions.
| Active Lead Abrasion Rig | Metrology | Dynamic Lead Insulation | {"stroke_freq":{"type":"number","units":"Hz"},"contact_force":{"type":"number","units":"N"},"saline_bath_temp":{"type":"number","units":"C"}} | TRUE | TRUE | TRUE | 6 | 6 | 24 | High | ISO 14708-2;ASTM F732 | insulation tester;pacer wear rig;abrasion tester | {} | PQ simulates 10 years of cardiac motion (360 million cycles).
| Diamond Turning Lathe (IOL) | Machining | Ultra-Precision Optical | {"surface_finish":{"type":"number","units":"Ra-nm"},"spindle_runout":{"type":"number","units":"um"},"pos_acc":{"type":"number","units":"nm"}} | TRUE | TRUE | TRUE | 12 | 6 | 24 | High | ISO 11979-3;ISO 13485 | lens lathe;IOL turner;optical CNC | {} | PQ verifies diopter accuracy and MTF (Modulation Transfer Function).
| Needle Bevel Vision Sizer | Metrology | 360-Degree Point Mapper | {"geometric_acc":{"type":"number","units":"um"},"scan_time":{"type":"number","units":"sec"},"camera_count":{"type":"integer"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Medium | ISO 13485;FDA 820 | point profiler;needle vision;bevel mapper | {} | OQ verifies bevel angle and lancet point sharpness repeatability.
| Suture Spool Tension Meter | Metrology | Inline Dynamic Sensor | {"tension_range":{"type":"range","units":"cN"},"sampling_freq":{"type":"number","units":"Hz"},"nist_traceable":{"type":"boolean"}} | TRUE | FALSE | FALSE | 12 |  |  | Medium | ISO 17025;FDA QMSR | thread tension;spool meter;suture gauge | {} | Standard to qualify automated suture winders during OQ/PQ.
| DES Solvent Recovery Unit | Utilities | Closed-Loop Aseptic Still | {"purity_threshold":{"type":"number","units":"%"},"exhaust_safe":{"type":"boolean"},"flow_rate":{"type":"number","units":"L/hr"}} | TRUE | TRUE | TRUE | 12 | 12 |  | Medium | ISO 13485;21 CFR 820 | solvent still;drug reclaimer;DES utility | {} | PQ verifies zero cross-contamination between different drug batches.
| Reagent Stopper-Inserter | Assembly | Vacuum-Assisted Capper | {"insertion_force":{"type":"number","units":"N"},"vacuum_acc":{"type":"number","units":"mbar"},"sterile_path":{"type":"boolean"}} | TRUE | TRUE | TRUE | 12 | 12 | 12 | High | ISO 13408-1;21 CFR 1271 | vial stopper;aseptic capper;IVD sealer | {} | PQ involves helium leak testing for container closure integrity (CCI).
| Qualified Walk-off Mat | Facilities | Cleanroom Entry Adhesive | {"peel_strength":{"type":"number","units":"N"},"particle_retention":{"type":"number","units":"%"},"surface_res":{"type":"number","units":"Ohm"}} | TRUE | FALSE | TRUE |  | 12 |  | Low | ISO 14644;FDA QMSR | tacky mat;cleanroom rug;entry mat | {} | PQ verifies particulate reduction in gowning area vs baseline.
| WFI Double-Tube-Plate Heat Exchanger | Utilities | Sanitary Thermal Transfer | {"heat_load":{"type":"number","units":"kW"},"delta_press_max":{"type":"number","units":"psi"},"leak_detect":{"type":"boolean"}} | TRUE | TRUE | TRUE | 12 | 12 |  | High | USP ;ASME BPE | sanitary HE;WFI cooler;sterile heater | {} | IQ verifies double-tube-plate design to prevent cross-contamination.
| Pacer Telemetry Range Tester | Metrology | RF Distance Bench | {"freq_range":{"type":"range","units":"MHz"},"snr_acc":{"type":"number","units":"dB"},"sim_tissue_loss":{"type":"number","units":"dB"}} | TRUE | TRUE | FALSE | 12 | 12 |  | High | IEC 60601-1-2;ISO 13485 | RF range;telemetry tester;pacer bench | {} | OQ verifies stable data link across the maximum clinical distance.
| Femoral Head Impact Tester | Metrology | Dynamic Drop-Weight Stand | {"drop_height":{"type":"number","units":"m"},"impact_force_acc":{"type":"number","units":"%"},"fixture_angle":{"type":"integer"}} | TRUE | TRUE | TRUE | 12 | 12 |  | High | ISO 7206-10;ISO 11491 | impact tester;ceramic head tester;drop stand | {} | PQ verifies survival of ceramic heads under surgical "miss-hit" simulation.
| Catheter Compatibility Bench | Metrology | Multi-Device Fitting Jig | {"luer_acc":{"type":"number","units":"um"},"taper_match":{"type":"boolean"},"swiveling_torque":{"type":"number"}} | TRUE | TRUE | TRUE | 12 | 12 |  | Medium | ISO 80369-7;ISO 11070 | guide-wire fit;compatibility jig;catheter match | {} | OQ verifies smooth passage of guidewire through catheter lumen.
| Automated Reagent Labeler | Labeling | Wraparound Vial Labelling | {"vpm_speed":{"type":"number"},"label_pos_acc":{"type":"number","units":"mm"},"vision_read":{"type":"boolean"}} | TRUE | TRUE | TRUE | 12 | 12 |  | Medium | 21 CFR 830;21 CFR Part 11 | vial labeler;IVD bottle labeler;batch labeler | {} | PQ verifies UDI legibility on rounded small-diameter vials.
| Tissue Valve Fixation Tank | HCT/P Processing | Glutaraldehyde Aseptic Tank | {"vessel_vol":{"type":"number","units":"L"},"temp_acc":{"type":"number","units":"C"},"agitation_acc":{"type":"number"}} | TRUE | TRUE | TRUE | 12 | 12 | 12 | High | 21 CFR 1271;ISO 5840 | fixation tank;crosslinking vessel;heart valve tank | {} | PQ verifies consistent tissue collagen cross-linking via shrink-temp.
| Cleanroom Hands-Free Intercom | Facilities | Flush-Mount Aseptic Audio | {"noise_cancel_db":{"type":"number"},"material_finish":{"type":"string"},"cleanability":{"type":"string"}} | TRUE | TRUE | FALSE |  | 12 |  | Low | ISO 14644-1;FDA QMSR | sterile phone;cleanroom audio;aseptic comms | {} | IQ verifies chemical resistance to VHP and bleach-based wipes.
| Cardiac Pulsatile Flow Generator | Cardiovascular Support | CPB Fluid Dynamics Sim | {"flow_pulse_rate":{"type":"range","units":"bpm"},"stroke_volume":{"type":"number","units":"mL"},"mean_pressure":{"type":"number","units":"mmHg"}} | TRUE | TRUE | TRUE | 12 | 6 | 12 | High | 21 CFR 870.4330;ISO 13485 | pulsatile pump;flow generator;cardiac sim | {} | Used to validate extracorporeal circuits and oxygenators under pulsatile load.
| Motorized Imaging C-Arm | Support: Imaging | Cath Lab Isocentric Arm | {"tilt_range":{"type":"range","units":"deg"},"angulation_speed":{"type":"number","units":"deg/s"},"anticollision":{"type":"boolean"}} | TRUE | TRUE | FALSE | 12 | 12 |  | High | 21 CFR 1020.30;IEC 60601-2-43 | C-arm;motorized gantry;imaging arm | {} | OQ verifies cranial/caudal and lateral angulation accuracy and safety stop.
| Low-Attenuation Carbon Fiber Table | Support: Imaging | Angiographic Procedural Table | {"max_patient_weight":{"type":"number","units":"kg"},"float_travel_x_y":{"type":"string"},"attenuation_factor":{"type":"number"}} | TRUE | TRUE | FALSE |  | 12 |  | Medium | 21 CFR 870.1360;ISO 13485 | cath table;angiography table;X-ray table | {} | OQ verifies motorized vertical movement and electromagnetic blocking logic.
| VA-ECMO Console | Cardiovascular Support | Venoarterial Bypass System | {"battery_reserve":{"type":"number","units":"min"},"flow_limit_alarm":{"type":"number","units":"L/min"},"integrated_blender":{"type":"boolean"}} | TRUE | TRUE | TRUE | 12 | 6 | 12 | High | 21 CFR 870.4150;IEC 60601-1 | ECMO;heart-lung support;bypass console | {} | Critical for cardiogenic shock. PQ involves 24-hour continuous run stability.
| IVUS Imaging Console | Diagnostic: Imaging | Intravascular Ultrasound | {"transducer_freq":{"type":"number","units":"MHz"},"pullback_speed":{"type":"number","units":"mm/s"},"image_res":{"type":"number","units":"um"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Medium | 21 CFR 870.2310;21 CFR Part 11 | IVUS;ultrasound console;vessel mapper | {} | OQ verifies real-time radiological synchronization and pullback accuracy.
| Optical Coherence Tomography (OCT) Station | Diagnostic: Imaging | High-Res Cardiac Imaging | {"scan_dia":{"type":"number","units":"mm"},"axial_res":{"type":"number","units":"um"},"frame_rate":{"type":"number","units":"fps"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Medium | 21 CFR 870.1230;ISO 13485 | OCT;cardiac OCT;vessel scanner | {} | Used for stent apposition analysis. OQ verifies image sharpness and noise levels.
| Intramedullary Nail Fatigue Rig | Metrology: Mechanical | Dynamic Bone Fixation Stand | {"axial_load":{"type":"number","units":"kN"},"flexion_moment":{"type":"number","units":"Nm"},"freq_acc":{"type":"number"}} | TRUE | TRUE | TRUE | 6 | 6 | 24 | High | ISO 5838-1;ISO 14630 | nail tester;IM nail fatigue;ortho nail rig | {} | PQ verifies survival under cyclic loading (e.g., 1 million cycles).
| Glenoid Locking Mechanism Shear Bench | Metrology: Mechanical | Static Anatomic Shoulder Stand | {"shear_force_cap":{"type":"number","units":"N"},"displacement_acc":{"type":"number","units":"um"},"fixture_angle":{"type":"integer"}} | TRUE | TRUE | TRUE | 12 | 12 |  | High | ISO 11491;ISO 13485 | glenoid shear;shoulder tester;joint lock tester | {} | Used for destructive testing of the locking mechanism in shoulder implants.
| Wrist Implant Arthroplasty Simulator | Metrology: Mechanical | Multi-Axis Joint Motion Stand | {"flex_ext_range":{"type":"range","units":"deg"},"radial_ulnar_dev":{"type":"range","units":"deg"},"cycle_rate":{"type":"number"}} | TRUE | TRUE | TRUE | 12 | 6 | 24 | High | ASTM F1357;ISO 14242 | wrist simulator;joint wear;wrist fatigue | {} | PQ involves gravimetric wear assessment over 2 million cycles.
| Suture Anchor Pull-out Tester | Metrology: Mechanical | Cyclic Soft-Tissue Bench | {"insertion_torque":{"type":"number","units":"Nm"},"pull_displacement":{"type":"number","units":"mm"},"peak_load":{"type":"number","units":"N"}} | TRUE | TRUE | TRUE | 12 | 12 |  | High | ASTM F3690;ISO 13485 | anchor puller;suture anchor tester;soft tissue test | {} | PQ verifies the clinical resistance to anchor "pull-out" from bone.
| Modular Implant Taper Disassembly Stand | Metrology: Mechanical | Axial Connection Bench | {"disassembly_force":{"type":"number","units":"N"},"crosshead_speed":{"type":"number","units":"mm/min"},"stroke":{"type":"number","units":"mm"}} | TRUE | TRUE | TRUE | 12 | 12 |  | High | ASTM F2009;ISO 13485 | taper tester;hip modular bench;disassembly rig | {} | Used to validate the axial disassembly force of modular orthopedic joints.
| Anesthesia Machine Calibration Analyzer | Metrology: Gas | High-Speed Vapor Analyzer | {"agent_id":{"type":"array"},"conc_acc":{"type":"number","units":"%"},"resp_rate_acc":{"type":"number"}} | TRUE | FALSE | FALSE | 6 |  |  | Medium | ISO 80601-2-13;ISO 17025 | anesthesia cal;vaporizer tester;gas analyzer | {} | Reference standard to qualify surgical anesthesia delivery systems.
| Breathing Circuit Integrity Bench | Metrology | Low-Pressure Leak Tester | {"test_pressure":{"type":"number","units":"cmH2O"},"leak_lod":{"type":"number","units":"mL/min"},"cycle_time":{"type":"number","units":"sec"}} | TRUE | TRUE | TRUE | 12 | 12 |  | High | ISO 5367;FDA QMSR | circuit leak;breathing tube tester;vent circuit bench | {} | PQ verifies 100% leak-free integrity of patient breathing circuits.
| Pulse Oximeter Functional Calibrator | Metrology: Electronics | Patient Simulator (SpO2) | {"spo2_range":{"type":"range","units":"%"},"pulse_acc":{"type":"number","units":"BPM"},"r_curve_support":{"type":"boolean"}} | TRUE | FALSE | FALSE | 12 |  |  | Medium | ISO 80601-2-61;ISO 17025 | spo2 simulator;oximeter cal;r-curve master | {} | Used to qualify the accuracy of pulse oximetry sensors and monitors.
| Bio-Conjugation Reaction Vessel | Reagent Mfg | Temp-Controlled Glass Reactor | {"vessel_vol":{"type":"number","units":"L"},"jacket_temp_acc":{"type":"number","units":"C"},"stirring_torque":{"type":"number","units":"Ncm"}} | TRUE | TRUE | TRUE | 12 | 12 | 12 | Medium | ISO 13485;21 CFR 820.75 | conjugation tank;antibody reactor;IVD vessel | {} | PQ verifies yield consistency for gold-conjugate or protein-binding.
| Automated Luer-Lock Taper Sizer | Metrology | Laser Dimensional Hub Gauge | {"taper_acc":{"type":"number","units":"um"},"concentricity":{"type":"number","units":"um"},"cycle_speed":{"type":"number"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Low | ISO 80369-7;ISO 17025 | luer sizer;hub gauge;laser hub checker | {} | High-speed in-process inspection of molded catheter hubs.
| Small-Bone Fracture Fixation Rig | Metrology: Mechanical | Metallic Plate Fatigue Stand | {"span_length":{"type":"number","units":"mm"},"flexural_stiffness":{"type":"number","units":"N/mm"},"sinusoidal_load":{"type":"boolean"}} | TRUE | TRUE | TRUE | 12 | 12 | 24 | High | ASTM F3437;ISO 13485 | plate fatigue;small bone rig;fracture stand | {} | PQ verifies survival of small fracture plates through 1 million cycles.
| Ceramic Head Impact Resistance Stand | Metrology | Dynamic Drop-Weight Tower | {"drop_height":{"type":"number","units":"m"},"impact_force":{"type":"number","units":"kN"},"sample_angle":{"type":"integer"}} | TRUE | TRUE | TRUE | 12 | 12 |  | High | ISO 11491;ASTM F2345 | ceramic head tester;femoral head impact;drop tower | {} | Used to validate ceramic femoral heads against surgical "miss-hit" impacts.
| Spinal Nucleus Replacment Simulator | Metrology: Mechanical | Flexural Fatigue Bench | {"flex_angle":{"type":"number","units":"deg"},"axial_compression":{"type":"number","units":"N"},"saline_immersion":{"type":"boolean"}} | TRUE | TRUE | TRUE | 6 | 6 | 24 | High | ISO 18192-2;ASTM F2789 | nucleus sim;spinal disc wear;intervertebral tester | {} | PQ requires gravimetric wear analysis over 5 million cycles.
| Ultrasonic Stent Expansion Bench | Assembly: Inspection | In-Situ Deployment Vision | {"expansion_dia_acc":{"type":"number","units":"um"},"radial_force_res":{"type":"number","units":"mN"},"immersion_tank":{"type":"boolean"}} | TRUE | TRUE | FALSE | 12 | 12 |  | High | ISO 25539-2;FDA 820 | stent expansion;radial force bench;deployment vision | {} | OQ verifies vision measurement accuracy in aqueous (37°C) environments.
| Robotic Surgical Effector Tester | Metrology: Surgical | Force/Motion Precision Stand | {"force_res":{"type":"number","units":"mN"},"axial_acc":{"type":"number","units":"um"},"DOF":{"type":"integer"}} | TRUE | TRUE | TRUE | 6 | 6 | 12 | High | IEC 60601-1;ASTM F3107 | surgical robot tester;da vinci tester;effector cal | {} | Critical for validating robotic instrument precision and safety stops.
| Endoscope Image Fiber Aligner | Production: Assembly | Fiber Optic Connector Bench | {"alignment_acc":{"type":"number","units":"um"},"insertion_loss":{"type":"number","units":"dB"},"vision_v":{"type":"string"}} | TRUE | TRUE | TRUE | 12 | 12 |  | High | ISO 13485;ISO 8600-1 | fiber aligner;scope optics bench;endoscope assembler | {} | PQ verifies optical resolution and color fidelity of the scope.
| Passive Air Sampler (Settle Plates) | Environment: Monitoring | Microbiology Exposure Stand | {"plate_dia":{"type":"number","units":"mm"},"exposure_time":{"type":"number","units":"hr"},"sterile_path":{"type":"boolean"}} | TRUE | FALSE | TRUE |  |  |  | High | ISO 14698;FDA Guidance | settle plate;passive monitor;micro sampler | {} | PQ involves validating the microbial recovery rate in classified areas.
| Contact Lens Hydration Tank | Production: Ophthalmic | Saline Buffering Station | {"ph_acc":{"type":"number"},"temp_acc":{"type":"number","units":"C"},"vessel_vol":{"type":"number","units":"L"}} | TRUE | TRUE | TRUE | 6 | 12 | 12 | Medium | ISO 18369;ISO 13485 | lens hydrator;saline tank;contact lens prep | {} | PQ verifies final lens water content and dimensional stability.
| Orthopedic Implant Casting Furnace | Production: Forming | Investment Casting Unit | {"melt_temp":{"type":"number","units":"C"},"vacuum_acc":{"type":"number","units":"mbar"},"cooling_rate":{"type":"number"}} | TRUE | TRUE | TRUE | 6 | 6 | 12 | High | ISO 13485;ASTM F75 | CoCr furnace;implant caster;metal furnace | {} | PQ involves grain size analysis and alloy composition (ICP) checks.
| Automated Suture Needle Swager | Production: Assembly | High-Speed Pneumatic Joiner | {"swage_force":{"type":"number","units":"N"},"die_alignment":{"type":"number","units":"um"},"stroke_depth":{"type":"number"}} | TRUE | TRUE | TRUE | 12 | 12 | 12 | High | ISO 13485;21 CFR 820.75 | swage machine;needle attacher;suture joiner | {} | PQ involves mandatory USP pull-strength testing.
| Cleanroom Particle Counter (Liquid) | Metrology | HCT/P / IVD Particle Analyzer | {"size_range":{"type":"array"},"flow_rate":{"type":"number","units":"mL/min"},"lod":{"type":"number"}} | TRUE | TRUE | FALSE | 12 | 12 |  | High | USP ;ISO 13485 | liquid counter;vial particle tester;sub-visible analyzer | {} | Critical for validating cleanliness of sterile HCT/P containers.
| Dialysis Filter Integrity Tester | Metrology | Diffusion/Bubble Point Bench | {"test_pressure":{"type":"number","units":"bar"},"flow_sensitivity":{"type":"number","units":"sccm"},"vol_acc":{"type":"number"}} | TRUE | TRUE | TRUE | 12 | 12 |  | High | ISO 8637;FDA QMSR | dialyzer tester;integrity bench;filter leak | {} | OQ verifies the sensitivity to detect individual broken fibers.
| Hospital Cart Washer (Industrial) | Cleaning | Large-Format Spray Booth | {"wash_temp":{"type":"number","units":"C"},"spray_impact":{"type":"number","units":"N"},"cycle_record":{"type":"boolean"}} | TRUE | TRUE | TRUE | 12 | 6 | 12 | Medium | ISO 15883-7;AAMI ST98 | cart washer;booth cleaner;hospital washer | {} | PQ involves soil removal validation on surgical transport carts.
| MRI Implant Heating Simulator | Metrology: Safety | RF/Gradient Phantom Bench | {"sar_acc":{"type":"number","units":"W/kg"},"temp_res":{"type":"number","units":"C"},"phantom_vol":{"type":"number"}} | TRUE | TRUE | FALSE | 12 | 12 |  | High | ASTM F2182;ISO 10974 | MRI phantom;heating tester;safety simulator | {} | Used to qualify the MR-conditional safety of new metal implants.
| Orthopedic Bone Plate Bending Rig | Metrology: Mechanical | 3-Point Flexural Fatigue | {"load_range":{"type":"range","units":"N"},"span_length":{"type":"number","units":"mm"},"sampling_rate":{"type":"number"}} | TRUE | TRUE | TRUE | 12 | 12 |  | High | ASTM F382;ISO 13485 | plate bender tester;fatigue rig;bone plate fatigue | {} | PQ verifies survival of plates through 1 million flexural cycles.
| Drug-Eluting Balloon (DEB) Pleater | Production: Assembly | Thermal Folding Bench | {"pleat_pressure":{"type":"number","units":"psi"},"fold_temp":{"type":"number","units":"C"},"dwell_acc":{"type":"number"}} | TRUE | TRUE | TRUE | 6 | 12 | 12 | High | ISO 10555-4;FDA Guidance 1545 | balloon folder;DEB pleater;cardiac folder | {} | PQ verifies drug coating integrity during mechanical folding.
| Automated Micro-Slide Cover-slipper | Lab: Histology | Vacuum/Adhesive Applicator | {"mounting_vol":{"type":"number","units":"uL"},"slide_throughput":{"type":"number"},"bubble_detect":{"type":"boolean"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Low | ISO 13485;FDA QMSR | cover-slipper;histology bot;HCTP slide prep | {} | OQ verifies consistent resin dispensing and zero-bubble mounting.
| Nitinol Af Temperature Tester | Metrology | Bend and Free Recovery (BFR) | {"temp_ramp_rate":{"type":"number","units":"C/min"},"sensor_acc":{"type":"number","units":"C"},"fluid_vol":{"type":"number"}} | TRUE | TRUE | TRUE | 6 | 12 |  | High | ASTM F2082;ISO 13485 | nitinol tester;Af rig;BFR bench | {} | PQ verifies the Austenite finish temperature of shape-set alloys.
| Cleanroom Wall/Ceiling HEPA Vacuum | Facilities: Cleaning | ISO 4/5 Approved Filter Vac | {"hepa_efficiency":{"type":"number","units":"%"},"static_lift":{"type":"number","units":"inH2O"},"air_flow":{"type":"number"}} | TRUE | FALSE | FALSE |  | 12 |  | Low | ISO 14644-1;NSF 49 | HEPA vac;cleanroom vacuum;ISO vac | {} | IQ verifies exhaust filtration does not compromise cleanroom particles.
| Blood Pressure Cuff Leak Bench | Metrology | Differential Pressure Tester | {"test_pressure":{"type":"number","units":"mmHg"},"leak_lod":{"type":"number","units":"mmHg/min"},"cycle_time":{"type":"number"}} | TRUE | TRUE | TRUE | 12 | 12 |  | Medium | ISO 81060-1;FDA QMSR | cuff tester;BP leak bench;sphygmomanometer tester | {} | PQ verifies 100% leak-free integrity of therapeutic cuffs.
| IVD Bio-Printing Dispenser | Production: IVD | Piezo / Solenoid Inkjet | {"dispense_acc":{"type":"number","units":"%"},"drop_vol":{"type":"number","units":"nL"},"viscosity_max":{"type":"number"}} | TRUE | TRUE | TRUE | 6 | 12 | 12 | High | ISO 13485;FDA 820.75 | bio-jet;antibody printer;LFA printer | {} | PQ verifies protein activity retention after high-frequency jetting.
| Orthopedic Reamer Calibrator | Metrology | Rotational Torque Master | {"torque_acc":{"type":"number","units":"%"},"rpm_res":{"type":"number"},"sampling_rate":{"type":"number"}} | TRUE | FALSE | FALSE | 12 |  |  | Medium | ISO 17025;FDA 820 | reamer cal;ortho tool tester;rotary master | {} | Reference standard for qualifying surgical rotary tools.
| Guidewire Core Wire Grinder | Production: Machining | Micro-Diameter Centerless | {"od_tolerance":{"type":"number","units":"um"},"length_tol":{"type":"number","units":"mm"},"coolant_filter":{"type":"boolean"}} | TRUE | TRUE | TRUE | 12 | 6 | 24 | Medium | ISO 13485;21 CFR 820.70 | core grinder;guidewire maker;micro-grinder | {} | PQ verifies OD taper consistency across bulk production lots.
| Passive Particle Settle Plate Stand | Environment | Fixed Location Aseptic Stand | {"height_from_floor":{"type":"number","units":"mm"},"shielding":{"type":"boolean"},"material":{"type":"string"}} | TRUE | FALSE | FALSE |  | 12 |  | Low | ISO 14644-1;FDA Guidance | settle stand;micro plate holder;cleanroom stand | {} | IQ verifies positioning in "at-rest" and "in-operation" zones.
| Nitinol Af Temperature Tester | Metrology | ASTM F2082 Bench | {"temp_range":{"type":"range","units":"C"},"ramp_rate":{"type":"number","units":"C/min"},"sensor_acc":{"type":"number","units":"C"}} | TRUE | TRUE | TRUE | 6 | 12 |  | High | ASTM F2082;ISO 13485 | nitinol tester;Af rig;BFR bench | {} | PQ verifies the Austenite finish temperature of shape-set alloys.
| Cord Blood Automated Thawer | HCT/P Processing | Waterless Cell Thawer | {"thaw_time":{"type":"number","units":"sec"},"temp_uniformity":{"type":"number","units":"C"},"vial_cap":{"type":"integer"}} | TRUE | TRUE | TRUE | 12 | 12 | 12 | High | 21 CFR 1271;ISO 13485 | cell thawer;dry thaw;cord blood thawer | {} | Critical for cryopreserved HCT/P viability. [1, 2]
| IABP Performance Test Rig | Metrology | Counterpulsation Bench | {"pulse_freq":{"type":"range","units":"bpm"},"sync_acc":{"type":"number","units":"ms"},"pressure_acc":{"type":"number"}} | TRUE | TRUE | TRUE | 12 | 12 |  | High | 21 CFR 870.1025;IEC 60601-1 | IABP tester;balloon pump tester;CPB bench | {} | OQ verifies synchronization with ECG and arterial pressure triggers.
| WFI Loop Ozone Generator | Utilities | Electrolytic Sanitization | {"ozone_conc":{"type":"number","units":"ppb"},"destruction_cycle":{"type":"boolean"},"lamp_life":{"type":"integer"}} | TRUE | TRUE | TRUE | 6 | 6 | 12 | High | USP <1231>;ISO 13485 | ozone sanitization;water loop cleaner;sanitary ozone | {} | OQ verifies rapid ozone ramp-up and complete destruction before use.
| Active Battery Hermeticity Bench | Metrology | Helium Mass Spec Leak | {"lod_sensitivity":{"type":"number","units":"mbar.L/s"},"cycle_time":{"type":"number","units":"sec"},"sniff_mode":{"type":"boolean"}} | TRUE | TRUE | TRUE | 12 | 12 |  | High | ISO 13485;21 CFR 820 | battery leak;hermetic tester;mass spec leak | {} | 100% release test for lithium primary cells in cardiac implants. [3, 4]
| LFA Striper ZX1010 | IVD Production | Nano-Dispense System | {"dispense_vol":{"type":"range","units":"nL"},"line_width_acc":{"type":"number","units":"um"},"throughput":{"type":"number"}} | TRUE | TRUE | TRUE | 6 | 6 | 12 | High | ISO 13485;FDA QMSR | LFA striper;BioDot;ZX1010 | {} | PQ verifies antibody/antigen binding consistency on NC membranes.
| Ceramic Femoral Head Burst Tower | Metrology | Quasi-Static Compressive | {"axial_load_cap":{"type":"number","units":"kN"},"displacement_acc":{"type":"number","units":"um"},"fixture_v":{"type":"string"}} | TRUE | TRUE | TRUE | 12 | 12 |  | High | ISO 7206-10;ISO 11491 | burst tower;ceramic head burst;ortho crush | {} | Used for destructive safety qualification of ceramic joint components.
| Syringe Plunger Lubricity Bench | Metrology | Piston Friction Analyzer | {"breakloose_force":{"type":"number","units":"N"},"glide_force":{"type":"number","units":"N"},"speed":{"type":"number"}} | TRUE | TRUE | TRUE | 12 | 12 |  | Medium | ISO 7886-1;ISO 13485 | glide force;syringe lubricity;plunger tester | {} | Evaluates silicone oil coating performance on barrel walls.
| Nitinol Welding Glovebox | Production | Laser Shielded Chamber | {"o2_ppm_limit":{"type":"number","units":"ppm"},"h2o_ppm_limit":{"type":"number","units":"ppm"},"gas_type":{"type":"string"}} | TRUE | TRUE | TRUE | 12 | 12 | 12 | High | ISO 13485;ANSI Z136.1 | nitinol welder;laser box;shielding glovebox | {} | PQ verifies weld joint chemistry (oxidation levels) under inert gas.
| Suture Swage Pull-test Gauge | Metrology | Dynamic Force Analyzer | {"force_range":{"type":"range","units":"N"},"peak_hold":{"type":"boolean"},"grip_type":{"type":"string"}} | TRUE | TRUE | FALSE | 12 | 12 |  | High | USP <881>;ISO 13485 | swage tester;needle puller;suture gauge | {} | Reference standard to qualify needle swaging machines (USP pull test).
| High-Purity Gas Heat Exchanger | Utilities | WFI-to-N2 Thermal Unit | {"temp_setpoint":{"type":"number","units":"C"},"delta_p_max":{"type":"number","units":"psi"},"leak_detect":{"type":"boolean"}} | TRUE | TRUE | TRUE | 12 | 12 |  | Medium | ISO 8573-1;ISO 13485 | gas heater;N2 cooler;utility exchanger | {} | OQ verifies temperature stability of process gases in cleanrooms.
| Robotic Surgery Training Bench | Support | Surgical Effector Sim | {"DOF_acc":{"type":"number","units":"um"},"force_feedback":{"type":"boolean"},"latency":{"type":"number","units":"ms"}} | TRUE | TRUE | FALSE | 6 | 12 |  | High | IEC 60601-1;ISO 13485 | da vinci sim;robot bench;effector trainer | {} | Used to qualify software updates for robotic surgical tools.
| Micro-CT Volumetric Station | Metrology | 3D NDT Scanner | {"voxel_resolution":{"type":"number","units":"um"},"kvp_output":{"type":"number","units":"kV"},"sample_wt_max":{"type":"number"}} | TRUE | TRUE | FALSE | 12 | 12 |  | Medium | ISO 15708-3;ISO 17025 | Micro-CT;3D NDT;internal scanner | {} | Used to validate internal geometries of complex porous implants.
| Passive Settle Plate Stand | Facilities | ISO 5 Microbiology Stand | {"height":{"type":"number","units":"mm"},"shielding":{"type":"boolean"},"surface_finish":{"type":"string"}} | TRUE | FALSE | FALSE |  | 12 |  | Low | ISO 14698;FDA Guidance | settle stand;micro stand;exposure station | {} | IQ verifies placement logic for environmental monitoring points.
| Heart Valve Leaflet Sewing Rig | Assembly | Robotic Aseptic Sewing | {"stitch_acc":{"type":"number","units":"um"},"thread_tension":{"type":"number","units":"N"},"vision_v":{"type":"string"}} | TRUE | TRUE | TRUE | 6 | 6 | 12 | High | ISO 5840-1;ISO 13485 | valve sewer;leaflet sewer;robotic sewing | {} | PQ involves biochemical testing of tissue/scaffold bond integrity.
| Active Lead Tunneling Bench | Metrology | Insertion Force Sim | {"friction_acc":{"type":"number","units":"mN"},"travel_length":{"type":"number","units":"mm"},"bio-bath":{"type":"boolean"}} | TRUE | TRUE | TRUE | 12 | 12 |  | High | ISO 14708-2;ISO 13485 | tunneling rig;pacer insertion;friction bench | {} | PQ simulates lead passage through anatomical venous anatomy.
| IOL Badal Optometer | Metrology | Badal Optical Bench | {"diopter_acc":{"type":"number","units":"D"},"axis_res":{"type":"number","units":"deg"},"nist_traceable":{"type":"boolean"}} | TRUE | FALSE | FALSE | 12 |  |  | Medium | ISO 11979-2;ISO 17025 | optometer;badal bench;IOL master | {} | Reference standard to qualify intraocular lens power testers.
| Orthopedic Plate Fatigue Rig | Metrology | 3-Point Flexural Cycling | {"max_load":{"type":"number","units":"kN"},"cycle_freq":{"type":"number","units":"Hz"},"sampling_rate":{"type":"number"}} | TRUE | TRUE | TRUE | 6 | 6 | 24 | High | ASTM F382;ISO 13485 | plate fatigue;ortho fatigue;bone plate tester | {} | PQ verifies survival of trauma plates through 1 million cycles.
| Aseptic Pass-Through VHP Unit | Control | Bio-Decon Controller | {"dwell_time":{"type":"number","units":"min"},"h2o2_ppm_acc":{"type":"number","units":"%"},"audit_trail":{"type":"boolean"}} | TRUE | TRUE | TRUE | 6 | 6 | 12 | High | ISO 14937;ISO 14644-1 | hatch controller;VHP hatch;pass controller | {} | Software-driven control of aseptic material transfer logic. [5, 6]
| Cleanroom Static Decay Meter | Metrology | ESD Certification Probe | {"voltage_limit":{"type":"number","units":"V"},"decay_time_acc":{"type":"number","units":"ms"},"rh_compensation":{"type":"boolean"}} | TRUE | FALSE | FALSE | 12 |  |  | Low | IEC 61340;ISO 14644 | static meter;ESD probe;charge decay | {} | Used to qualify cleanroom flooring and workstations for active implants.