- PURPOSE

  - The purpose of this installation qualification is to define the qualification plan for the American BioTech Supply model ABT-HC-LS-49 Refrigerator. The refrigerator will be utilized for storing incoming supplies that are required to conduct operations activities at StimLabs. These incoming supplies must be stored in a refrigerated environment prior to utilization.

  - This validation is a deliverable for PCR21-045 “Addition of a Refrigerator to Warehouse 1.”

- SCOPE

  - This validation applies to the American BioTech Supply model ABT-HC-LS-49 Refrigerator that will be utilized for storing incoming supplies required to conduct operations activities at StimLabs. These incoming supplies must be stored in a refrigerated environment prior to utilization. This validation shall serve as the initial qualification of this equipment.

  - This equipment is located in Warehouse 1 at the StimLabs Roswell, GA facility.

- BACKGROUND

  - StimLabs is a tissue bank involved in the processing of human derived cells and tissues into products for human use. StimLabs receives supplies that require storage in a refrigerated environment prior to utilization during operations activities. These incoming supplies must be stored in a refrigerated environment to preserve their characteristics until they are utilized.

  - To maintain a refrigerated environment, a refrigerator must maintain its temperature between 1°C and 10°C. To ensure the refrigerator is fit for the intended purpose, StimLabs will perform an installation qualification and continue to monitor the performance of the refrigerator while it is being used to store supplies that require refrigeration through an operational qualification.

  - The American BioTech Supply model ABT-HC-LS-49 Refrigerator will be calibrated prior to operational requirement testing. The set point may be adjusted prior to operational requirement testing if necessary.

- REFERENCES

  - Food and Drug Administration 21 CFR Part 1271

  - American Association for Tissue Banks Standards for Tissue Banking

  - American BioTech Supply Refrigerator Owner’s Manual.

  - American BioTech Supply ABT-HC-LS-49 Spec sheet.

  - American BioTech Supply ABT-HC-LS-49 Exterior Cutsheet.

  - American BioTech Supply ABT-HC-LS-49 Interior Cutsheet.

  - FDA Guidance for Industry–Process Validation: General Principles and Practices

  - GHTF/SG3/N99-10:2004 Quality Management Systems–Process Validation Guidance

  - VAP21-029-00 ABT-HC-LS-49 Warehouse Refrigerator (PVP)

  - VAP21-029-01 ABT-HC-LS-49 Warehouse Refrigerator (IQ)

  - VAP21-029-02 ABT-HC-LS-49 Warehouse Refrigerator (OQ)

  - PCR21-045 “Addition of a Refrigerator to Warehouse 1.”

- DEFINITIONS

  - Acceptance Criteria: Predetermined requirements whereby pass/fail decisions are determined. Acceptance criteria are specific and measurable.

  - Critical to Quality: Attribute of a part, assembly, sub-assembly, product, or process that has a direct and significant impact on its actual or perceived quality to the customer or end user.

  - Deviation: The action of departing from an established method, procedure, or acceptance criteria.

  - Installation Qualification (IQ): Establishing by objective evidence that all key aspects of the process equipment and ancillary system installation adhere to the manufacturer’s approved specification(s).

  - Operational Qualification (OQ): Establishing by objective evidence the process control limits and action levels, which result in a product that meets or exceeds predetermined requirements, are within the operating limits required by the process.

  - Process Validation: A mechanism to ensure that a process, whose output is not fully verifiable by inspection and testing, is capable of consistently (repeatedly and reliably) producing a result or product that meets or exceeds predetermined requirements. Process validation proves that a process is under control and is used to determine the process variables, acceptable limits for those variables, and establish appropriate in-process controls.

  - Protocol: A pre-approved and prospective document that states the requirements and procedures for conducting a validation.

  - Validation/Study Leader: The individual who has responsibility and authority over the operation of the equipment or process being validated.

  - Verification: Confirmation by examination, and with provision of documented evidence, that a specific requirement has been fulfilled.

- RESPONSIBILITIES

| Function | Name | Author Protocol & Report | Review Protocol & Report | Execute Protocol | Update WI, SOP, Forms |
| --- | --- | --- | --- | --- | --- |
| Validation/Study Leader | Alexander Waters | X |  |  |  |
| MwER/MR/HOQ/TBD | Melissa O’Connor |  | X |  |  |
| Director of Quality | Stacy Harmon |  | X |  |  |
| Continuous Improvement Manager | Justin Junnier |  | X |  |  |
| Operations Management | Michael Baird |  | X |  | X |
| Leads/Operators/Engineers | Trained Personnel |  |  | X |  |

- SAFETY CONSIDERATIONS

  - Do not store volatile or flammable substances in this unit except in sealed containers.

  - Never use an acid-based cleaning solution

  - Before conducting repairs outside the scope of routine maintenance and inspection, the equipment should be powered off and disconnected from its utility connection(s).

- EQUIPMENT

  - Manufacturing Equipment Description

    - Manufacturer: American BioTech Supply (ABS)

    - Serial Number: ABS-20148447-2109

    - Model Number: ABT-HC-LS-49

    - External Dimensions:

      - Height: 81.75” with casters; 78.25” without casters.

      - Width: 54”

      - Depth: 34.75”

    - Location:  Warehouse 1

  - Manufacturing Equipment Parameters

    - The factory set point is 3℃. “SET” is the temperature set point (compressor off point) and may be adjusted. “Hy” is the temperature differential and may be adjusted. The target operation temperature range should have the low end of the interval equal to the “SET”, while the high end of the interval will equal “SET” + “Hy”. To maintain the operation temperature between 3°C and 7°C, you would set “SET” =3°C, and “Hy” =4°C.

    - Electronic Temperature Control

    - Serial plate voltage rating: 115V.

    - Serial plate frequency rating: 60 Hz.

  - Test Equipment

    - Calibrated equipment will be used for the execution of this protocol where necessary. Equipment ID’s, serial numbers and certificates of calibration will be documented in the data collection documentation where applicable.

    - Digital multimeter/voltage meter

    - Data logger(s)

    - Tape measure

- PROCEDURE

  - Verify Equipment Identification

    - Verify and document the manufacturer in the data collection sheets.

    - Verify and document the model number in the data collection sheets.

    - Verify and document the serial number in the data collection sheets.

  - Verify and Record the External Dimensions

    - Height: Measure and record the height of the equipment.

    - Width: Measure and record the width of the equipment.

    - Depth: Measure and record the depth of the equipment.

  - Verify Environmental Requirements

    - Temperature: Record the temperature in the area where the unit is installed.

    - Relative Humidity: Record the relative humidity in the area where the unit is installed.

    - Positioning: Verify that the unit is not exposed to direct sunlight or other sources of heat.

  - Verify the Utility Connection(s)

    - Voltage: Record the voltage of the equipment’s power supply.

    - Frequency: Record the frequency of the equipment’s power supply.

  - Verify Proper Installation

    - Ensure the equipment is appropriately leveled to facilitate proper articulation of the equipment doors.

    - Verify the casters are properly installed to keep the cabinet appropriately leveled and ensure the unit will not rock, lean, or display any symptoms of foundational instability.

    - Verify the bottle that houses the equipment’s thermocouple probe, in the equipment’s internal chamber, contains the appropriate thermal buffer medium.

    - Verify and document the shelves/trays are appropriately installed in the internal chamber of the equipment.

    - Verify and document that any light emission apparatuses, i.e., lamps, LEDs, bulbs, etc., are properly installed and illuminate appropriately.

    - Verify and document that adequate spacing, that meets or exceeds the minimum clearances expressed by the manufacturer, exists around the equipment. A clearance space of 2-3 inches around the refrigerator is specified by the manufacturer.

  - Verify Equipment Controls

    - With the equipment’s set point at 3℃, place the data logger probe in the centermost position of the refrigerator’s internal chamber. Allow the internal chamber and data logger probe to equilibrate for a minimum of one hour.

      - Verify and document a set point of 3℃ and an operating temperature range equal to the set point (3℃) -1℃ to +4℃ can be measured on the data logger positioned in the centermost position of the equipment’s internal chamber.

      - Verify and document the average temperature measured by the data logger agrees with the set point on the digital temperature controller display within -1℃ or +4℃.

    - Verify and document visual and audible alarms are triggered when the equipment’s internal chamber temperature exceeds 9℃; tripping the alarms can be achieved via opening the door of the refrigerator for approximately 5 minutes.

  - Verify Data Logger

    - Verify that the data logger is programmed.

    - Verify the data logger is placed in the center of the chamber.

    - After the unit is turned on, verify the data logger is recording the temperature. Monitor for at least one hour. Record the start time and stop time of the data logger and the average temperature measurement. The average temperature measured by the data logger should agree with the set point within a range of -1°C to +4℃.

    - Remove the data logger after the test, download and print the data, initial and date the printed report. Include VAP21-029-01 in the data logger report description and include a copy in the IQ report.

- LOT SIZE & SAMPLING PLAN

  - None

- ACCEPTANCE CRITERIA

  - Equipment Identification

    - Manufacturer: American BioTech Supply (ABS)

    - Serial Number: ABS-20148447-2109

    - Model Number: ABT-HC-LS-49

  - External Dimensions

    - Height: 81.75” with casters; 78.25” without casters.

    - Width: 54”

    - Depth: 34.75”

  - Environmental Requirements

    - Temperature must be between 5℃ and 40℃

    - Relative Humidity must be ≤ 80%

    - Verify unit is not installed in direct sunlight or other heat sources.

  - Utility Connection(s)

    - Voltage: 115V ±10%

    - Frequency: 60Hz ±10%

  - Installation

    - Verify that the equipment casters are properly installed and appropriately adjusted such that appropriate leveling is attained as evidenced by proper articulation of the equipment doors.

    - Verify that, upon achieving appropriate installation of the casters and subsequent appropriate leveling, the equipment has also achieved stability as evidenced by the equipment not rocking, leaning, or displaying any other type of foundational instability.

    - Verify that the bottle, which houses the equipment’s thermocouple probe, in the equipment’s internal chamber, contains the appropriate thermal buffer medium.

    - Verify that the door gasket is clean, intact, free of defects, and forms a seal with the equipment cabinet when in the closed position.

    - Verify that any light emission apparatuses, i.e., lamps, LEDs, bulbs, etc., are properly installed and illuminate appropriately.

    - Verify that adequate spacing, that meets or exceeds the minimum clearances expressed by the manufacturer, exists around the equipment. The manufacturer specifies that a clearance space of at least 2 to 3 inches is required around the refrigerator.

  - Equipment Controls

    - Set point: Verify and document a set point of 3℃ and an operating temperature range equal to the set point (3℃) -1℃ to +4℃ can be averaged on the data logger positioned in the centermost region of the equipment’s empty chamber over at least one hour.

      - The averaged data recorded by the data logger, the equipment set point, and the temperature displayed on the equipment’s digital temperature controller must all agree within a range of-1℃ to +4℃ to the set point of 3℃.

    - Alarm: Visible and audible alarms should initiate after the door has been intentionally left open to cause the internal chamber temperature to exceed 9℃.

  - Data Logger

    - Over the minimum timespan of one hour, the averaged data recorded by the data logger, the equipment set point, and the temperature displayed on the equipment’s digital temperature controller must all agree within a range of-1℃ to +4℃ of a set point of 3℃.

- MATERIAL DISPOSITION

  - None

- DOCUMENTATION REQUIREMENTS

  - Results shall be summarized in an IQ Report.

  - Various control plans, including the preventative maintenance (PM) schedule, safety/PPE requirements, environmental requirements, and re-validation criteria shall be documented in the installation qualification report.

  - Raw data sheets shall be included in the appendix of the IQ report and/or a reference to the location of the raw data shall be included in the IQ report.

  - Calibration records, where applicable, shall be included in the appendix of the IQ report.

  - Signed IQ Protocol Training forms will be included in the appendix of the IQ report.

- DEVIATIONS

  - Deviations to the IQ will be reviewed for approval using FRM020-012, Protocol Deviation Form, and included in the IQ report.

- PRE-REQUISITE STEPS

  - Approved Process Validation Plan (PVP) or justification for omission.

  - The IQ Protocol must be approved before validation activities may begin.

  - Completed Operator Training.

- APPENDIX

16.1–DATA COLLECTION SHEETS

STOP! Read Instructions.

Instructions:

- Verify the pre-requisite criteria have been met on this data collection sheet prior to testing. Note that the Pre-requisite criteria only need to be verified once for the validation.

- For attribute data: Circle “Pass” or “Fail” or write the words “Pass” or “Fail”

- For variable data: Record actual results in the units specified or add units if the units are not specified.

- Verification must include signature of the individual performing the step and the date performed. Verification must be completed concurrently with performance.

- Follow Good Documentation Practice (GDP).

Pre-requisite Criteria (must be completed and verified prior to start of IQ activities):

| Pre-requisite | Complete (Circle One) | Verification Signature/Date |
| --- | --- | --- |
| Process Validation Plan is approved. | Yes No |  |
| IQ Protocol is signed/approved. | Yes No |  |
| Operator training is complete and documented. | Yes No |  |

Pre-requisites verified by (signature & date): _____________________________________________________________

Quality reviewed by (signature & date): ______________________________________________________________

Section A–Equipment Identification

| Equipment Name | Manufacturer |  | Model Number | Serial Number |  | Stimlabs ID Number |  |  | Verification Signature/Date |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Specification or Acceptance Criteria |  | Result |  |  |  |  |  | Verification Signature/Date |  |  |
| External Dimensions |  | Height Actual: ____________ Manufacture: 81.75” w/ casters; 78.25” w/o casters |  |  | Pass / Fail |  |  |  |  |  |
|  |  | Width Actual: ____________ Manufacture: 54” |  |  | Pass / Fail |  |  |  |  |  |
|  |  | Depth Actual: ____________ Manufacture: 34.75” |  |  | Pass / Fail |  |  |  |  |  |
| Equipment is marked with StimLabs Equipment ID. |  | Pass / Fail |  |  |  |  |  |  |  |  |
| Equipment is marked with a calibration sticker and is within current calibration. |  | Pass / Fail |  |  |  |  |  |  |  |  |
| A logbook has been created for the equipment and is available for use. |  | Pass / Fail |  |  |  |  |  |  |  |  |

Section A completed by (signature & date): ________________________________________________________________

Quality reviewed by (signature & date): ______________________________________________________________

Section B-Installation

| Specification or Acceptance Criteria | Result |  | Verification Signature/Date |
| --- | --- | --- | --- |
| The equipment is appropriately leveled. | Pass / Fail |  |  |
| The equipment does not rock or lean and is stable. | Pass / Fail |  |  |
| The internal thermocouple probe bottle is filled with the appropriate thermal buffer medium. | Pass / Fail |  |  |
| The equipment door gasket is clean, intact, free of defects, and forms a proper seal to the cabinet. | Pass / Fail |  |  |
| Equipment shelves/trays are appropriately installed in the internal chamber. | Pass / Fail |  |  |
| Light emission apparatuses are properly installed and illuminate appropriately. | Pass / Fail |  |  |
| Clearance space is ≥ 2-3” around the refrigerator. | Left Side: | Pass / Fail |  |
|  | Right Side: | Pass / Fail |  |
|  | Back: | Pass / Fail |  |

Section B completed by (signature & date): ________________________________________________________________

Quality reviewed by (signature & date): ______________________________________________________________

Section C–Utility Connection

| StimLabs ID Number | Serial Number | Utility | Requirement | Result | Verification Signature/Date |
| --- | --- | --- | --- | --- | --- |
|  |  | Voltage | 115V±10% Actual: | Pass Fail |  |
|  |  | Frequency | 60 Hz ±10% Actual: | Pass Fail |  |

Record Test/Measurement Equipment Used and Calibration Due Date (attach calibration certificate to final report): __________________________________________________________________________________________________________________

__________________________________________________________________________________________________________________

Section C completed by (Signature & date): ________________________________________________________________

Quality reviewed by (signature & date): ______________________________________________________________

Section D-Environmental Requirements

| Equipment Description | Serial Number | Equipment ID Number or StimLabs ID Number | Result |  | Verification Signature/Date |
| --- | --- | --- | --- | --- | --- |
|  |  |  | Humidity | Temperature |  |
|  |  |  | Range: 80% Actual: Pass / Fail | Range: 5°C to 40°C Actual: Pass / Fail |  |
|  |  | Cal Due Date: |  |  |  |

Record Test/Measurement Equipment Used and Calibration Due Date (attach calibration certificate to final report): __________________________________________________________________________________________________________________

__________________________________________________________________________________________________________________

__________________________________________________________________________________________________________________

| Specification and Acceptance Criteria | Result | Verification Signature/Date |
| --- | --- | --- |
| The equipment is positioned away from direct sunlight or other sources of heat. | Pass / Fail |  |

Section D completed by (signature & date): ________________________________________________________________

Quality reviewed by (signature & date): ______________________________________________________________

Section E–Operational Requirements

| Specification and Acceptance Criteria | Reading/Results | Result | Verification Signature/Date |
| --- | --- | --- | --- |
| Set Point Verification – Unit turned on with the set point at 3°C. Data logger placed on shelf near the center of chamber. |  |  |  |
| Visual and audible alarm initiates after the doors have been intentionally left open. | Audible alarm temp: ____________________°C Visual alarm temp: ____________________°C | Pass / Fail Pass / Fail |  |
| 3℃ (-1°C /+4℃) must be averaged by the data logger over at least one hour in the empty chamber. | Enter Set-Point _____________℃ | N/A |  |
|  | Time data logger placed: _____________ | N/A |  |
|  | Time data logger checked (at least one hour after time place in): ______________ | N/A |  |
| The averaged data logger temperature measurement, the set-point, and the display temperature must agree within 3℃ (-1°C /+4℃) after a minimum of one (1) hour has elapsed. | Set-Point reading after 1 hour _____________℃ | Pass / Fail |  |
|  | Display reading after 1 hour _____________℃ | Pass / Fail |  |
|  | Results documented on data logger report | Pass / Fail |  |

Section E completed by (signature & date): ________________________________________________________________

Quality reviewed by (signature & date): ___________________________________________________________________