- PURPOSE

  - The purpose of this Installation Qualification (IQ) is to define the validation scheme for the So-Low Cleanroom Refrigerator. The refrigerator will be utilized for storing tissue and processing supplies that are required to be maintained in a refrigerated environment prior to processing.

  - This validation is a deliverable for PCR21-025 New Cleanroom Construction and Support Lab Expansion.

- SCOPE

  - This validation applies to the So-Low Cleanroom Refrigerator for storing tissue and processing supplies that are required to be maintained in a refrigerated environment prior to processing. This validation shall serve as the initial qualification of this equipment.

  - This equipment is located in the newly constructed Cleanroom at the StimLabs Roswell, GA facility.

- BACKGROUND

  - StimLabs is a tissue bank involved in the processing of human derived tissues into products for human use.  Stimlabs has constructed a new ISO Class 7 cleanroom for the processing of human tissue at its facility in Roswell, GA.  During the tissue processing, it is necessary to refrigerate tissue and solutions at various stages in order to preserve the characteristics of the tissue over time.

  - To maintain a refrigerated environment, a refrigerator must maintain its temperature between 1°C and 10°C. The set point drift for the temperature controller must not exceed 2°C in a given time span.  To ensure the refrigerator is fit for the intended purpose, StimLabs will perform an installation qualification and continue to monitor the performance of the refrigerator while it is being used to store human tissue and supplies through an operational qualification.

  - The So-Low Cleanroom Refrigerator has been calibrated prior to the start of the validation activities.  The set point may be adjusted prior to the start of the validation if desired.  Loading the unit will cause the temperature to rise as the warmer product is introduced into the chamber.  If a large amount of product is to be introduced, the manufacturer recommends loading in stages, allowing several hours between stages to allow temperature to stabilize again before the introduction of additional warm product. The manufacture also recommends:

    - Minimum clearances must be maintained.

    - Avoid blocking the evaporator fans on the top shelf. The should be four inches of clearance around the fan to allow proper air flow.

    - Do not overload the unit. The maximum suggested load is about 75% of the chamber volume.

    - Minimize the time the door is left open.

- REFERENCES

  - Food and Drug Administration 21 CFR Part 1271

  - American Association for Tissue Banks Standards for Tissue Banking

  - So-Low Cleanroom Refrigerator Equipment Manual

  - FDA Guidance for Industry – Process Validation: General Principles and Practices

  - GHTF/SG3/N99-10:2004 Quality Management Systems – Process Validation Guidance

  - VAP21-010-00 So-Low Cleanroom Refrigerator (PVP)

- DEFINITIONS

  - Acceptance Criteria:  Pre-determined requirements upon which pass/fail decisions are based.  Acceptance criteria are specific and measureable.

  - Critical to Quality:  Attribute of a part, assembly, sub-assembly, product or process that has direct and significant impact on its actual or perceived quality to the customer or end user.

  - Deviation:  The action of departing from an established method, procedure, or acceptance criteria.

  - Installation Qualification (IQ):  Establishing by objective evidence that all key aspects of the process equipment and ancillary system installation adhere to the manufacturer’s approved specification(s).

  - Process Validation:  A mechanism to ensure that a process whose output is not fully verifiable by inspection and test is capable of consistently (repeatedly and reliably) producing a result or product that meets pre-determined requirements.  Process validation proves that a process is under control and is used to determine the process variables, acceptable limits for those variables, and to establish appropriate in-process controls.

  - Protocol:  A pre-approved and prospective document that states the requirements and procedures for conducting a validation.

  - Validation/Study Leader: The individual who has the responsibility and authority over the operation of the equipment or process being validated.

  - Verification:  Confirmation by examination with a provision of documented evidence that a specific requirement has been fulfilled.

- RESPONSIBILITIES

| Function | Name | Author Protocol | Review & Approve Protocol | Execute Protocol | Execute Testing |
| --- | --- | --- | --- | --- | --- |
| Validation/Study Leader | Michael Baird | X | X |  |  |
| MwER/MR/HOQ/TBD | Melissa O’Connor |  | X |  |  |
| Operations Management | Michael Baird |  | X |  |  |
| Quality Management | Stacy Harmon |  | X |  |  |
| Other | Rob Frye | X | X |  |  |
| Other | Trained Technician |  |  | X | X |

- SAFETY CONSIDERATIONS

  - Use correct voltage

  - Do not expose to water or moisture; use the controller only when the operating limits avoid sudden temperature changes with high atmospheric humidity to prevent formation of condensation.

  - Warning: Disconnect all electrical connections before any kind of maintenance.

  - Fit the probe where it is not accessible by the end user.

  - Consider the maximum current which can be applied to each relay.

  - Ensure that the wires for probes, loads, and the power supply are separated and far enough from each other, without crossing or intertwining.

  - Improperly installed lamps may cause damage to the lighting circuit.

- EQUIPMENT

  - Manufacturing Equipment Description

    - Manufacturer: So-Low

    - Serial Number(s): will be identified on the data collection sheets for the IQ

    - Model Number: DH4-24GD-TM-CR

    - Location:  New Cleanroom

  - Manufacturing Equipment Parameters

    - Set point is 5℃ ± 2℃ (over 24 hours)

    - Electronic Temperature Control

    - 6 Amp Power Source

  - Test Equipment

    - Calibrated equipment shall be used for the execution of this protocol.  Equipment ID/serial numbers and calibration information will be documented on the data collection sheet and final report

    - Voltage Meter

    - Tape Measure

    - Data logger (temperature)

- PROCEDURE

  - Verify Equipment Identification

    - Confirm the model identified in this protocol.

    - Document serial number on the data collection sheet.

  - Verify Environmental Requirement

    - Temperature – Measure the temperature in the area where the unit is installed

    - Relative Humidity – Measure the relative humidity in the area where the unit is installed.

  - Verify Critical Dimensions

    - Verify the dimensional specification of the unit.

  - Verify Proper Installation

    - Leveling – make sure the unit is leveled properly so the doors operate correctly. Verify the unit has a slight tilt from front to back.

    - Door Hinge – Verify that the door has been installed according to manufacturer’s direction in the equipment manual and that the door can open and close properly. If the door does not operate properly, adjust the level of the unit.

    - Shelving – Verify that the shelves are securely attached according to the manufacturer’s recommendation in the equipment manual.

    - Lamps – verify the lamps are properly installed, illuminated, and fully engaged in the lamp holders.

    - Verify wall space per manufacturer’s recommendations – verify that there is at least 3” of space between the cabinet and any adjacent wall or fixture. Verify there is at least 12” of space above the unit.

  - Verify Utility Connections

    - Measure amps being drawn by unit.

  - Verify Equipment Controls

    - Temperature – Allow the unit to run for at least 4 hours. Verify the set-point of 5℃ ± 2℃ can be measured after 4 hours. Place one data logger with the probe positioned in the middle of the chamber to check the temperature of the chamber. Verify a set-point 5℃ ± 2℃ can be measured on the data logger. Verify the temperature measured by the data logger is in agreement with the set point and display temperature on the temperature controller. (within 2℃).

    - Verify alarm with the door open – After reaching set point, open the door of the refrigerator and wait at least 5 minutes. Verify the refrigerator displays visual and audible alarms when the temperature is not controlled.

  - Verify Chart Recorder

    - Verify that the chart recorder is installed.

    - Verify Chart Recorder paper is installed.

    - Verify the pen is making contact with the paper.

    - After unit is turned on, verify the Chart Recorder pen is recording the temperature on the paper. Monitor for at least one hour by recording the start time and temperature measurements the recorder recorded. The temperature measured by the Chart Recorder should be in the agreement with the data logger within 2℃.

    - Remove the Chart Recorder paper after the test, initial, and date the paper and write “VAP21-010-01” on the paper and include a copy in the IQ final report.

- LOT SIZE & SAMPLING PLAN

  - None

- ACCEPTANCE CRITERIA

  - Equipment Verification

    - Must match Manufacturing Equipment Description section above.

  - Environment Verification

    - Temperature is for reference only

    - Relative Humidity is for reference only

  - Critical Dimensions

    - The dimensions of the chamber should match the equipment specification within the tolerance described in the manufacturer’s specification. Exterior unit dimensions are 28.35” x 33.5” x 79.25”

  - Installation

    - Leveling –The unit has a slight tilt from front to back upon visual examination.

    - Door Hinge – the door should open and close easily with no resistance. The door should remain close.

    - Shelving –The shelves are securely attached.

    - Lamps – Lamp should be securely snapped into place.

    - Wall Space- there is at least 3” of space between the cabinet and any adjacent wall or fixture. Verify there is at least 12” of space above the unit.

  - Utility Connections

    - Amperage

      - Verify amperage is 6.0 Amps (±5%)

  - Equipment Controls

    - Set-point

      - 5℃ ± 2℃ must be measured by the data logger (after 4 hour ramp) in the empty chamber.

      - The data logger (average over at least one hour), set-point, and display temperature must be in agreement within 2℃.

    - Alarm

      - Visible and audible alarm should sound after the door has been left open for a period of five minutes.

  - Chart Recorder

    - Temperature recorded should match the data logger temperature accurately over a time span of at least upon hour ± 2℃.

- MATERIAL DISPOSITION

  - None

- DOCUMENTATION REQUIREMENTS

  - Results shall be summarized in an IQ Report.

  - Various Control Plans, including the Preventative Maintenance (PM) Schedule, safety/PPE requirements, environmental requirements, and re-validation criteria shall be documented in the final report.

  - Raw data sheets shall be included in the appendix of the IQ report and/or a reference to the location of the raw data shall be included in the IQ report.

  - Calibration records (if applicable) shall be attached to the IQ report (in the Appendix).

  - Signed IQ Protocol Training forms will be attached to the IQ report (in the Appendix).

- DEVIATIONS

  - Deviations to the IQ will be pre-approved using FRM020-012 Protocol Deviation Form and explained in the IQ report.

  - Copies of signed Protocol Deviation Forms will be attached to the IQ report (in the Appendix).

- PRE-REQUISITE STEPS

  - Approved Process Validation Plan (PVP) or justification for omission.

  - The IQ Protocol must be signed before validation activities may begin.

  - Completed Operator Training.

- APPENDIX

  - Blank data collection sheets

APPENDIX 16.1 – DATA COLLECTION SHEETS

STOP!  Read Instructions.

Instructions:

- Complete one data collection sheet for each So-Low Cleanroom Refrigerator.

- Verify the Pre-requisite criteria have been met on this data collection sheet prior to testing.  Note that the Pre-requisite criteria only need to be verified once for the validation.

- For attribute data: Circle “Pass” or “Fail” or write the words “Pass” or “Fail”

- For variable data: Record actual results in the units specified or add units if the units are not specified.

- Verification must include signature of the individual performing the step and the date performed.  Verification must be completed concurrently with performance.

- Follow Good Documentation Practice (GDP).

Pre-requisite Criteria (must be complete and verified prior to start of IQ activities):

| Pre-requisite | Complete (Circle One) | Verification Signature/Date |
| --- | --- | --- |
| Process Validation Plan is approved. | Yes No |  |
| IQ Protocol is signed and approved. | Yes No |  |
| Operator Training is complete and documented. | Yes No |  |

Pre-Requisites Verified By (signature & date): _____________________________________________________________

Quality Reviewed By (signature & date): ______________________________________________________________

Section A – Equipment Identification

| Equipment Name | Model Number | Manufacturer | Serial Number | Stimlabs ID Number | Verification Signature/Date |
| --- | --- | --- | --- | --- | --- |

| Specification or Acceptance Criteria | Result |  | Verification Signature/Date |
| --- | --- | --- | --- |
| External Dimensional Specifications Match Protocol | Length: Actual: ____________ Manufacture: 28.35” | Pass / Fail |  |
|  | Height: Actual: ____________ Manufacture: 79.25” | Pass / Fail |  |
|  | Width: Actual: ____________ Manufacture: 33.5” | Pass / Fail |  |
| Equipment is marked with StimLabs Equipment ID | Pass / Fail |  |  |
| Equipment is marked with a calibration sticker and is within current calibration | Pass / Fail |  |  |
| A logbook has been created for the equipment and is available for use | Pass / Fail |  |  |

Section A Completed By (signature & date): ________________________________________________________________

Quality Reviewed By (signature & date): ______________________________________________________________

Section B - Installation

| Specification or Acceptance Criteria | Result |  | Verification Signature/Date |
| --- | --- | --- | --- |
| Equipment has a slight tilt from front to back. | Pass / Fail |  |  |
| Door open and closed properly | Pass / Fail |  |  |
| Shelves are securely attached | Pass / Fail |  |  |
| Lamps installed , illuminated, and properly locked in place. | Pass / Fail |  |  |
| Wall space has at a minimum 3” between unit and adjacent walls | Left Side: | Pass / Fail |  |
|  | Right Side: | Pass / Fail |  |
|  | Back: | Pass / Fail |  |

Section B Completed By (signature & date): ________________________________________________________________

Quality Reviewed By (signature & date): ______________________________________________________________

Section C – Utility Connection

| StimLabs ID Number | Serial Number | Requirement | Result | Verification Signature/Date |
| --- | --- | --- | --- | --- |
|  |  | 6 Amps Actual: | Pass Fail |  |

Record Test/Measurement Equipment Used and Calibration Due Date (attach calibration certificate to final report): __________________________________________________________________________________________________________________

__________________________________________________________________________________________________________________

Section C Completed By (Signature & date): ________________________________________________________________

Quality Reviewed By (signature & date): ______________________________________________________________

Section D - Environmental Requirements

| StimLabs ID Number | Serial Number | Chart Recorder ID Number | Result |  | Verification Signature/Date |
| --- | --- | --- | --- | --- | --- |
|  |  |  | Humidity | Temperature |  |
|  |  |  | Range: Reference only Actual: Pass / Fail | Range: Reference only Actual: Pass / Fail |  |
|  |  | Cal Due Date: |  |  |  |

Record Test/Measurement Equipment Used and Calibration Due Date (attach calibration certificate to final report): __________________________________________________________________________________________________________________

__________________________________________________________________________________________________________________

__________________________________________________________________________________________________________________

Section D Completed By (signature & date): ________________________________________________________________

Quality Reviewed By (signature & date): ______________________________________________________________

Section E – Operational Requirements

| Specification and Acceptance Criteria | Reading/Results | Result | Verification Signature/Date |
| --- | --- | --- | --- |
| Set Point Verification – Unit turned on and set to 5°C. Data logger placed on shelf near the center of chamber. |  |  |  |
| Open the door and wait at least 5 minutes. Enter time audible and visual alarm initiates. Greater than 5 minutes is acceptable | Audible alarm time: Start:____________ Stop:____________ Duration:_______ Visual alarm time: Start:____________ Stop:____________ Duration:_______ | Pass / Fail Pass / Fail |  |
| 5℃ ± 2℃ must be measured by the data logger (average over at least one hour) in the empty chamber. | Enter Set-Point _____________ ℃ | N/A |  |
|  | Enter Time Empty unit was turned on: __________________ | N/A |  |
|  | Enter Time set point was checked (4 hours after unit turned on): ____________________ | N/A |  |
|  | Display reading after 4 hours _____________ ℃ | 5℃ ± 2℃ |  |
|  | Time data logger placed ( 4 hours after unit turned on): _____________ | N/A |  |
|  | Time data logger checked (at least one hour after time place in): ______________ | N/A |  |
|  | Set-Point reading after 1 hour _____________ ℃ | Pass / Fail |  |
| The data logger (average over at least one hour), set-point, and display temperature must be in agreement within 2℃. | Display reading after 1 hour _____________ ℃ | Pass / Fail |  |
| Temperature recorded should match the data logger temperature accurately over a time span of at least upon hour ± 2℃. | Results documented on data logger report | Pass / Fail |  |

Section E Completed By (signature & date): ________________________________________________________________

Quality Reviewed By (signature & date): ___________________________________________________________________