- PURPOSE

  - The purpose of this installation qualification update the temperature settings that will be tested for the newly installed incubator at StimLabs’ facility used to store tissue processing solutions.

  - This validation is a deliverable for PCR21-025 New Cleanroom Construction and Support Lab Expansion.

- SCOPE

  - This validation applies to the Incucell incubator.  During tissue processing, Stimlabs utilizes heated solution in order to clean placental tissue.  Incucell incubator will be used to store the processing solutions so they will maintain the desired temperature prior to use.

  - This equipment is located in the new cleanroom at the StimLabs Roswell, GA facility

- BACKGROUND

  - StimLabs is a tissue bank involved in the processing of human derived tissues into products for transplantation.  During tissue processing, StimLabs utilizes processing solutions that have been warmed in order to increase the efficiency of the cleaning rinses.  StimLabs has purchased the incubator unit for the storage of these processing solutions at controlled temperatures.

  - Successful completion of this validation study will ensure continued accuracy and reliability of the functions provided by the equipment.

- REFERENCES

  - Food and Drug Administration 21 CFR Part 1271

  - American Association for Tissue Banks Standards for Tissue Banking

  - Incucell Incubator Equipment Manual

  - FDA Guidance for Industry – Process Validation: General Principles and Practices

  - GHTF/SG3/N99-10:2004 Quality Management Systems – Process Validation Guidance

  - VAP21-013-00 Incucell Incubator (PVP)

- DEFINITIONS

  - Acceptance Criteria:  Pre-determined requirements upon which pass/fail decisions are based.  Acceptance criteria are specific and measurable.

  - Critical to Quality:  Attribute of a part, assembly, sub-assembly, product or process that has direct and significant impact on its actual or perceived quality to the customer or end user.

  - Deviation:  The action of departing from an established method, procedure, or acceptance criteria.

  - Installation Qualification (IQ):  Establishing by objective evidence that all key aspects of the process equipment and ancillary system installation adhere to the manufacturer’s approved specification(s).

  - Process Validation:  A mechanism to ensure that a process whose output is not fully verifiable by inspection and test is capable of consistently (repeatedly and reliably) producing a result or product that meets pre-determined requirements.  Process validation proves that a process is under control and is used to determine the process variables, acceptable limits for those variables, and to establish appropriate in-process controls.

  - Protocol:  A pre-approved and prospective document that states the requirements and procedures for conducting a validation.

  - Validation/Study Leader: The individual who has the responsibility and authority over the operation of the equipment or process being validated.

  - Verification:  Confirmation by examination with a provision of documented evidence that a specific requirement has been fulfilled.

- RESPONSIBILITIES

| Function | Name | Author Protocol | Review & Approve Protocol | Execute Protocol | Execute Testing |
| --- | --- | --- | --- | --- | --- |
| Validation/Study Leader | Michael Baird | X | X |  |  |
| MwER/MR/HOQ/TBD | Melissa O’Connor |  | X |  |  |
| Operations Management | Michael Baird |  | X |  |  |
| Quality Management | Stacy Harmon |  | X |  |  |
| Other | Rob Frye |  | X |  |  |
| Other | Trained Technician |  |  | X | X |

- SAFETY CONSIDERATIONS

  - Use the correct voltage

  - Ensure that the unit is installed on a level, sturdy floor

  - Do not install unit in a flammable or volatile location

  - Do not install the unit in a humid place or in a place where it is likely to be splashed by water

  - Always disconnect the power supply to the unit prior to any repair or maintenance

- EQUIPMENT

  - Manufacturing Equipment Description

    - Manufacturer: BMT Medical Technology

    - Serial Number(s): will be identified on the data collection sets for the IQ

    - Model Number: Incucell Eco 404

  - Manufacturing Equipment Parameters

    - Target Temperature for Operations: 39°C ± 3.9°C

    - Electronic Temperature Control

    - Power Source: 115 Volts +10, -5

  - Test Equipment

    - Calibrated equipment shall be used for the execution of this protocol.  Equipment ID/serial numbers and calibration information will be documented on the data collection sheet.

    - Voltage Meter

    - Data Logger

- PROCEDURE

  - Verify Equipment Identification

    - Confirm the model identified in this protocol.

    - Document serial number on the data collection sheet.

  - Verify Environmental Requirement

    - Temperature – Measure the temperature inthe area where the unit is installed

    - Relative Humidity – Measure the relative humidity in the area where the unit is installed.

  - Verify Critical Dimensions

    - Verify the external dimensional specification of the unit.

  - Verify Proper Installation

    - Stabilization - The unit should not rock, lean, or be otherwise unstable.

    - Shelving – verify that shelves are securely attached.

    - Verify wall space per manufacturer’s recommendations – verify that there is at least 4” of space between the cabinet and any adjacent wall or fixture.

  - Verify Utility Connections

    - Measure the voltage of the refrigerator power source.

    - Measure max amp breaker

  - Verify Equipment Controls

    - Power Key

      - Activates the digital display when pressed.

      - Activates stand by mode when held for 1 second.

    - Temperature 39°C – With the set-point set to 39°C, place data loggers with the probe positioned in the top, middle, and bottom of the chamber to check the temperature of the chamber. Allow the temperature to equilibrate for at least one hour after placing the probe. Data loggers will begin recording data after the one-hour time period. Verify a set-point 39°C ± 3.9°C can be measured on the data logger for at least 4 hours. Verify the temperature measured by the data logger is in agreement with the set point. 39°C ± 3.9°C is the standard operating temperature in which this incubator is going to be utilized.

- LOT SIZE & SAMPLING PLAN

  - None

- ACCEPTANCE CRITERIA

  - Equipment Verification

    - Must match Manufacturing Equipment Description section above.

  - Environment Verification

    - Temperature must be between 5°C to 40°C

    - Relative Humidity less 80%

  - Critical Dimensions

    - The dimensions of the chamber should match the equipment specification within the tolerance described in the manufacturer’s specification. Dimensions include door, handle, legs or casters.

      - Length – 29.9”

      - Depth – 31.1”

      - Height – 75.2”

  - Installation

    - Stabilization

      - The unit should not rock, lean, or be otherwise unstable.

    - Wall Space

      - 4” of spacing between the unit and adjacent walls.

  - Utility Connections

    - Power

      - Verify power is 115V +10,-5V

    - Max Amperage

      - Verify minimum of a 15 Amp Breaker.

  - Equipment Controls

    - Power Key

      - Activates the digital display when pressed.

      - Activates stand by mode when held for 1 second.

    - Set-point

      - Set point of 39°C ± 3.9°C must be obtained in the empty chamber. The data logger (average over at least four hours), set-point must be in agreement within ± 10% of set point.

- MATERIAL DISPOSITION

  - None

- DOCUMENTATION REQUIREMENTS

  - Results shall be summarized in an IQ Report.

  - Various Control Plans, including the Preventative Maintenance (PM) Schedule, safety/PPE requirements, environmental requirements, and re-validation criteria shall be documented in the final report.

  - Raw data sheets shall be included in the appendix of the IQ report and/or a reference to the location of the raw data shall be included in the IQ report.

  - Calibration records (if applicable) shall be attached to the IQ report (in the Appendix).

  - Signed IQ Protocol Training forms will be attached to the IQ report (in the Appendix).

- DEVIATIONS

  - Deviations to the IQ will be pre-approved using FRM020-012 Protocol Deviation Form and explained in the IQ report.

  - Copies of signed Protocol Deviation Forms will be attached to the IQ report (in the Appendix).

- PRE-REQUISITE STEPS

  - Approved Process Validation Plan (PVP) or justification for omission.

  - The IQ Protocol must be signed before validation activities may begin.

  - Completed Operator Training.

- APPENDIX

  - Blank data collection sheets

APPENDIX 16.1 – DATA COLLECTION SHEETS

STOP!  Read Instructions.

Instructions:

- Complete one data collection sheet for each Incucell Incubator.

- Verify the Pre-requisite criteria have been met on this data collection sheet prior to testing.  Note that the Pre-requisite criteria only need to be verified once for the validation.

- For attribute data: Circle “Pass” or “Fail” or write the words “Pass” or “Fail”

- For variable data: Record actual results in the units specified or add units if the units are not specified.

- Verification must include signature of the individual performing the step and the date performed.  Verification must be completed concurrently with performance.

- Follow Good Documentation Practice (GDP).

Pre-requisite Criteria (must be complete and verified prior to start of IQ activities):

| Pre-requisite | Complete (Circle One) | Verification Signature/Date |
| --- | --- | --- |
| Process Validation Plan is approved. | Yes No |  |
| IQ Protocol is signed and approved. | Yes No |  |
| Operator Training is complete and documented. | Yes No |  |

Pre-Requisites Verified By (signature & date): _____________________________________________________________

Quality Reviewed By (signature & date): ______________________________________________________________

Section A – Equipment Identification

| Equipment Name | Model Number | Manufacturer | Serial Number | Stimlabs ID Number | Verification Signature/Date |
| --- | --- | --- | --- | --- | --- |

| Specification or Acceptance Criteria | Result |  | Verification Signature/Date |
| --- | --- | --- | --- |
| External Dimensional Specifications Match Protocol | Depth: Actual: ____________ Manufacture: 31.1” | Pass / Fail |  |
|  | Height: Actual: ____________ Manufacture: 75.2” | Pass / Fail |  |
|  | Length: Actual: ____________ Manufacture: 29.9” | Pass / Fail |  |
| Equipment is marked with StimLabs Equipment ID | Pass / Fail |  |  |
| Equipment is marked with a calibration sticker and is within current calibration | Pass / Fail |  |  |
| A logbook has been created for the equipment and is available for use | Pass / Fail |  |  |

Section A Completed By (signature & date): ________________________________________________________________

Quality Reviewed By (signature & date): ______________________________________________________________

Section B - Installation

| Specification or Acceptance Criteria | Result |  | Verification Signature/Date |
| --- | --- | --- | --- |
| Equipment does not rock, lean, or otherwise unstable. | Pass / Fail |  |  |
| Wall space has at a minimum 4” between unit and adjacent walls | Left Side: | Pass / Fail |  |
|  | Right Side: | Pass / Fail |  |
|  | Back: | Pass / Fail |  |

Section B Completed By (signature & date): ________________________________________________________________

Quality Reviewed By (signature & date): ______________________________________________________________

Section C – Utility Connection

| StimLabs ID Number | Serial Number | Utility | Requirement | Result | Verification Signature/Date |
| --- | --- | --- | --- | --- | --- |
|  |  |  | 115 +10,-5 Volt Actual: | Pass Fail |  |
|  |  |  | Minimum 15A breaker Actual: | Pass Fail |  |

Record Test/Measurement Equipment Used and Calibration Due Date (attach calibration certificate to final report): __________________________________________________________________________________________________________________

__________________________________________________________________________________________________________________

Section C Completed By (Signature & date): ________________________________________________________________

Quality Reviewed By (signature & date): ______________________________________________________________

Section D - Environmental Requirements

| StimLabs ID Number | Serial Number | Chart Recorder ID Number | Result |  | Verification Signature/Date |
| --- | --- | --- | --- | --- | --- |
|  |  |  | Humidity | Temperature |  |
|  |  |  | Range: 80% Actual: Pass / Fail | Range: 5°C to 40°C Actual: Pass / Fail |  |
|  |  | Cal Due Date: |  |  |  |

Record Test/Measurement Equipment Used and Calibration Due Date (attach calibration certificate to final report): __________________________________________________________________________________________________________________

__________________________________________________________________________________________________________________

__________________________________________________________________________________________________________________

Section D Completed By (signature & date): ________________________________________________________________

Quality Reviewed By (signature & date): ______________________________________________________________

Section E – Operational Requirements

| Specification and Acceptance Criteria | Reading/Results | Result | Verification Signature/Date |
| --- | --- | --- | --- |
| On/Off Key | On – Digital display lights up | Pass / Fail |  |
|  | Off – Hold key for 1 second – the unit is in stand by mode | Pass / Fail |  |
| Set Point Verification 39°C – Unit turned on and set to temperature below. |  |  |  |
| 39°C ± 3.9°C must be measured by the data logger (average over at least four hour) in the empty chamber. After 1 hour of equilibrate period. | Enter Set-Point _____________ °C | N/A |  |
|  | Data Logger ID placed: Top_____________ Middle_____________ Bottom_____________ | N/A |  |
|  | Time data placed in chamber ______________ | N/A |  |
|  | Time data logger removed (at least 5 hours after time place in): ______________ | N/A |  |
|  | Set-Point reading after 5 hours _____________ °C | Pass / Fail |  |
| Temperature recorded should match the data logger temperature accurately over a time span of at least one hour ± 10%. | Results documented on data logger report | Pass / Fail |  |

Section E Completed By (signature & date): ________________________________________________________________

Quality Reviewed By (signature & date): ___________________________________________________________________