- PURPOSE

  - The purpose of this validation is to provide individuals who’s responsibility is to approve the installation qualification protocol that did not approve REV 01 for the newly installed heat sealer units at StimLabs’ facility used to seal inner and outer pouches for StimLabs tissue products.

  - This validation is a deliverable for PCR21-025 New Cleanroom Construction and Support Lab Expansion.

- SCOPE

  - This validation applies to the Van der Stahl heat sealer.  A heat sealer is used to seal inner and outer product pouches.  This validation includes the qualification of the heat sealer equipment and heat sealing process.  Process setting for both inner and outer pouch shall remain the same validated under VAP/VAR16-005. The inner pouch will maintain a sterile barrier.  Sterility maintenance was validated under VAP/VAR16-002.

  - This equipment is located in the new cleanroom at the StimLabs Roswell, GA facility.

  - Inner pouch, vial pouch, and outer pouch are constructed of different material, thus necessitating a different set of parameters to achieve an acceptable seal.  The sealers will be validated to define the process parameters for sealing both inner and outer pouches using the validated process setting.

- BACKGROUND

  - StimLabs is a tissue bank involved in the processing of human derived tissues into products for transplantation.  In order to provide the tissue products in a manner that complies with sterile transplantation procedures, StimLabs packages the tissue product in a dual pouch configuration, including both an inner and outer pouch.

  - To ensure the pouch sealer is fit for the intended purpose, StimLabs will perform an installation qualification and continue to monitor the performance of the pouch sealer while it is being used to seal product pouches.

- REFERENCES

  - Food and Drug Administration 21 CFR Part 1271

  - American Association for Tissue Banks Standards for Tissue Banking

  - Van Der Stahl Heat Sealer Equipment Manual

  - FDA Guidance for Industry – Process Validation: General Principles and Practices

  - GHTF/SG3/N99-10:2004 Quality Management Systems – Process Validation Guidance

  - VAP21-014-00 Van Der Stahl Heat Sealer PVP

- DEFINITIONS

  - Acceptance Criteria:  Pre-determined requirements upon which pass/fail decisions are based.  Acceptance criteria are specific and measureable.

  - Critical to Quality:  Attribute of a part, assembly, sub-assembly, product or process that has direct and significant impact on its actual or perceived quality to the customer or end user.

  - Deviation:  The action of departing from an established method, procedure, or acceptance criteria.

  - Installation Qualification (IQ):  Establishing by objective evidence that all key aspects of the process equipment and ancillary system installation adhere to the manufacturer’s approved specification(s).

  - Process Validation:  A mechanism to ensure that a process whose output is not fully verifiable by inspection and test is capable of consistently (repeatedly and reliably) producing a result or product that meets pre-determined requirements.  Process validation proves that a process is under control and is used to determine the process variables, acceptable limits for those variables, and to establish appropriate in-process controls.

  - Protocol:  A pre-approved and prospective document that states the requirements and procedures for conducting a validation.

  - Validation/Study Leader: The individual who has the responsibility and authority over the operation of the equipment or process being validated.

  - Verification:  Confirmation by examination with a provision of documented evidence that a specific requirement has been fulfilled.

- RESPONSIBILITIES

| Function | Name | Author Protocol | Review & Approve Protocol | Execute Protocol | Execute Testing |
| --- | --- | --- | --- | --- | --- |
| Validation/Study Leader | Michael Baird | X | X |  |  |
| MwER/MR/HOQ/TBD | Melissa O’Connor |  | X |  |  |
| Operations Management | Michael Baird |  | X |  |  |
| Quality Management | Stacy Harmon |  | X |  |  |
| Other | Rob Frye |  | X |  |  |
| Other | Trained Technician |  |  | X | x |

- SAFETY CONSIDERATIONS

  - Use of the correct voltage.

- EQUIPMENT

  - Manufacturing Equipment Description

    - Manufacturer: Van der Stahl

    - Serial number: will be identified in the IQ data sheets

    - Model Number: MS350

    - Location: New Cleanroom

  - Manufacturing Equipment Parameters

    - Heating Temperature: 140°F to 400°F (in 1°F increments)

    - Heating Time: 0.0 to 5.0 seconds (in 0.1 second increments)

    - Cooling temperature: 100°F to the set heating temperature (in 1°F increments)

    - Temperature Precision: ±10°F

    - Electric Temperature Control

    - Power Source AC 110V/220V

  - Test Equipment

    - Calibrated equipment shall be used for the execution of this protocol.  Equipment ID/serial numbers and calibration information will be documented on the data collection sheet.

    - Voltage Meter

    - Tape Measure

    - Data logger (temperature)

- MATERIALS

  - N/A

- TEST METHODS

  - N/A

- PROCEDURE

  - Verify Equipment Identification

    - Confirm the model identified in this protocol.

    - Document serial number on the data collection sheet.

  - Verify Utility Connections

    - Measure the voltage of the refrigerator power source.

    - Minimum breaker measurement

  - Verify Proper Installation

    - Machine installed on a flat, secure working place.

  - Calibration Verification

    - Calibration completed prior to start of validation

  - Verify Safety Requirements

    - Verify micro-switch does not turn on and lever automatically reverses to original position when foreign object is pinched

    - Cycle ended with activation of ABORT button

  - Equipment Operation

    - On/Off switch works properly

    - Start/Stop

      - Press each start button with the machine in the ready position and confirm the machine starts.

      - Press stop button after the machine has started and confirm that the operation is aborted.

    - Temperature Controller

      - Set temperature at 400°F and start machine to confirm operation.

      - Set temperature to 400°F and start machine to confirm operation.

    - Timer

      - Set time to 0.8 seconds, start machine, and confirm the machine operates.

      - Set time to 5 seconds, start machine, and confirm that machine operates.

- LOT SIZE & SAMPLING PLAN

  - N/A

- ACCEPTANCE CRITERIA

  - Equipment Verification

    - Must match Manufacturing Equipment Description section above.

  - Utility Connections

    - Power

      - 110 ±10% volts

    - Minimum Breaker

      - Verify minimum 15A breaker.

  - Proper Installation

    - Machine installed on a flat, secure working place.

  - Calibration Verification

    - Calibration is complete and machine is labeled with calibration due date.

  - Verify Safety Requirements

    - Micro-switch does not turn on and lever automatically reverses to original position when foreign object is pinched.

    - Cycle ended with activation of ABORT button.

  - Equipment Operation

    - On/Off switch works properly

      - Digital display lights up with on switch.

      - Digital display turns off when off switch is pressed.

    - Start/Stop

      - Machine starts when foot pedal is pressed.

      - Machine starts when start button is pressed.

      - Operation is aborted when stop button is pressed.

    - Temperature Controller

      - Temperature reaches 400°F when machine is started.

      - Temperature reaches 140°F when machine is started.

    - Timer

      - Machine operates to 0.8 seconds when started.

      - Machine operates to 5 seconds when started.

- MATERIAL DISPOSITION

  - N/A

- DOCUMENTATION REQUIREMENTS

  - Results shall be summarized in an IQ Report.

  - Various Control Plans, including the Preventative Maintenance (PM) Schedule, safety/PPE requirements, environmental requirements, and re-validation criteria shall be documented in the final report.

  - Raw data sheets shall be included in the appendix of the IQ report and/or a reference to the location of the raw data shall be included in the IQ report.

  - Calibration records (if applicable) shall be attached to the IQ report (in the Appendix).

  - Signed IQ Protocol Training forms will be attached to the IQ report (in the Appendix).

- DEVIATIONS

  - Deviations to the IQ will be pre-approved using FRM020-012 Protocol Deviation Form and explained in the IQ report.

  - Copies of signed Protocol Deviation Forms will be attached to the IQ report (in the Appendix).

- PRE-REQUISITE STEPS

  - Approved Process Validation Plan (PVP) or justification for omission.

  - The IQ Protocol must be signed before validation activities may begin.

  - Completed Operator Training.

- APPENDIX

  - Blank data collection sheets

APPENDIX 18.1 – DATA COLLECTION SHEETS

STOP!  Read Instructions.

Instructions:

- Complete one data collection sheet (pages 7-11) for each Van der Stahl Heat Sealer.

- Verify the Pre-requisite criteria have been met on this data collection sheet prior to testing.  Note that the Pre-requisite criteria only need to be verified once for the validation.

- For attribute data: Circle “Pass” or “Fail” or write the words “Pass” or “Fail”

- For variable data: Record actual results in the units specified or add units if the units are not specified.

- Verification must include signature of the individual performing the step and the date performed.  Verification must be completed concurrently with performance.

- Follow Good Documentation Practice (GDP).

Pre-requisite Criteria (must be complete and verified prior to start of IQ activities):

| Pre-requisite | Complete (Circle One) | Verification Signature/Date |
| --- | --- | --- |
| Process Validation Plan is approved. | Yes No |  |
| IQ Protocol is signed and approved. | Yes No |  |
| Operator Training is complete and documented. | Yes No |  |

Pre-Requisites Verified By (signature & date): _____________________________________________________________

Quality Reviewed By (signature & date): ______________________________________________________________

Section A – Equipment Identification

| Equipment Name | Model Number | Serial Number | Stimlabs ID Number | Temperature Controller Model | Verification Signature/Date |
| --- | --- | --- | --- | --- | --- |

Section A Completed By (signature & date): ________________________________________________________________

Quality Reviewed By (signature & date): ______________________________________________________________

Section B - Calibration Verification

| Specification or Acceptance Criteria | Result | Verification Signature/Date |
| --- | --- | --- |
| Equipment is marked with StimLabs Equipment ID | Pass / Fail |  |
| Equipment is marked with a calibration sticker and is within current calibration | Pass / Fail |  |
| A logbook has been created for the equipment and is available for use | Pass / Fail |  |

Section B Completed By (signature & date): ________________________________________________________________

Quality Reviewed By (signature & date): ______________________________________________________________

Section C - Installation

| Specification or Acceptance Criteria | Result | Verification Signature/Date |
| --- | --- | --- |
| Machine installed on a flat, secure working place. | Pass / Fail |  |

Section C Completed By (signature & date): ________________________________________________________________

Quality Reviewed By (signature & date): ______________________________________________________________

Section D – Utility Connection

| StimLabs ID Number | Serial Number | Utility | Requirement | Result | Verification Signature/Date |
| --- | --- | --- | --- | --- | --- |
|  |  |  | 110 (±10%) Volt Actual: | Pass Fail |  |
|  |  |  | Minimum Breaker 15A Actual: | Pass Fail |  |

Record Test/Measurement Equipment Used and Calibration Due Date (attach calibration certificate to final report): __________________________________________________________________________________________________________________

__________________________________________________________________________________________________________________

Section D Completed By (Signature & date): ________________________________________________________________

Quality Reviewed By (signature & date): ______________________________________________________________

Section E – Safety Requirements

| Specification or Acceptance Criteria | Result | Verification Signature/Date |
| --- | --- | --- |
| Micro-switch does not turn on and lever automatically reverses to original position when foreign object is pinched. | Pass / Fail |  |
| Cycle ended with activation of ABORT button | Pass / Fail |  |

Section E Completed By (signature & date): ________________________________________________________________

Quality Reviewed By (signature & date): ______________________________________________________________

Section F – Equipment Controls

| Specification and Acceptance Criteria | Reading/Results | Result | Verification Signature/Date |
| --- | --- | --- | --- |
| Digital display lights up with on switch. | N/A | Pass / Fail |  |
| Digital display turns off when off switch is pressed. | N/A | Pass / Fail |  |
| Machine starts when foot pedal is pressed. | N/A | Pass / Fail |  |
| Machine starts when start button is pressed. | N/A | Pass / Fail |  |
| Operation is aborted when stop button is pressed. | N/A | Pass / Fail |  |
| Temperature Controller |  |  |  |
| Temperature set to 400°F | Enter Set-Point _____________ °F | N/A |  |
| Press start button and confirm 400°F is reached | Temperature reached: _____________ °F | Pass / Fail |  |
| Temperature set to 140°F | Enter Set-Point _____________ °F | N/A |  |
| Press start button and confirm 140°F is reached | Temperature reached: _____________ °F | Pass / Fail |  |
| Timer |  |  |  |
| Timer set to 0.8 seconds | Enter Set-Point _____________ seconds | N/A |  |
| Press start button and confirm machine operates for 0.8 seconds | Operating time _____________ seconds | Pass / Fail |  |
| Timer set to 5 seconds | Enter Set-Point _____________ seconds | N/A |  |
| Press start button and confirm machine operates for 5 seconds | Operating time _____________ seconds | Pass / Fail |  |

Record Test/Measurement Equipment Used and Calibration Due Date (attach calibration certificate to final report): __________________________________________________________________________________________________________________

__________________________________________________________________________________________________________________

Section F Completed By (signature & date): ________________________________________________________________

Quality Reviewed By (signature & date): ___________________________________________________________________