- PURPOSE

  - The purpose of this installation qualification (IQ) protocol is to define the installation qualification plan for the initial installation and commissioning of the Incucell IC 222 ECO incubator, which will be installed at the StimLabs facility in Roswell, GA. This unit is being validated as a mobile unit for utilization in any ISO 7 cleanroom at the StimLabs facility in Roswell, GA.

  - This validation is a deliverable for PCR 21-036 New Incubator.

- SCOPE

  - This installation qualification (IQ) protocol applies to the Incucell IC 222 ECO incubator. The incubator is used for the controlled storage of tissue processing solutions.

  - This equipment is located in the existing cleanroom at the StimLabs facility located in Roswell, GA.

- BACKGROUND

  - StimLabs is a tissue bank involved in the processing of human tissues for the purpose of producing finished products for transplantation. During human tissue processing, StimLabs utilizes processing solutions to perform tissue cleaning rinses. StimLabs purchased the incubator that is subject to this validation to enable the controlled storage of processing solutions and to facilitate their readiness for use in the processing of human tissues.

- REFERENCES

  - Food and Drug Administration 21 CFR Part 1271

  - American Association for Tissue Banks Standards for Tissue Banking

  - Equipment Manual

  - FDA Guidance for Industry–Process Validation: General Principles and Practices

  - GHTF/SG3/N99-10:2004 Quality Management Systems–Process Validation Guidance

  - Process Validation Plan (PVP) VAP21-020-00 Incucell IC 222 ECO Incubator

  - PCR21-036-New Incubator

- DEFINITIONS

  - Acceptance Criteria: Predetermined requirements upon which pass/fail decisions are based. Acceptance criteria are specific and measurable.

  - Critical to Quality: Attribute of a part, assembly, sub-assembly, product, or process that has direct and significant impact on its actual or perceived quality to the customer or end user.

  - Deviation: The action of departing from an established method, procedure, or acceptance criteria.

  - Installation Qualification (IQ): Establishing by objective evidence that all key aspects of the process equipment and ancillary system installation adhere to the manufacturer’s approved specification(s).

  - Process Validation: A mechanism to ensure that a process whose output is not fully verifiable by inspection and test is capable of consistently (repeatedly and reliably) producing a result or product that meets predetermined requirements. Process validation proves that a process is under control and is used to determine the process variables, acceptable limits for those variables, and to establish appropriate in-process controls.

  - Protocol: A pre-approved and prospective document that states the requirements and procedures for conducting a validation.

  - Validation/Study Leader: The individual who has the responsibility and authority over the operation of the equipment or process being validated.

  - Verification: Confirmation by examination, with provision of documented evidence, that a specific requirement has been fulfilled.

- RESPONSIBILITIES

| Function | Name | Author Protocols and Reports | Approve Protocols and Reports | Execute Protocol | WI, SOP’s, and Forms |
| --- | --- | --- | --- | --- | --- |
| Validation/Study Leader | Alexander Waters | X | X |  |  |
| Validation Reviewer | Rob Frye |  | X |  |  |
| MwER/MR/HOQ/TBD | Melissa O’Connor |  | X |  |  |
| Director of Quality | Stacy Harmon |  | X |  |  |
| Continuous Improvement/Quality Engineering Manager | Justin Junnier |  | X |  |  |
| Operations Management | Michael Baird |  | X |  |  |
| Operations Management | Dana McCullough |  |  |  | X |
| Operations | Trained Technicians |  |  | X |  |

- SAFETY CONSIDERATIONS

  - Use the correct voltage.

  - Ensure the unit is installed in an area where the floors, ceilings, walls, covers, doors, and windows are completely constructed, clean, painted, stable, without cracks, and not peeling off or falling off.

  - The area of installation should be well ventilated with low air movement, minimal traffic flow and away from doors and/or heat vents. The device is not intended for use in an atmosphere with possible danger of flammable or explosive substances.

  - Do not install the unit in a humid place or in a place where it is likely to be splashed by water.

  - Always disconnect the power supply to the device prior to any repair or maintenance.

- EQUIPMENT

  - Manufacturing Equipment Description

    - Device: Drying Oven

    - Equipment Name/Device Trade Name: Incucell

    - Manufacturer: BMT Medical Technology s.r.o.–MMM Group

    - Serial Number: H181858

    - Model: IC 222 ECO

    - Location: Existing cleanroom

  - Manufacturing Equipment Parameters

    - Volume: 222 Liters

    - Working temperature: The device’s working temperature ranges from 5℃ above ambient temperature to 100℃.

    - Target temperature for operation: 39℃±10%

    - Electronic temperature control.

    - Power source: 115 Volts ±10%

    - Frequency: 50/60 Hz ± 5%

    - Current Protection (A): Minimum is ≥ 15A

  - Test Equipment

    - Calibrated equipment shall be used for the execution of this protocol where applicable. Equipment ID/serial number and calibration information will be recorded in the data collection sheets where applicable.

    - (3) Data loggers

    - Digital multimeter

    - Flat head and Phillip’s head screwdrivers

    - Tape measure

- PROCEDURE

  - Verify Equipment Identification

    - Verify the equipment name, device, manufacturer, and model identified in this protocol and record the information in the data collection sheets.

    - Verify the serial number identified in this protocol and record the information in the data collection sheets.

  - Verify Environmental Requirements

    - Temperature: Measure the temperature of the area where the device will be utilized and record the result in the data collection sheets.

    - Relative Humidity: Measure the relative humidity of the area where the device will be utilized and record the result in the data collection sheets.

  - Verify Critical Dimensions

    - Verify the external dimensions of the unit and record the results in the data collection sheets.

  - Verify Proper Installation

    - Stabilization: The unit should not rock, lean, or be otherwise unstable.

    - Shelving: Verify that all shelving is securely and appropriately installed.

    - Verify that the incubator is installed with at least 100 mm (3.94”) of space between walls/surrounding objects and the device’s walls at the side and rear and record the results in the data collection sheets.

    - Verify the distance from the ceiling to the device’s upper wall is sufficient to allow easy access to the validation port and record the result in the data collection sheets.

    - The device shall be accessible from the front. The dimensions of the free access space for fully opening the chamber door shall be at least equal to the width of the door.

  - Verify Utility Connections

    - Measure the voltage of the incubator power source and record the result in the data collection sheets.

    - Measure the frequency of the incubator power source and record the result in the data collection sheets.

    - Measure the current protection in amperes (A) and record the result in the data collection sheets.

  - Verify Critical Components, Devices, and Systems

    - ON/OFF Push Button

      - Activates the device when pushed to the ON position. Deactivates the device when pushed to the OFF position.

    - Temperature 39℃: With the set point at 39℃, position three data logger probes in the top, middle, and bottom of the chamber to monitor its temperature. Allow the temperature to equilibrate for at least one hour after placing the probes. Verify 39℃ ± 10% is averaged on the data loggers over at least twenty-four hours. Verify the temperatures measured by each data logger agree with the set point and record the results in the data collection sheets.

    - Temperature 55℃: With the set point at 55℃, position three data logger probes in the top, middle, and bottom of the chamber to monitor its temperature. Allow the temperature to equilibrate for at least one hour after placing the probes. Verify 55℃ ± 10% is averaged on the data loggers over at least twenty-four hours. Verify the temperatures measured by each data logger agree with the set point and record the results in the data collection sheets.

- LOT SIZE & SAMPLING PLAN

  - The period of time selected to monitor processing solution temperature in the incubator is twenty-four hours. Twenty-four hours was selected based on estimated typical operations processing solution turnover rates. It is not uncommon for the stocking of a processing solution incubator to occur at the end of operations for a given day, whereby the processing solutions will be stored in the incubator overnight and, based-upon processing throughput for a given period, can provide sufficient inventory to sustain the following day’s processing operations; This can result in some of the processing solution from a single incubator stocking cycle to remain in the incubator for periods of time that can reach twenty-four hours.

  - Monitoring three zones of the incubator for twenty-four hours will ensure that the device will provide uniform, accurate, and reliable temperature regulation for periods of time that represent the high end of the spectrum of potential processing solution incubator storage times.

- ACCEPTANCE CRITERIA

  - Equipment Verification

    - Device information must match the information in the preceding Manufacturing Equipment Description section.

  - Environment Verification

    - Temperature must be between 5℃ and 40℃.

    - Relative humidity must be ≤ 80%.

  - Critical Dimensions

    - The external dimensions of the device should match the equipment specification within the tolerance described in the manufacturer’s specification, if applicable. When verifying during the execution of the IQ, if the dimensions are out of specification justification can be provided if it has no impact to functionality of the equipment or impact to the manufacturing process. Dimensions include the door with the handle and the device feet.

      - Height: 1090 mm (42.9”)

      - Width: 760 mm (29.9”)

      - Depth: 790 mm (31.1”)

  - Installation

    - Stabilization

      - The unit should not rock, lean, or be otherwise unstable.

    - Shelving

      - All shelving should be securely and appropriately installed.

    - Device Clearance

      - Distance of the device’s rear and side walls to surrounding walls/objects should be ≥ 100 mm (3.94”).

      - Distance from the device’s upper wall to the ceiling should be sufficient to provide easy access to the validation port.

      - The device shall be accessible from the front. The dimensions of the free access space for fully opening the chamber door shall be at least equal to the width of the door.

  - Utility Connections

    - Power

      - Verify power is 115V ±10%.

    - Frequency

      - Verify frequency is 50/60 Hz ± 5%.

    - Current Protection (A)

      - Verify minimum current protection is ≥ 15A.

  - Critical Components, Devices, and Systems

    - ON/OFF Push Button

      - Verify the push button activates the device when pushed to the ON position and deactivates the device when pushed to the OFF position.

    - Set Point

      - Set points of 39℃ ± 3.9℃ and 55℃ ± 5.5℃ must be obtained in the empty chamber. The data loggers’ averages over a minimum of four hours must agree, within 10%, with their respective set points.

- MATERIAL DISPOSITION

  - None

- DOCUMENTATION REQUIREMENTS

  - Results shall be summarized in the IQ report.

  - Various control plans, including the preventive maintenance (PM) schedule, safety/PPE requirements, environmental requirements, and re-validation criteria shall be documented in the IQ report.

  - Raw data sheets shall be included in the appendix of the IQ report and/or a reference to the location of the raw data shall be included in the IQ report.

  - Calibration records, where applicable, shall be included in the appendix of the IQ report.

  - Completed IQ protocol training forms will be included in the appendix of the IQ report.

- DEVIATIONS

  - Deviations from this IQ protocol will be recorded using FRM020-012, Protocol Deviation Form, and explained/justified in the IQ report.

  - Copies of completed protocol deviation forms will be included in the IQ report appendix.

- PRE-REQUISITE STEPS

  - Approved Process Validation Plan (PVP) or justification for omission.

  - The IQ protocol must be signed before validation activities may begin.

  - Completion of operator training.

- APPENDIX: DATA COLLECTION SHEETS

STOP! Read Instructions.

Instructions:

- Complete one data collection sheet for each Incucell Incubator.

- Verify the pre-requisite criteria have been met on this data collection sheet prior to testing. Note that the pre-requisite criteria only need to be verified once for the validation.

- For attribute data: Indicate “Pass” or “Fail” or write the words “Pass” or “Fail.”

- For variable data: record actual results in the units specified or add units if the units are not specified.

- Verification must include the signature of the individual performing the step and the date performed. Performance of testing and verification must be completed concurrently.

- Follow Good Documentation Practices (GDP).

Pre-requisite Criteria (must be complete and verified prior to the start of IQ activities):

| Pre-requisite | Complete (Circle One) | Verification Signature/Date |
| --- | --- | --- |
| Process Validation Plan (PVP) is approved. | Yes No |  |
| IQ Protocol is approved. | Yes No |  |
| Operator training is complete and documented. | Yes No |  |

Pre-requisites verified by (signature & date): _____________________________________________________________

Quality reviewed by (signature & date): ______________________________________________________________

Section A–Equipment Identification

| Equipment Name | Device | Model | Manufacturer | Serial Number | StimLabs ID Number | Verification Signature/Date |
| --- | --- | --- | --- | --- | --- | --- |

| Specification or Acceptance Criteria | Result |  | Verification Signature/Date |
| --- | --- | --- | --- |
| External Dimensions | Height Actual: ____________ Technical parameter: 1090 mm (42.9”) | Pass / Fail |  |
|  | Width Actual: ____________ Technical parameter: 760 mm (29.9”) | Pass / Fail |  |
|  | Depth Actual: ____________ Technical parameter: 790 mm (31.1”) | Pass / Fail |  |
| Equipment is marked with StimLabs Equipment ID. | Pass / Fail |  |  |
| Equipment is marked with a valid calibration sticker. | Pass / Fail |  |  |
| A logbook has been created for the equipment and is available for use. | Pass / Fail |  |  |

Section A completed by (signature & date): ________________________________________________________________

Quality reviewed by (signature & date): ______________________________________________________________

Section B–Installation

| Specification or Acceptance Criteria | Result |  | Verification Signature/Date |
| --- | --- | --- | --- |
| All shelving is securely and appropriately installed. | Pass / Fail |  |  |
| Equipment does not rock or lean and is not otherwise unstable. | Pass / Fail |  |  |
| There are at least 100 mm (3.94”) of space between walls/surrounding objects and the device’s walls at the side and rear. | Left Side: | Pass / Fail |  |
|  | Right Side: | Pass / Fail |  |
|  | Rear: | Pass / Fail |  |
| Distance from the device’s upper wall to the ceiling is sufficient to provide easy access to the validation port. | Pass / Fail |  |  |
| The device is accessible from the front. Sufficient space exists to allow the chamber door to open. | Pass / Fail |  |  |

Section B completed by (signature & date): ________________________________________________________________

Quality reviewed by (signature & date): ______________________________________________________________

Section C–Utility Connection

| StimLabs ID Number | Serial Number | Utility | Requirement | Result | Verification Signature/Date |
| --- | --- | --- | --- | --- | --- |
|  |  | Voltage | 115V ± 10% Actual: | Pass Fail |  |
|  |  | Frequency (Hz) | 50/60 Hz ± 5% Actual: | Pass Fail |  |
|  |  | Current Protection | ≥15A Actual: | Pass Fail |  |

Record test/measurement equipment used and calibration due dates. Attach certificates of calibration to the final report: __________________________________________________________________________________________________________________

__________________________________________________________________________________________________________________

Section C completed by (signature & date): ________________________________________________________________

Quality reviewed by (signature & date): ______________________________________________________________

Section D-Environmental Requirements

| StimLabs Asset ID | Serial Number | Chart Recorder ID Number | Result |  | Verification Signature/Date |
| --- | --- | --- | --- | --- | --- |
|  |  |  | Humidity | Temperature |  |
|  |  |  | Range: £80% Actual: Pass / Fail | Range: 5°C to 40°C Actual: Pass / Fail |  |
|  |  | Cal. Due Date: |  |  |  |

Record test/measurement equipment used and calibration due dates. Attach certificates of calibration to the final report: __________________________________________________________________________________________________________________

__________________________________________________________________________________________________________________

Section D completed by (signature & date): ________________________________________________________________

Quality reviewed by (signature & date): ______________________________________________________________

Section E–Operational Requirements

| Specification and Acceptance Criteria | Reading/Results | Result | Verification Signature/Date |
| --- | --- | --- | --- |
| ON/OFF Push Button | ON: Activates the device. | Pass / Fail |  |
|  | OFF: Deactivates the device. | Pass / Fail |  |
| Set Point Verification 39°C – Unit turned on and set to temperature below. |  |  |  |
| After a one-hour equilibration period, 39°C ± 10% must be averaged over at least twenty-four hours by the data loggers in the empty chamber. | Incubator set point _____________°C | N/A |  |
|  | Data logger ID placed: Top_____________ Middle_____________ Bottom_____________ | N/A |  |
|  | Time data loggers placed in chamber: ______________ | N/A |  |
|  | Time data loggers removed: ______________ | N/A |  |
|  | Incubator set point reading after five hours: _____________°C | Pass / Fail |  |
| Twenty-four-hour average of data loggers. | Data logger ID: Top_____________ Result: _____________°C | Pass / Fail |  |
|  | Data logger ID: Middle_____________ Result: _____________°C | Pass / Fail |  |
|  | Data logger ID: Bottom_____________ Result: _____________°C | Pass / Fail |  |
| Set Point Verification 55°C – Unit turned on and set to temperature below. |  |  |  |
| After a one-hour equilibration period, 55°C ± 10% must be averaged over at least twenty-four hours by the data loggers in the empty chamber. | Incubator set point _____________°C | N/A |  |
|  | Data Logger ID placed: Top_____________ Middle_____________ Bottom_____________ | N/A |  |
|  | Time data loggers placed in chamber: ______________ | N/A |  |
|  | Time data loggers removed: ______________ | N/A |  |
|  | Incubator set point reading after five hours: _____________°C | Pass / Fail |  |
| Twenty-four-hour average of data loggers. | Data logger ID: Top_____________ Result: _____________°C | Pass / Fail |  |
|  | Data logger ID: Middle_____________ Result: _____________°C | Pass / Fail |  |
|  | Data logger ID: Bottom_____________ Result: _____________°C | Pass / Fail |  |

Record test/measurement equipment used and calibration due dates. Attach certificates of calibration to the final report: __________________________________________________________________________________________________________________

__________________________________________________________________________________________________________________

Section E completed by (signature & date): ________________________________________________________________

Quality reviewed by (signature & date): ___________________________________________________________________