- PURPOSE

  - The purpose of this installation qualification protocol is to define the installation qualification scheme for the initial installation and commissioning of the VERIFY™ S3082 Incubator that is to be installed as a mobile unit at the Roswell, GA location of StimLabs. The equipment is to be validated as a mobile unit so that it may be utilized in any ISO 7 classified cleanroom at StimLabs.

- SCOPE

  - This IQ protocol applies to the VERIFY™ S3082 Incubator. The VERIFY™ S3082 Incubator is designed for use with VERIFY™ Self-Contained Biological Indicators (SCBIs). This equipment is utilized to provide objective evidence for determining the efficacy of sterilization processes conducted at StimLabs.

  - This equipment is located in the existing cleanroom at the StimLabs facility in Roswell, GA.

- BACKGROUND

  - StimLabs is a tissue bank involved in processing human derived tissues to generate products for transplantation. The VERIFY™ S3082 Incubator is designed to incubate VERIFY™ Self-Contained Biological Indicators (SCBIs) at 57℃ following a steam sterilization cycle.

  - Successful completion of this validation study will ensure conditions during tissue processing are met to kill a specified number of microorganisms for given sterilization processes and subsequently provide a level of confidence in the efficacy of the sterilization processes.

  - To ensure the incubator is fit for the intended purpose(s), StimLabs will perform an installation qualification and continue to monitor the performance of the incubator while it is being used to incubate VERIFY™ Self-Contained Biological Indicators (SCBIs).

  - The equipment is to be calibrated by an independent contractor that is an approved supplier prior to initiation of validation activities.

- REFERENCES

  - Food and Drug Administration 21 CFR Part 1271

  - American Association for Tissue Banks Standards for Tissue Banking

  - Equipment Manual

  - FDA Guidance for Industry – Process Validation: General Principles and Practices

  - GHTF/SG3/N99-10:2004 Quality Management Systems – Process Validation Guidance

- VAP21-025-00 Verify S3082 Incubator Process Validation Plan (PVP)

  - PCR21-036-New Incubator

- DEFINITIONS

  - Acceptance Criteria: Predetermined requirements upon which pass/fail decisions are based. Acceptance criteria are specific and measurable.

  - Biological Indicator (BI): Test systems containing viable microorganisms providing a defined resistance to a specific sterilization process.

  - Critical to Quality: Attribute of a part, assembly, sub-assembly, product or process that has direct and significant impact on its actual or perceived quality to the customer or end user.

  - Deviation: The action of departing from an established method, procedure, or acceptance criteria.

  - Installation Qualification (IQ): Establishing by objective evidence that all key aspects of the process equipment and ancillary system installation adhere to the manufacturer’s approved specification(s).

  - Process Validation: A mechanism to ensure that a process, whose output is not fully verifiable by inspection and testing, is capable of consistently (repeatedly and reliably) producing a result or product that meets predetermined requirements. Process validation proves that a process is under control and is used to determine the process variables, acceptable limits for those variables, and to establish appropriate in-process controls.

  - Protocol: A pre-approved and prospective document that states the requirements and procedures for conducting a validation.

  - Validation/Study Leader: The individual who has responsibility and authority over the operation of the equipment or process being validated.

  - Verification: Confirmation by examination, with provision of documented evidence, that a specific requirement has been fulfilled.

- RESPONSIBILITIES

| Function | Name | Author Protocols & Reports | Approve Protocols & Reports | Execute Protocol | WI, SOP’s, FORMS |
| --- | --- | --- | --- | --- | --- |
| Validation/Study Leader | Alexander Waters | X | X |  |  |
| Validation Reviewer | Rob Frye |  | X |  |  |
| MwER/MR/HOQ/TBD | Melissa O’Connor |  | X |  |  |
| Director of Quality | Stacy Harmon |  | X |  |  |
| Continuous Improvement/Quality Engineering Manager | Justin Junnier |  | X |  |  |
| Operations Management | Michael Baird |  | X |  |  |
| Operations Management | Dana McCullough |  |  |  | X |
| Operations | Trained Technicians |  |  | X |  |

- SAFETY CONSIDERATIONS

  - Do not use the equipment in the presence of flammable or combustible chemicals or gasses.

  - Do not place the equipment near air conditioning or heating vents or sources.

  - Do not install the equipment in a humid place or an area where it may be splashed by water.

  - Unplug the incubator from its power supply prior to conducting any maintenance or repair activities.

  - Ensure the equipment is installed on a level sturdy surface.

  - Use the correct utility connection(s).

- EQUIPMENT

  - Manufacturing Equipment Description

    - Manufacturer: Steris

    - Model: S3082

    - Serial number: 19041647

  - Manufacturing Equipment Parameters

    - Factory set operating range: 55℃-59℃; setpoint is 57℃.

    - Power requirements: 120V±10%; 50/60Hz

    - Ambient operating environment: 17℃-27℃, maximum 95% RH.

    - Incubator heating block(s) comprised of a precision sensor element, electronic controller, and temperature indicating liquid crystal display (LCD) strip.

  - Test Equipment

    - Calibrated equipment shall be used for the execution of this protocol where applicable. Equipment ID/serial numbers and calibration information will be documented in the data collection sheets.

    - Data logger

    - Tape measure

Digital multimeter/ voltage meter

- PROCEDURE

  - Verify Equipment Identification

    - Confirm the manufacturer, model and serial number specified in the Manufacturing Equipment Description section of this protocol. Document findings in the data collection sheets.

  - Verify Environmental Requirements

    - Record the temperature of the area of installation in the data collection sheets.

    - Record the relative humidity of the area of installation in the data collection sheets.

  - Verify Critical Dimensions

    - Verify the external dimensions of the equipment. Record the results in the data collection sheets.

  - Verify Proper Installation

    - Confirm the equipment is properly stabilized and does not rock or lean.

    - Confirm the equipment is placed away from sources of heat or cold. Confirm the area of installation is not drafty.

    - Confirm the equipment has adequate space for proper articulation of the equipment cover.

  - Verify Utility Connections

    - Record the voltage of the equipment’s power supply.

    - Record the frequency of the equipment’s power supply.

  - Verify Equipment Controls

    - Power Switch

      - After plugging the incubator into an appropriate power supply, depressing the power switch to the “ON” position results in the temperature light remaining until the temperature is within 2℃ of the setpoint, at which time the light will begin flashing. Unloaded heat up time from a cold start is ≤ 60 minutes.

      - Depressing the power switch to the “OFF” position results in the equipment de-energizing the heating block(s).

    - ???????????????????????????????????????

- LOT SIZE & SAMPLING PLAN

- ACCEPTANCE CRITERIA

- MATERIAL DISPOSITION

- DOCUMENTATION REQUIREMENTS

  - Results will be summarized in an installation qualification report.

  - Various Control Plans, including the Preventative Maintenance (PM) Schedule, safety/PPE requirements, environmental requirements, and re-validation criteria shall be documented in the final report.

  - Raw data sheets shall be included in the appendix of the IQ report and/or a reference to the location of the raw data shall be included in the IQ report.

  - Calibration records (if applicable) shall be attached to the IQ report (in the Appendix).

  - Signed IQ Protocol Training forms will be attached to the IQ report (in the Appendix).

- DEVIATIONS

  - Deviations to the IQ will be pre-approved using FRM020-012 Protocol Deviation Form and explained in the IQ report.

  - Copies of signed Protocol Deviation Forms will be attached to the IQ report (in the Appendix).

- PRE-REQUISITE STEPS

  - Approved Process Validation Plan (PVP) or justification for omission.

  - The IQ Protocol must be signed before validation activities may begin.

  - Completed Operator Training.

- APPENDIX

  - Blank training form (or use approved form)

  - Blank data collection sheets