- PURPOSE

  - The purpose of this installation qualification is to define the activities required to install the ABS Refrigerator model: ABT-HC-23-TS. The refrigerator will be utilized for storing tissue and processing supplies that are required to be maintained in a refrigerated environment prior to processing.

  - This validation is a deliverable for PCR21-028 New Critical Supplier Northside Hospital.

- SCOPE

  - This validation applies to the ABS Refrigerator model: ABT-HC-23-TS for storing placental and umbilical cord donors after a cesarean birth that are required to be maintained in a refrigerated environment prior to processing.  This validation shall serve as the initial qualification of this equipment.

  - This equipment is located at the Women’s Center at Northside Hospital Forsyth, GA.  See photo below in Figure 2.2.

Figure 2.2- ABS Refrigerator model: ABT-HC-23-TS

- BACKGROUND

  - StimLabs is a tissue bank involved in the processing of human derived cells and tissues into products for human use.  StimLabs receives tissue and other supplies that are required to be stored in a refrigerated environment prior to processing.  During the waiting period, the tissue or supplies must be held in a refrigerated environment in order to preserve their characteristics.

  - To maintain a refrigerated environment, a refrigerator must maintain its internal operating temperature between 2℃ and 10℃. Target operating temperature range for the refrigerator is 2℃ to 7℃. The set point drift must not surpass (-1℃ or +4℃) in a given timespan with the set point at 3℃ and temperature differential set to 4℃. To ensure the refrigerator is fit for the intended purpose, StimLabs will perform an installation qualification and continue to monitor the performance of the refrigerator while it is being used to store human tissue.

  - The Refrigerator will be calibrated prior to the start of the validation activities.  The set point may be adjusted prior to the start of the validation if required.

- REFERENCES

  - Food and Drug Administration 21 CFR Part 1271

  - American Association for Tissue Banks Standards for Tissue Banking

  - ABS Refrigerator model: ABT-HC-23-TS Manual -21636

  - ABS Refrigerator model: ABT-HC-23-TS Technical Specifications Sheet

  - ABS Refrigerator model: ABT-HC-23-TS Exterior Dimensions Sheet

  - ABS Refrigerator model: ABT-HC-23-TS Interior Dimensions Sheet

  - FDA Guidance for Industry – Process Validation: General Principles and Practices

  - GHTF/SG3/N99-10:2004 Quality Management Systems – Process Validation Guidance

  - VAP21-026-00 ABS Refrigerator model: ABT-HC-23-TS PVP

  - QSP-020-001Process and Equipment Validation

  - PCR21-028 Critical Supplier Northside Hospital

- DEFINITIONS

  - Acceptance Criteria:  Pre-determined requirements upon which pass/fail decisions are based.  Acceptance criteria are specific and measurable.

  - Critical to Quality:  Attribute of a part, assembly, sub-assembly, product or process that has direct and significant impact on its actual or perceived quality to the customer or end user.

  - Deviation:  The action of departing from an established method, procedure, or acceptance criteria.

  - Installation Qualification (IQ):  Establishing by objective evidence that all key aspects of the process equipment and ancillary system installation adhere to the manufacturer’s approved specification(s).

  - Process Validation:  A mechanism to ensure that a process whose output is not fully verifiable by inspection and test is capable of consistently (repeatedly and reliably) producing a result or product that meets pre-determined requirements.  Process validation proves that a process is under control and is used to determine the process variables, acceptable limits for those variables, and to establish appropriate in-process controls.

  - Protocol:  A pre-approved and prospective document that states the requirements and procedures for conducting a validation.

  - Validation/Study Leader: The individual who has the responsibility and authority over the operation of the equipment or process being validated.

  - Verification:  Confirmation by examination with a provision of documented evidence that a specific requirement has been fulfilled.

- RESPONSIBILITIES

| Function | Name | Author Protocol & Report | Review & Approve Protocols | Execute Protocol | Provide Services |
| --- | --- | --- | --- | --- | --- |
| Validation/Study Leader | Justin Junnier & Alex Waters | X | X | X |  |
| MwER/MR/HOQ/TBD | Melissa O’Connor |  | X |  |  |
| Head of Quality | Stacy Harmon |  | X |  |  |
| Operations Management | Michael Baird |  | X |  |  |
| CI/QE Management | Justin Junnier |  | X |  |  |
| Donor Services (access to hospital) | John Scott or other StimLab Rep. |  |  |  | X |
| Calibration | Qualified Supplier |  |  |  | X |
| Installation | Qualified Supplier |  |  |  | X |

- SAFETY CONSIDERATIONS

  - Do not store volatile or flammable substances in this unit except in sealed containers.

  - Never use an acid-based cleaning solution

  - Before proceeding with maintenance or checking of the unit, set the power switch to off and disconnect the power

- EQUIPMENT

  - Manufacturing Equipment Description

    - Manufacturer: ABS (Horizon Scientific)

    - Serial Number: ABS-10281607-2110

    - Model Number: ABT-HC-23-TS

    - Interior Volume: 23 cu ft

    - Temperature Range: 1°C to 10°C

    - External Dimensions: (27” wide x 82.625” height x 30.5” depth)

    - Location: Women’s Center, Northside Hospital Forsyth, GA

  - Manufacturing Equipment Parameters

    - Equipment Temperature Range: 1°C to 10°C

    - The factory set point is 3℃. “SET” is the temperature set point (compressor off point) and may be adjusted. “Hy” is the temperature differential and may be adjusted. The target operation temperature range should have the low end of the interval equal to the “SET”, while the high end of the interval will equal “SET” + “Hy”. To maintain the operation temperature between 3°C and 7°C, you would set “SET” =3°C, and “Hy” =4°C. See owner’s guide.

    - Electronic Temperature Control

    - Continuous Internal Data Logging with USB port for download

    - Voltage Rating:115V

    - Frequency Rating: 60Hz

  - Test Equipment

    - Calibrated equipment shall be used for the execution of this protocol.  Equipment ID/serial numbers and calibration information will be documented on the data collection sheet.

    - Voltage Meter (QTY: 1)

    - Tape Measure (QTY:1)

    - Tape Measure (QTY:1)

    - Cryopak iMINI, MX-HE-S-16-L Data logger (QTY:1)

      - This data logger will be used to verify the following:

        - Temperature of the refrigerator

        - The refrigerator’s Data logging capabilities. Data during the validation will be extracted and compared to the data the Cryopak iMINI to verify data retention is accurate.

- PROCEDURE

  - Verify Equipment Identification

    - Confirm the model identified in this protocol.

    - Document serial number on the data collection sheet.

  - Verify Environmental Requirement

    - Temperature – Measure the temperature in the area where the unit is installed with a Cryopak iMINI, MX-HE-S-16-L Data logger

    - Relative Humidity – Measure the relative humidity in the area where the unit is installed with a Cryopak iMINI, MX-HE-S-16-L Data logger.

    - Positioning – Verify that the unit is not exposed to direct sunlight or other heat sources.

  - Verify Critical Dimensions

    - Verify the dimensional specification of the unit.

  - Verify Proper Installation

    - Verify and document that the equipment is appropriately leveled so the door functions properly.

    - Verify and document that the casters are properly installed and adjusted in such a way that facilitates appropriate leveling of the equipment and ensures the equipment will not rock, lean, or be unstable.

    - Verify that the bottle, which houses the equipment’s thermocouple probe, in the equipment’s internal chamber, contains the appropriate thermal buffer medium.

    - Verify and document that the door gasket is clean, intact, free of defects, and forms a proper seal to the cabinet.

    - Verify and document that any shelves/trays are appropriately installed in the internal chamber of the equipment.

    - Verify and document that any light emission apparatuses, i.e., lamps, LEDs, bulbs, etc., are properly installed and illuminate appropriately.

    - Verify and document that adequate spacing, that meets or exceeds the minimum clearances expressed by the manufacturer, exists around the equipment. A clearance space of 2-3 inches around the refrigerator is specified by the manufacturer. See owner’s guide.

  - Verify Utility Connections

    - Measure the voltage of the refrigerator at the power source.

  - Verify Equipment Controls

    - Temperature – With the set-point set to 3℃, place the data logger probe positioned in the chamber to check the temperature.  Allow the temperature to equilibrate for at least one hour.  The probe inside the chamber is where simulated product temperature will be most accurate.

      - Verify a set-point 3℃ (-1℃ or +4℃), can be measured on the data logger positioned in the chamber.

      - Verify the temperature measured by the data logger agrees with the set point and display temperature on the temperature controller.

    - Verify alarm with the door open – After reaching set point, open the door of the refrigerator and wait at least 5 minutes.  Verify the refrigerator displays visual and audible alarms when the temperature exceeds 7℃.

- LOT SIZE & SAMPLING PLAN

  - None

- ACCEPTANCE CRITERIA

  - Equipment Verification

    - Must match Manufacturing Equipment Description section above.

  - Environment Verification

    - Temperature must be between 65℉ to 85℉ (18.3℃ to 29.4℃)

    - Verify unit is not installed in direct sunlight or other heat sources.

  - Dimensions

    - The dimensions of the exterior cabinet should match the equipment specification within ±1.0".

  - Installation

    - Verify that the casters are properly installed and appropriately adjusted such that appropriate leveling is attained as evidenced by proper articulation of the equipment door.

    - Verify that, upon achieving appropriate adjustment of the casters and subsequent appropriate leveling, the equipment has also achieved stability as evidenced by the equipment not rocking, leaning, or displaying any other type of foundational instability.

    - Verify that the bottle, which houses the equipment’s thermocouple probe, in the equipment’s internal chamber, contains the appropriate thermal buffer medium.

    - Verify that the data logger probe is situated in the centermost position of the equipment’s internal chamber.

    - Verify that the door gasket is clean, intact, free of defects, and forms a seal with the equipment cabinet when in the closed position.

    - Verify that any light emission apparatuses, i.e., lamps, LEDs, bulbs, etc., are properly installed and illuminate appropriately.

    - Verify that adequate spacing, that meets or exceeds the minimum clearances expressed by the manufacturer, exists around the equipment. The manufacturer specifies that a clearance space of at least 2 to 3 inches is required around the refrigerator.

  - Utility Connections

    - Power

      - Verify power is 115V ± 5%

    - Powers up

      - The interior light and displays turn on.

  - Equipment Controls

    - Set-point

      - Verify and document a set point of 3℃ and an operating temperature range equal to the set point 3℃ (-1℃ or +4℃) can be averaged on the data logger positioned in the centermost region of the equipment’s empty chamber over at least one hour the data logger (average over at least one hour), set-point, and display temperature must agree within 2℃.

        - The averaged data recorded by the data logger, the equipment set point, and the temperature displayed on the equipment’s digital temperature controller must all agree within a range of (-1℃ to +4℃) of a set point of 3℃.

    - Alarm

      - Visible and audible alarm should sound after the door has been left open and the temperature reaches 7℃. The test fails if the alarm does not sound when the display temperature exceeds 9℃.

  - Data Logger

    - Over the minimum timespan of one hour, the averaged data recorded by the data logger, the equipment set point, and the temperature displayed on the equipment’s digital temperature controller must all agree within a range of (-1℃ to +4℃) of a set point of 3℃.

- MATERIAL DISPOSITION

  - None

- DOCUMENTATION REQUIREMENTS

  - Results shall be summarized in an IQ Report.

  - Various Control Plans, including the Preventative Maintenance (PM) Schedule, safety/PPE requirements, environmental requirements, and re-validation criteria shall be documented in the final report.

  - Raw data sheets shall be included in the appendix of the IQ report and/or a reference to the location of the raw data shall be included in the IQ report.

  - Calibration records shall be attached to the IQ report (in the Appendix).

  - Signed IQ Protocol Training forms will be attached to the IQ report (in the Appendix).

- DEVIATIONS

  - Deviations to the IQ will be pre-approved using FRM020-012 Protocol Deviation Form and explained in the IQ report.

  - Copies of signed Protocol Deviation Forms will be attached to the IQ report (in the Appendix).

- PRE-REQUISITE STEPS

  - Approved Process Validation Plan (PVP).

  - The IQ Protocol must be signed before validation activities may begin with the exception of initial installation to place unit in place and verify no damage was discovered to the refrigerator upon unpacking.

- APPENDIX

  - Blank data collection sheets

APPENDIX 16.1 – DATA COLLECTION SHEETS

STOP!  Read Instructions.

Instructions:

- Complete one data collection sheet for each ABS Refrigerator model: ABT-HC-23-TS.

- Verify the Pre-requisite criteria have been met on this data collection sheet prior to testing.  Note that the Pre-requisite criteria only need to be verified once for the validation.

- For attribute data: Circle “Pass” or “Fail” or write the words “Pass” or “Fail”

- For variable data: Record actual results in the units specified or add units if the units are not specified.

- Verification must include signature of the individual performing the step and the date performed.  Verification must be completed concurrently with performance.

- Follow Good Documentation Practice (GDP).

Pre-requisite Criteria (must be complete and verified prior to start of IQ activities):

| Pre-requisite | Complete (Circle One) | Verification Signature/Date |
| --- | --- | --- |
| Process Validation Plan is approved. | Yes No |  |
| IQ Protocol is signed and approved. | Yes No |  |

Pre-Requisites Verified By (signature & date): _____________________________________________________________

Quality Reviewed By (signature & date): ___________________________________________________________________

Section A – Equipment Identification

| Equipment Name | Model Number | Manufacturer | Serial Number | Stimlabs ID Number | Verification Signature/Date |
| --- | --- | --- | --- | --- | --- |

| Specification or Acceptance Criteria | Result |  | Verification Signature/Date |
| --- | --- | --- | --- |
| External Dimensional Specifications- tolerance of ±1.0 inch is acceptable for external cabinets. | Depth: Actual: ____________ Manufacture: 30.5ʺ | Pass / Fail |  |
|  | Height: Actual: ____________ Manufacture: 82.625ʺ | Pass / Fail |  |
|  | Width: Actual: ____________ Manufacture: 27ʺ | Pass / Fail |  |
| Equipment is marked with StimLabs Equipment ID | Pass / Fail |  |  |
| Equipment is marked with a calibration sticker and is within current calibration | Pass / Fail |  |  |

Section A Completed By (signature & date): ________________________________________________________________

Quality Reviewed By (signature & date): ______________________________________________________________

Section B - Installation

| Specification or Acceptance Criteria | Result |  | Verification Signature/Date |
| --- | --- | --- | --- |
| Refrigerator is level and stable after installation of casters, so the door operates as intended. | Pass / Fail |  |  |
| Equipment does not rock, lean, or otherwise unstable Powers up: the interior light, and displays are on) | Pass / Fail |  |  |
| The internal thermocouple probe bottle is filled with the appropriate thermal buffer medium. | Pass / Fail |  |  |
| Equipment shelves/trays are appropriately installed in the internal chamber. | Pass / Fail |  |  |
| The equipment door gasket is clean, intact, free of defects, and forms a proper seal to the cabinet. | Pass / Fail |  |  |
| Light emission apparatuses are properly installed and illuminate appropriately. | Pass / Fail |  |  |
| Clearance space is ≥ 2-3” around the refrigerator. | Left Side | Pass / Fail |  |
|  | Right Side | Pass / Fail |  |
|  | Back | Pass / Fail |  |

Section B Completed By (signature & date): ________________________________________________________________

Quality Reviewed By (signature & date): ___________________________________________________________________

Section C – Utility Connection

| StimLabs ID Number | Serial Number | Utility | Requirement | Result | Verification Signature/Date |
| --- | --- | --- | --- | --- | --- |
|  |  | Power Supply Voltage | 115V ± 5% Actual: | Pass Fail |  |
|  |  | Power Supply Frequency | 60Hz Actual: | Pass Fail |  |

Record Test/Measurement Equipment Used and Calibration Due Date (attach calibration certificate to final report): __________________________________________________________________________________________________________________

__________________________________________________________________________________________________________________

Section C Completed By (Signature & date): ________________________________________________________________

Quality Reviewed By (signature & date): ______________________________________________________________

Section D - Environmental Requirements

| StimLabs ID Number | Serial Number | Cryopak Data Logger ID Number | Result | Verification Signature/Date |
| --- | --- | --- | --- | --- |
|  |  |  | Temperature |  |
|  |  |  | Range: 65℉-85℉/ 18.3℃-29.4℃ Actual: ________________ Pass / Fail |  |
|  |  | Cal Due Date: |  |  |

Record Test/Measurement Equipment Used and Calibration Due Date (attach calibration certificate(s) to final report): __________________________________________________________________________________________________________________

__________________________________________________________________________________________________________________

__________________________________________________________________________________________________________________

| Specification and Acceptance Criteria | Result | Verification Signature/Date |
| --- | --- | --- |
| The refrigerator is located away from any source to of direct heat. | Pass / Fail |  |

Section D Completed By (signature & date): ________________________________________________________________

Quality Reviewed By (signature & date): ______________________________________________________________

Section E – Operational Requirements

| Specification and Acceptance Criteria | Reading/Results | Result | Verification Signature/Date |
| --- | --- | --- | --- |
| Set Point Verification – Unit turned on and set to 3˚C. Data logger placed in Chamber |  |  |  |  |
| Visual and audible alarm initiates after the door has been left open and the temperature reaches 7˚C. | Audible alarm temp: ____________________ ˚C Visual alarm temp: ____________________ ˚C | Pass / Fail Pass / Fail |  |
| The averaged data measured by the data logger, over at least one hour in the empty chamber, agrees, within a range of -1°C to +4℃, with a set point of 3℃. | Set point reading at start of time period: _____________ ℃ | N/A |  |
|  | Time data logger placed in chamber: _____________ | N/A |  |
|  | Time data logger removed from chamber (min. one hour from time placed in chamber): ______________ | N/A |  |
|  | Set point reading at conclusion of time period: _____________℃ | Pass / Fail |  |
| The averaged data recorded by the data logger, the equipment set point, and the temperature displayed on the equipment’s digital temperature controller all agree within a range of-1℃ to +4℃ of a set point of 3℃. | Set point reading at conclusion of time period: _____________ ℃ | Pass / Fail |  |
|  | Digital temperature controller display reading at conclusion of time period: _____________℃ | Pass / Fail |  |
| Temperatures recorded should match the data logger temperature accurately over a time span of at least one hour. | Results are documented on the data logger report. | Pass / Fail |  |

Section E Completed By (signature & date): ________________________________________________________________

Quality Reviewed By (signature & date): ___________________________________________________________________