- PURPOSE

  - The purpose of this validation document is to outline the Installation Qualification (IQ) protocol for the ViewPoint Cloud Continuous Monitoring System (VPC CMS or VPC) according to the manufacturer’s (Mesa Labs, Inc. or MLI) and the customer’s (Stimlabs, LLC or Stimlabs) specifications, procedures, and requirements.

- SCOPE

  - The scope of this validation protocol will verify all VPC CMS components are appropriately located, are properly installed, and correctly displayed in the VPC system software per the equipment manufacturer’s (MLI) and the customer’s (Stimlabs) specifications, procedures, and requirements.

  - Validation of the software associated with the VPC CMS is considered out of scope. The software associated with the VPC CMS is defined as Off-the-Shelf Software (OTS) where the application was developed and validated by the manufacturer (MLI). Customizations requiring additional validation testing of the OTS are unnecessary because configurability functions, which are considered standard features of the validated OTS, will be utilized by the manufacturer in configuring the software to Stimlabs’ requirements. The VPC software is a cloud-hosted, web-based application and requires no installations on Stimlabs’ network infrastructure. A compliance certification attesting the software has been sufficiently tested and shown to possess all features and functionalities required for compliance with FDA 21CFR Part 11 will be attached as objective evidence of software validation by the manufacturer in this validation package.

- BACKGROUND

  - Stimlabs is a tissue bank involved in processing human-derived cells and tissues to generate Human Cells, Tissues, and Cellular and Tissue-Based Products (HCT/P’s). Stimlabs is responsible for human-derived cells and tissues that require specific environmental conditions throughout production of HCT/P’s. To facilitate the achievement, maintenance, and traceability of production environment requirements, Stimlabs will utilize the VPC CMS for continuous, local, and remote monitoring and management of environmentally controlled areas and equipment.

- REFERENCES

  - Food and Drug Administration 21 CFR Part 1271

  - American Association for Tissue Banks Standards for Tissue Banking

  - CAPA22-003 Data lag identified in Cobex Temperature chart recorders

  - CM-000407 Product Spec Sheet – ViewPoint Differential Pressure-Digital-R02

  - D1701 Probe Accuracy Data Sheet and Calibration (CRYO RTD)

  - DV1720 - CM-000250 – ViewPoint 900 MHz Access Point

  - DV1731 – CM-000261 – 900 MHz VPx Sensor with LCD - RTD

  - DV1732 – CM000262 – 900 MHz VPx Sensor with LCD – I2C Snap Temp-Humidity

  - FDA Guidance for Industry – Process Validation: General Principles and Practices

  - GHTF/SG3/N99-10:2004 Quality Management Systems – Process Validation Guidance

  - PCR22-007 Installation of Continuous Environmental Monitoring System

  - QSP014-001 Environmental Control and Monitoring

  - QSP020-001 Process and Equipment Validation

  - QSP022-001 Storage of Tissue and Supplies

  - SL-081212-P Performance Qualification Protocol of the Stimlabs Production Areas

  - SL-081212-R Performance Qualification Report of the Stimlabs Production Areas

  - SOP013-001 Master Safety Plan

  - SOP015-002 ViewPoint Cloud Continuous Monitoring System

  - VAP22-006-00 ViewPoint Cloud Continuous Monitoring System Process Validation Plan

  - VAP22-006-02 ViewPoint Cloud Continuous Monitoring System Operational Qualification Protocol

  - VAP22-006-03 ViewPoint Cloud Continuous Monitoring System Performance Qualification Protocol

  - VAR22-006-01 ViewPoint Cloud Continuous Monitoring System Installation Qualification Report

  - VAR22-006-02 ViewPoint Cloud Continuous Monitoring System Operational Qualification Report

  - VAR22-006-03 ViewPoint Cloud Continuous Monitoring System Performance Qualification Report

  - ViewPoint Cloud Software User Manual

  - ViewPoint Cloud v1.1.2 Compliance Certification

  - VP-00133 Installation and Operational Qualification Protocol for the ViewPoint Cloud Continuous Monitoring Systems

  - VP-00153 Performance Qualification Protocol for the ViewPoint Cloud Continuous Monitoring System

  - VPx Sensor User Manual

  - WI015-035 Operation of the ViewPoint Cloud Continuous Monitoring System

- DEFINITIONS

  - Acceptance Criteria: Predetermined requirements upon which pass/fail decisions are based. Acceptance criteria are specific and measurable.

  - Critical to Quality: Attribute of a part, assembly, sub-assembly, product, or process that has direct and significant impact on its actual or perceived quality to the customer or end user.

  - Chart recorder: Equipment used to monitor temperature and/or humidity in a specific area by recording information onto a paper disc or paper strip.

  - Controlled Environment: Defined zone in which sources of contamination are controlled by specified means.

  - Deviation: The action of departing from an established method, procedure, or acceptance criteria.

  - Environmental Control: Activities performed to control the environment for the purpose of minimizing the potential for contamination and cross-contamination of tissue.

  - Environmental Monitoring: Activities performed to systematically observe and record data to characterize the environment to identify conditions under which the potential may exist for contamination or cross-contamination of tissue.

  - HCT/Ps: Human Cells, Tissues, and Cellular and Tissue-Based Products.

  - In-process Controls: Any tests, samples, evaluations, monitoring, or measurements performed during processing or preservation that are designed to evaluate the processing or preservation procedure of the tissue subjected to processing or preservation for conformance to specification(s).

  - Installation Qualification (IQ): Establishing by objective evidence that all key aspects of the process equipment and ancillary system installation adhere to the manufacturer’s approved specification(s).

  - Management with Executive Responsibility (a.k.a. Management with Executive Responsibility/Management Representative/Head of Quality/Tissue Bank Director): Those senior employees of an establishment who have the authority to establish or make changes to the establishment’s quality policy and quality system.

  - Objective Evidence: Information, records, or statements of fact which are based on observation, measurement, or tests that can be verified.

  - Operational Qualification (OQ): Establishing by objective evidence the process control limits and action levels which result in a product that meets predetermined requirements and are within the operating limits required by the process.

  - Off-the-Shelf Software (OTS): Applications developed by a vendor for general installations. These applications may or may not be validated by the vendor for release. At a minimum, a Risk Assessment should be done to determine if the program requires additional validation testing. However, any customization of these applications may require full validation and change control based on criticality and complexity of the customization/configuration (e.g., spreadsheet macros, changes to default parameters).

  - Performance/Product Qualification (PQ): Establishing by objective evidence that the process, under anticipated conditions, consistently (repeatedly and reliably) produces a product which meets predetermined requirements.

  - Preventative Maintenance: Maintenance that is regularly performed on a piece of equipment to lessen the likelihood of it failing. Preventative maintenance is performed while the equipment is still working, so that it does not break down unexpectedly.

  - Preventive Action: Action to eliminate the cause of a potential nonconformity or other undesirable situation in order to prevent occurrence; may be performed in conjunction with corrective action(s).

  - Process Validation: Establishing by objective evidence that the output of a process consistently meets predetermined specification.

  - Protocol: A pre-approved and prospective document that states the requirements and procedures for conducting a validation.

  - Subject Matter Expert (SME): These are individuals who provide authority in a particular area. These individuals will respond to investigator’s questions pertaining to specific areas when requested. The head of each department, or his/her designee, shall serve as the subject matter expert for the department. Each expert must understand the expectations associated with participation and availability during a regulatory inspection.

  - Validation/Study Leader: The individual who has the responsibility and authority over the operation of the equipment or process being validated.

  - Verification: Confirmation by examination with a provision of documented evidence that a specific requirement has been fulfilled.

- RESPONSIBILITIES

| Function | Name | Title | Develop, Author, and Pre-Approve VP-00133 and VP-00153 | Execute, Review and Approve VP-00133 and VP-00153 | Author Stimlabs Protocols and Reports | Execute Stimlabs Protocols | Review and Approve Stimlabs Protocols and Reports |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Validation/Study Leader | Alexander Waters | Quality Improvement Engineer |  | X* | X | X** | X |
| Subject Matter Expert | Mesa Labs | VPC CMS Manufacturer/Vendor Personnel | X | X |  |  |  |
| Management with Executive Responsibility | Melissa O’ Connor | Chief Quality & Regulatory Officer |  |  |  |  | X |
| Operations Management | Joel Cummings | Vice President of Operations |  |  |  |  | X |
| Quality Management | Justin Junnier | Continuous Improvement Manager |  |  |  |  | X |

*Stimlabs personnel will assist during the execution of the pre-written protocol, VP-00133, in resolution of observations and deviations, access to secure areas, and support for the overall completion of the qualification effort as detailed in the protocol. Objective evidence of training of utilized personnel will be attached.

**Stimlabs personnel will execute all elements of Stimlabs’ protocols. Objective evidence of training of utilized personnel will be attached.

- ENVIRONMENTAL REQUIREMENTS

  - Prior to initiating protocol execution activities for this validation, environmental considerations were addressed in the temperature mapping study. Refer to the Appendix for MLI documents associated with the study. The temperature mapping study establishes the required evidence that must be documented and verified in order to demonstrate that Stimlabs production areas can maintain the necessary storage requirements under normal operating conditions according to QSP22-001, Storage of Tissue and Supplies.

  - The access points that will be installed have an operating humidity range of 0-85% RH. Calibrated data loggers and/or chart recorders will be utilized to verify the area of installation exhibits acceptable relative humidity conditions.

- SAFETY CONSIDERATIONS

  - Standard cleanroom, office, laboratory, and warehouse safety hazards that may be encountered during this validation will be addressed according to the Stimlabs Master Safety Plan, SOP013-001.

- EQUIPMENT

  - Manufacturing Equipment Description

    - Equipment Verification: Due to the scale of this IQ, equipment manufacturer, serial number, model number, sensor type (where applicable), component location, equipment association (where applicable), Stimlabs asset ID, and calibration information (calibration date and calibration expiration date) (where applicable) for the VPC CMS components utilized in this validation will be recorded in the following manner to facilitate component visibility and traceability in accordance with Stimlabs’ procedures:

      - Communication Manager: In addition to recording the Communication Manager’s Device Information in the pre-written protocol provided by the manufacturer (VP-00133 Installation and Operational Qualification Protocol for the ViewPoint Cloud Continuous Monitoring Systems), the Communication Manager’s manufacturer, model number, serial number, version, IP Address, gateway, subnet, component location, and Stimlabs asset ID will be recorded in the data collection sheets of this IQ protocol. The Communication Manager information recorded in VP-00133, this IQ protocol’s data collection sheets, and displayed in the VPC at the conclusion of IOQ execution will be cross-referenced for agreement whereby the results will be documented in Stimlabs’ IQ report for this phase of validation.

      - VPx 900MHz Access Points: In addition to recording each Access Point’s Device Information in the pre-written protocol provided by the manufacturer (VP-00133 Installation and Operational Qualification Protocol for the ViewPoint Cloud Continuous Monitoring Systems), each Access Point’s manufacturer, model number, serial number, IP Address, gateway, subnet, component location, and Stimlabs asset ID will be recorded in the data collection sheets of this IQ protocol. The Access Point information recorded in VP-00133, this IQ protocol’s data collection sheets, and displayed in the VPC at the conclusion of IOQ execution will be cross-referenced for agreement whereby the results will be documented in Stimlabs’ IQ report for this phase of validation.

      - VPx, 900MHz, Sensor LCD - I2C Snap TEMP/HUMIDITY, KIT – Low: In addition to recording the I2C Snap TEMP/HUMIDITY sensor type and quantities in the pre-written protocol provided by the manufacturer (VP-00133 Installation and Operational Qualification Protocol for the ViewPoint Cloud Continuous Monitoring Systems), each I2C Snap TEMP/HUMIDITY sensor will have its manufacturer, sensor type, serial number, component location, equipment association (where applicable), Stimlabs asset ID, and calibration information (calibration date and calibration expiration date) recorded in the Data Collection Sheet(s) of this IQ protocol. The I2C Snap TEMP/HUMIDITY sensor information recorded in VP-00133, this IQ protocol’s data collection sheets, and displayed in the VPC at the conclusion of IOQ execution will be cross-referenced for agreement whereby the results will be documented in Stimlabs’ IQ report for this phase of validation.

      - VPx, 900MHz, Sensor LCD - RTD, KIT: In addition to recording the RTD sensor type and quantities in the pre-written protocol provided by the manufacturer (VP-00133 Installation and Operational Qualification Protocol for the ViewPoint Cloud Continuous Monitoring Systems), each RTD sensor will have its manufacturer, sensor type, serial number, component location, equipment association (where applicable), Stimlabs asset ID, and calibration information (calibration date and calibration expiration date) recorded in the Data Collection Sheet(s) of this IQ protocol. The RTD sensor information recorded in VP-00133, this IQ protocol’s data collection sheets, and displayed in the VPC at the conclusion of IOQ execution will be cross-referenced for agreement whereby the results will be documented in Stimlabs’ IQ report for this phase of validation.

      - VPx, 900MHz, Sensor, with LCD - Snap Digital Differential Pressure: In addition to recording the Snap Digital Differential Pressure sensor type and quantities in the pre-written protocol provided by the manufacturer (VP-00133 Installation and Operational Qualification Protocol for the ViewPoint Cloud Continuous Monitoring Systems), each Snap Digital Differential Pressure sensor will have its manufacturer, sensor type, serial number, component location, Stimlabs asset ID, and calibration information (calibration date and calibration expiration date) recorded in the Data Collection Sheet(s) of this IQ protocol. The Snap Digital Differential Pressure sensor information recorded in VP-00133, this IQ protocol’s data collection sheets, and displayed in the VPC at the conclusion of IOQ execution will be cross-referenced for agreement whereby the results will be documented in Stimlabs’ IQ report for this phase of validation.

      - VPx, 900MHz, Sensor LCD - CRYO RTD, KIT: In addition to recording the CRYO RTD sensor type and quantities in the pre-written protocol provided by the manufacturer (VP-00133 Installation and Operational Qualification Protocol for the ViewPoint Cloud Continuous Monitoring Systems), each CRYO RTD sensor will have its manufacturer, sensor type, serial number, component location, equipment association (where applicable), Stimlabs asset ID, and calibration information (calibration date and calibration expiration date) recorded in the Data Collection Sheet(s) of this IQ protocol. The CRYO RTD sensor information recorded in VP-00133, this IQ protocol’s data collection sheets, and displayed in the VPC at the conclusion of IOQ execution will be cross-referenced for agreement whereby the results will be documented in Stimlabs’ IQ report for this phase of validation.

  - Manufacturing Equipment Parameters

    - ViewPoint Cloud Continuous Monitoring System (VPC CMS)

      - VPC is a cloud-hosted, web-based application and any standard network-enabled computer, tablet or smartphone with a web browser can access the system.

      - Communication Manager and Display Dongle Kit

      - VPx 900MHz Access Point

        - 915 MHz ISM Band, Region 2 (North & South America)

        - Full Band 906.12 MHz to 924.12 MHz

        - Low Band 902.62 MHz to 914.87 MHz

        - High Band 914.87 to 927.62 MHz

        - 12.5 dBm RF transmit power

        - Operating temperature -10 degrees Celsius<35 degrees Celsius

        - Operating humidity 0-85% RH

        - 6V DC line power connection or Power Over Ethernet (POE)

        - Static or DHCP IP configuration

      - VPx, 900MHz, Sensor LCD - I2C Snap TEMP/HUMIDITY, KIT – Low

        - Application range:10 degrees Celsius to 60 degrees Celsius and 10% RH to 80% RH

        - Accuracy: ±0.4 degrees Celsius (across full application range)

        - Accuracy: ±2.00% RH (Range: 10% ≤% RH ≤ 80% RH) (across the application range)

        - Uncertainty: ±0.25 degrees Celsius and ± 3% RH (across the application range)

        - 6V DC line or (2) 3.6V lithium-ion battery powered

          - 1 year expected battery life (estimate based on 5-minute reporting interval and dependent upon RF environment)

      - VPx, 900MHz, Sensor LCD - RTD, KIT

        - Application Range: -25 degrees Celsius to 140 degrees Celsius

        - Accuracy: ±0.35 degrees Celsius (across full application range)

        - 6V DC line or (2) 3.6V lithium-ion battery powered

          - 1 year expected battery life (estimate based on 5-minute reporting interval and dependent upon RF environment)

      - VPx, 900MHz, Sensor, with LCD - Snap Digital Differential Pressure

        - Pressure Measurement Range: -0.50 to 0.50 in H2O (inches H2O)

        - Temperature Application Range: -20 degrees Celsius to 85 degrees Celsius

        - Zero-point accuracy: 0.00032 in H2O

        - Span accuracy: 3% of reading

        - Zero-point repeatability: 0.00016 in H2O

        - Span repeatability: 0.5% of reading

        - Span shift due to temperature variation: < 0.5% of reading per 10 degrees Celsius

        - Offset stability: < 0.0002 inH2O/year

        - Flow step response time: < 3ms

      - VPx, 900MHz, Sensor LCD - CRYO RTD, KIT

        - Application Range: -200 degrees Celsius to -20 degrees Celsius

        - 6V DC line or (2) 3.6V lithium-ion battery powered

          - 1 year expected battery life (estimate based on 5-minute reporting interval and dependent upon RF environment)

  - Test Equipment

    - Refer to the Appendix for documentation on the temperature mapping study. Calibrated equipment was used in the temperature mapping study that precedes execution of the IQ portion of the protocol and provides objective evidence that directly impacts the nature of the activities performed. All calibrated equipment used in the temperature mapping study is contained in the Data Logger Calibration Verification test attachments of the study.

    - Calibrated chart recorders and/or data loggers will be utilized to verify acceptable relative humidity conditions in areas of installation where VPC components have specified operating humidity ranges. Information will be documented in the data collection sheets where applicable and supporting raw data will be included in the validation package.

- MATERIALS

  - Any standard network-enabled computer, tablet, or smartphone with a web browser.

  - Refer to the Appendix. All materials utilized in the execution of this phase of validation are documented in the data collection sheets and VP-00133.

  - Calibrated data loggers and/or chart recorders.

- TEST METHODS

  - The IQ and OQ portions of this validation are pre-written by MLI in one contiguous IOQ protocol (VP-00133 Installation and Operational Qualification Protocol for the ViewPoint Cloud Continuous Monitoring Systems). To follow Stimlabs validation procedure, QSP020-001 Process and Equipment Validation, IQ and OQ portions of the contiguous protocol will be evaluated individually for Stimlabs’ reporting purposes. Test attachments pertaining to specific phases of validation will be considered in the documentation for their respective portions and test attachments pertaining to both the IQ and OQ phases of validation will be considered in the documentation for both. Subsequently, all test attachments of the IOQ protocol will be executed sequentially and performed as written by the manufacturer. The following outlines the components of this validation that the IQ will be comprised of:

    - Installation Qualification (IQ): VP-00133 has been pre-written by MLI and will be executed by MLI with assistance from Stimlabs personnel. Refer to the Appendix for the temperature mapping study documentation. The temperature mapping study provides objective evidence that all Stimlabs areas included in the VPC installation can successfully maintain the required temperature conditions during normal operating conditions. Refer to the Appendix for protocol/data collection sheets that pertain to the IQ. The test methods performed by MLI pertinent to the IQ are listed below.

      - Protocol Execution Signature List

      - Equipment Verification

      - Installation Verification

      - Additional Information

      - Protocol Execution Summary and Acceptance

      - Deviations and Corrective Action Form Log

      - Deviation and Corrective Action Form

    - Additional equipment verification will be conducted and documented in the data collection sheets of this document.

- PROCEDURE

  - Refer to VP-00133 in the Appendix. The procedure for this phase of the validation is outlined in VP-00133 and will be executed by the manufacturer as written by the manufacturer with assistance, as needed, from Stimlabs.

  - In addition to the procedure outlined in VP-00133, the following will be performed during execution:

    - Equipment verification information will be recorded in the data collection sheets contained in the Appendix.

    - Applicable equipment verification information captured in the data collection sheets will be cross-referenced with the equipment verification information recorded in VP-00133.

    - Applicable equipment verification information from the data collection sheets and VP-00133 will be cross-referenced with equipment information available in the VPC software.

  - Areas where access points, which have specific operating humidity ranges of 0-85% RH will be monitored with calibrated data loggers and/or chart recorders during installation.

- LOT SIZE & SAMPLING PLAN

  - All components of the VPC CMS being installed will be analyzed – a sampling plan or statistical rationale is not applicable.

- ACCEPTANCE CRITERIA

  - Refer to VP-00133 in the Appendix section. The acceptance criteria for this phase of the validation is outlined in VP-00133 and will be evaluated and documented by MLI and Stimlabs therein.

  - In addition to the acceptance criteria outlined in VP-00133, the following acceptance criteria will be observed in regard to the supplemental equipment verification to be performed in this protocol:

    - The applicable equipment verification information that is cross-referenced amongst VP-00133, this protocol’s data collection sheets, and the VPC software must agree.

  - Acceptable relative humidity conditions that are within the operating humidity of the access points must be exhibited.

- MATERIAL DISPOSITION

  - There is no material disposition associated with this validation.

- DOCUMENTATION REQUIREMENTS

  - Results shall be summarized in an IQ Report.

  - Various Control Plans, including the Preventative Maintenance (PM) Schedule, safety/PPE requirements, environmental requirements, and re-validation criteria shall be documented in the final report.

  - Raw data sheets shall be included in the Appendix of the IQ report and/or a reference to the location of the raw data shall be included in the IQ report.

  - Calibration records (if applicable) shall be attached to the IQ report (in the Appendix).

  - Signed IQ Protocol Training forms will be attached to the IQ report (in the Appendix) if applicable.

- DEVIATIONS

  - Deviations to the IQ will be pre-approved using FRM020-012, Protocol Deviation Form, in addition to being documented in the Deviations and Corrective Action Form Logs and Deviation and Corrective Action Forms of the protocol provided by the manufacturer.

  - Copies of signed Protocol Deviation Forms will be attached to the IQ report (in the Appendix).

- PRE-REQUISITE STEPS

  - Approved Process Validation Plan (PVP) or justification for omission.

  - The IQ Protocol must be approved before validation activities begin.

  - The documented results of the Temperature Mapping Study must be reviewed and approved prior to initiation of IQ activities.

  - VP-00133 must be approved before validation activities may begin.

  - Completed operator training if applicable.

- APPENDIX

  - DATA COLLECTION SHEETS

  - CM-000407 Product Spec Sheet – ViewPoint Differential Pressure-Digital-R02

  - D1701 Probe Accuracy Data Sheet and Calibration (CRYO RTD)

  - DV1720 - CM-000250 – ViewPoint 900 MHz Access Point

  - DV1731 – CM-000261 – 900 MHz VPx Sensor with LCD - RTD

  - DV1732 – CM000262 – 900 MHz VPx Sensor with LCD – I2C Snap Temp-Humidity

  - SL-081212-P Performance Qualification Protocol of the Stimlabs Production Areas

  - SL-081212-R Performance Qualification Report of the Stimlabs Production Areas

  - ViewPoint Cloud Software User Manual

  - ViewPoint Cloud v1.1.2 Compliance Certification

  - VP-00133 Installation and Operational Qualification Protocol for the ViewPoint Cloud Continuous Monitoring Systems

  - VP-00153 Performance Qualification Protocol for the ViewPoint Cloud Continuous Monitoring System

  - VPx Sensor User Manual

  - Training Forms (if applicable)

  - Certificates of calibration for test equipment (if applicable)

  - Raw data sheets for test equipment (if applicable)

- REVISION HISTORY

| Revision Number | Description of Change | Changed By |
| --- | --- | --- |
| 01 | Original Issue | Alexander Waters |

DATA COLLECTION SHEETS

| Equipment Verification-Communication Manager |  |  |  |
| --- | --- | --- | --- |
| Instructions: Record the specified equipment information in the table below. If a section is not applicable to the VPC system being installed, mark it as N/A. Record any additional information in the Comments section. |  |  |  |
| ☐N/A Communication Manager |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |
|  | Model #: |  |  |
|  | Serial #: |  |  |
|  | Version: |  |  |
|  | IP Address: |  |  |
|  | Gateway: |  |  |
|  | Subnet: |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Stimlabs Asset ID: |  |  |

| Comments: |  |  |  |
| --- | --- | --- | --- |
| Verification |  | Signature | Date |
| Performed By: |  |  |  |
| Quality Reviewed By: |  |  |  |

| Equipment Verification-VPx 900MHz Access Points |  |  |  |  |
| --- | --- | --- | --- | --- |
| Instructions: Record the specified equipment information in the table below. If a section is not applicable to the VPC system being installed, mark it as N/A. Record any additional information in the Comments section. |  |  |  |  |
| ☐N/A Access Point |  |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |  |
|  | Model #: |  |  |  |
|  | Serial #: |  |  |  |
|  | IP Address: |  |  |  |
|  | Gateway: |  |  |  |
|  | Subnet: |  |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta |  | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Stimlabs Asset ID: |  |  |  |
| CHR or DAT # |  | Relative Humidity (0-85%): |  |  |
| ☐N/A Access Point |  |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |  |
|  | Model #: |  |  |  |
|  | Serial #: |  |  |  |
|  | IP Address: |  |  |  |
|  | Gateway: |  |  |  |
|  | Subnet: |  |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta |  | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Stimlabs Asset ID: |  |  |  |
| CHR or DAT # |  | Relative Humidity (0-85%): |  |  |

| Comments: |  |  |  |
| --- | --- | --- | --- |
| Verification |  | Signature | Date |
| Performed By: |  |  |  |
| Quality Reviewed By: |  |  |  |

| Equipment Verification-VPx 900MHz Access Points |  |  |  |  |
| --- | --- | --- | --- | --- |
| Instructions: Record the specified equipment information in the table below. If a section is not applicable to the VPC system being installed, mark it as N/A. Record any additional information in the Comments section. |  |  |  |  |
| ☐N/A Access Point |  |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |  |
|  | Model #: |  |  |  |
|  | Serial #: |  |  |  |
|  | IP Address: |  |  |  |
|  | Gateway: |  |  |  |
|  | Subnet: |  |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta |  | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Stimlabs Asset ID: |  |  |  |
| CHR or DAT # |  | Relative Humidity (0-85%): |  |  |
| ☐N/A Access Point |  |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |  |
|  | Model #: |  |  |  |
|  | Serial #: |  |  |  |
|  | IP Address: |  |  |  |
|  | Gateway: |  |  |  |
|  | Subnet: |  |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta |  | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Stimlabs Asset ID: |  |  |  |
| CHR or DAT # |  | Relative Humidity (0-85%): |  |  |

| Comments: |  |  |  |
| --- | --- | --- | --- |
| Verification |  | Signature | Date |
| Performed By: |  |  |  |
| Quality Reviewed By: |  |  |  |

| Equipment Verification-VPx 900MHz Access Points |  |  |  |  |
| --- | --- | --- | --- | --- |
| Instructions: Record the specified equipment information in the table below. If a section is not applicable to the VPC system being installed, mark it as N/A. Record any additional information in the Comments section. |  |  |  |  |
| ☐N/A Access Point |  |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |  |
|  | Model #: |  |  |  |
|  | Serial #: |  |  |  |
|  | IP Address: |  |  |  |
|  | Gateway: |  |  |  |
|  | Subnet: |  |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta |  | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Stimlabs Asset ID: |  |  |  |
| CHR or DAT # |  | Relative Humidity (0-85%): |  |  |
| ☐N/A Access Point |  |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |  |
|  | Model #: |  |  |  |
|  | Serial #: |  |  |  |
|  | IP Address: |  |  |  |
|  | Gateway: |  |  |  |
|  | Subnet: |  |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta |  | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Stimlabs Asset ID: |  |  |  |
| CHR or DAT # |  | Relative Humidity (0-85%): |  |  |

| Comments: |  |  |  |
| --- | --- | --- | --- |
| Verification |  | Signature | Date |
| Performed By: |  |  |  |
| Quality Reviewed By: |  |  |  |

| Equipment Verification-VPx 900MHz Access Points |  |  |  |  |
| --- | --- | --- | --- | --- |
| Instructions: Record the specified equipment information in the table below. If a section is not applicable to the VPC system being installed, mark it as N/A. Record any additional information in the Comments section. |  |  |  |  |
| ☐N/A Access Point |  |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |  |
|  | Model #: |  |  |  |
|  | Serial #: |  |  |  |
|  | IP Address: |  |  |  |
|  | Gateway: |  |  |  |
|  | Subnet: |  |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta |  | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Stimlabs Asset ID: |  |  |  |
| CHR or DAT # |  | Relative Humidity (0-85%): |  |  |
| ☐N/A Access Point |  |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |  |
|  | Model #: |  |  |  |
|  | Serial #: |  |  |  |
|  | IP Address: |  |  |  |
|  | Gateway: |  |  |  |
|  | Subnet: |  |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta |  | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Stimlabs Asset ID: |  |  |  |
| CHR or DAT # |  | Relative Humidity (0-85%): |  |  |

| Comments: |  |  |  |
| --- | --- | --- | --- |
| Verification |  | Signature | Date |
| Performed By: |  |  |  |
| Quality Reviewed By: |  |  |  |

| Equipment Verification-VPx 900MHz Access Points |  |  |  |  |
| --- | --- | --- | --- | --- |
| Instructions: Record the specified equipment information in the table below. If a section is not applicable to the VPC system being installed, mark it as N/A. Record any additional information in the Comments section. |  |  |  |  |
| ☐N/A Access Point |  |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |  |
|  | Model #: |  |  |  |
|  | Serial #: |  |  |  |
|  | IP Address: |  |  |  |
|  | Gateway: |  |  |  |
|  | Subnet: |  |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta |  | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Stimlabs Asset ID: |  |  |  |
| CHR or DAT # |  | Relative Humidity (0-85%): |  |  |
| ☐N/A Access Point |  |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |  |
|  | Model #: |  |  |  |
|  | Serial #: |  |  |  |
|  | IP Address: |  |  |  |
|  | Gateway: |  |  |  |
|  | Subnet: |  |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta |  | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Stimlabs Asset ID: |  |  |  |
| CHR or DAT # |  | Relative Humidity (0-85%): |  |  |

| Comments: |  |  |  |
| --- | --- | --- | --- |
| Verification |  | Signature | Date |
| Performed By: |  |  |  |
| Quality Reviewed By: |  |  |  |

| Equipment Verification-VPx 900MHz Access Points |  |  |  |  |
| --- | --- | --- | --- | --- |
| Instructions: Record the specified equipment information in the table below. If a section is not applicable to the VPC system being installed, mark it as N/A. Record any additional information in the Comments section. |  |  |  |  |
| ☐N/A Access Point |  |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |  |
|  | Model #: |  |  |  |
|  | Serial #: |  |  |  |
|  | IP Address: |  |  |  |
|  | Gateway: |  |  |  |
|  | Subnet: |  |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta |  | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Stimlabs Asset ID: |  |  |  |
| CHR or DAT # |  |  | Relative Humidity (0-85%): |  |  |
| ☐N/A Access Point |  |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |  |
|  | Model #: |  |  |  |
|  | Serial #: |  |  |  |
|  | IP Address: |  |  |  |
|  | Gateway: |  |  |  |
|  | Subnet: |  |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta |  | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Stimlabs Asset ID: |  |  |  |
| CHR or DAT # |  | Relative Humidity (0-85%): |  |  |

| Comments: |  |  |  |
| --- | --- | --- | --- |
| Verification |  | Signature | Date |
| Performed By: |  |  |  |
| Quality Reviewed By: |  |  |  |

| Equipment Verification- VPx 900MHz I2C Snap TEMP/HUMIDITY |  |  |  |
| --- | --- | --- | --- |
| Instructions: Record the specified equipment information in the table below. If a section is not applicable to the VPC system being installed, mark it as N/A. Record any additional information in the Comments section. |  |  |  |
| ☐N/A I2C Snap TEMP/HUMIDITY |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |
|  | Sensor Type: |  |  |
|  | Serial #: |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Equipment Association, e.g., FRI008, OVN001, FRE004, etc., if applicable. |  |  |
|  | Stimlabs Asset ID: |  |  |
|  | Calibration Date: |  |  |
|  | Calibration Expiration Date: |  |  |
| ☐N/A I2C Snap TEMP/HUMIDITY |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |
|  | Sensor Type: |  |  |
|  | Serial #: |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Equipment Association, e.g., FRI008, OVN001, FRE004, etc., if applicable. |  |  |
|  | Stimlabs Asset ID: |  |  |
|  | Calibration Date: |  |  |
|  | Calibration Expiration Date: |  |  |

| Comments: |  |  |  |
| --- | --- | --- | --- |
| Verification |  | Signature | Date |
| Performed By: |  |  |  |
| Quality Reviewed By: |  |  |  |

| Equipment Verification- VPx 900MHz I2C Snap TEMP/HUMIDITY |  |  |  |
| --- | --- | --- | --- |
| Instructions: Record the specified equipment information in the table below. If a section is not applicable to the VPC system being installed, mark it as N/A. Record any additional information in the Comments section. |  |  |  |
| ☐N/A I2C Snap TEMP/HUMIDITY |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |
|  | Sensor Type: |  |  |
|  | Serial #: |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Equipment Association, e.g., FRI008, OVN001, FRE004, etc., if applicable. |  |  |
|  | Stimlabs Asset ID: |  |  |
|  | Calibration Date: |  |  |
|  | Calibration Expiration Date: |  |  |
| ☐N/A I2C Snap TEMP/HUMIDITY |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |
|  | Sensor Type: |  |  |
|  | Serial #: |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Equipment Association, e.g., FRI008, OVN001, FRE004, etc., if applicable. |  |  |
|  | Stimlabs Asset ID: |  |  |
|  | Calibration Date: |  |  |
|  | Calibration Expiration Date: |  |  |

| Comments: |  |  |  |
| --- | --- | --- | --- |
| Verification |  | Signature | Date |
| Performed By: |  |  |  |
| Quality Reviewed By: |  |  |  |

| Equipment Verification- VPx 900MHz I2C Snap TEMP/HUMIDITY |  |  |  |
| --- | --- | --- | --- |
| Instructions: Record the specified equipment information in the table below. If a section is not applicable to the VPC system being installed, mark it as N/A. Record any additional information in the Comments section. |  |  |  |
| ☐N/A I2C Snap TEMP/HUMIDITY |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |
|  | Sensor Type: |  |  |
|  | Serial #: |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Equipment Association, e.g., FRI008, OVN001, FRE004, etc., if applicable. |  |  |
|  | Stimlabs Asset ID: |  |  |
|  | Calibration Date: |  |  |
|  | Calibration Expiration Date: |  |  |
| ☐N/A I2C Snap TEMP/HUMIDITY |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |
|  | Sensor Type: |  |  |
|  | Serial #: |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Equipment Association, e.g., FRI008, OVN001, FRE004, etc., if applicable. |  |  |
|  | Stimlabs Asset ID: |  |  |
|  | Calibration Date: |  |  |
|  | Calibration Expiration Date: |  |  |

| Comments: |  |  |  |
| --- | --- | --- | --- |
| Verification |  | Signature | Date |
| Performed By: |  |  |  |
| Quality Reviewed By: |  |  |  |

| Equipment Verification- VPx 900MHz I2C Snap TEMP/HUMIDITY |  |  |  |
| --- | --- | --- | --- |
| Instructions: Record the specified equipment information in the table below. If a section is not applicable to the VPC system being installed, mark it as N/A. Record any additional information in the Comments section. |  |  |  |
| ☐N/A I2C Snap TEMP/HUMIDITY |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |
|  | Sensor Type: |  |  |
|  | Serial #: |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Equipment Association, e.g., FRI008, OVN001, FRE004, etc., if applicable. |  |  |
|  | Stimlabs Asset ID: |  |  |
|  | Calibration Date: |  |  |
|  | Calibration Expiration Date: |  |  |
| ☐N/A I2C Snap TEMP/HUMIDITY |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |
|  | Sensor Type: |  |  |
|  | Serial #: |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Equipment Association, e.g., FRI008, OVN001, FRE004, etc., if applicable. |  |  |
|  | Stimlabs Asset ID: |  |  |
|  | Calibration Date: |  |  |
|  | Calibration Expiration Date: |  |  |

| Comments: |  |  |  |  |
| --- | --- | --- | --- | --- |
| Verification |  | Signature |  | Date |
| Performed By: |  |  |  |  |
| Quality Reviewed By: |  |  |  |  |
| Equipment Verification- VPx 900MHz I2C Snap TEMP/HUMIDITY |  |  |  |  |
| Instructions: Record the specified equipment information in the table below. If a section is not applicable to the VPC system being installed, mark it as N/A. Record any additional information in the Comments section. |  |  |  |  |
| ☐N/A I2C Snap TEMP/HUMIDITY |  |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |  |
|  | Sensor Type: |  |  |  |
|  | Serial #: |  |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta |  | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Equipment Association, e.g., FRI008, OVN001, FRE004, etc., if applicable. |  |  |  |
|  | Stimlabs Asset ID: |  |  |  |
|  | Calibration Date: |  |  |  |
|  | Calibration Expiration Date: |  |  |  |
| ☐N/A I2C Snap TEMP/HUMIDITY |  |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |  |
|  | Sensor Type: |  |  |  |
|  | Serial #: |  |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta |  | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Equipment Association, e.g., FRI008, OVN001, FRE004, etc., if applicable. |  |  |  |
|  | Stimlabs Asset ID: |  |  |  |
|  | Calibration Date: |  |  |  |
|  | Calibration Expiration Date: |  |  |  |

| Comments: |  |  |  |  |
| --- | --- | --- | --- | --- |
| Verification |  | Signature |  | Date |
| Performed By: |  |  |  |  |
| Quality Reviewed By: |  |  |  |  |
| Equipment Verification- VPx 900MHz I2C Snap TEMP/HUMIDITY |  |  |  |  |
| Instructions: Record the specified equipment information in the table below. If a section is not applicable to the VPC system being installed, mark it as N/A. Record any additional information in the Comments section. |  |  |  |  |
| ☐N/A I2C Snap TEMP/HUMIDITY |  |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |  |
|  | Sensor Type: |  |  |  |
|  | Serial #: |  |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta |  | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Equipment Association, e.g., FRI008, OVN001, FRE004, etc., if applicable. |  |  |  |
|  | Stimlabs Asset ID: |  |  |  |
|  | Calibration Date: |  |  |  |
|  | Calibration Expiration Date: |  |  |  |
| ☐N/A I2C Snap TEMP/HUMIDITY |  |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |  |
|  | Sensor Type: |  |  |  |
|  | Serial #: |  |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta |  | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Equipment Association, e.g., FRI008, OVN001, FRE004, etc., if applicable. |  |  |  |
|  | Stimlabs Asset ID: |  |  |  |
|  | Calibration Date: |  |  |  |
|  | Calibration Expiration Date: |  |  |  |

| Comments: |  |  |  |  |
| --- | --- | --- | --- | --- |
| Verification |  | Signature |  | Date |
| Performed By: |  |  |  |  |
| Quality Reviewed By: |  |  |  |  |
| Equipment Verification- VPx 900MHz I2C Snap TEMP/HUMIDITY |  |  |  |  |
| Instructions: Record the specified equipment information in the table below. If a section is not applicable to the VPC system being installed, mark it as N/A. Record any additional information in the Comments section. |  |  |  |  |
| ☐N/A I2C Snap TEMP/HUMIDITY |  |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |  |
|  | Sensor Type: |  |  |  |
|  | Serial #: |  |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta |  | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Equipment Association, e.g., FRI008, OVN001, FRE004, etc., if applicable. |  |  |  |
|  | Stimlabs Asset ID: |  |  |  |
|  | Calibration Date: |  |  |  |
|  | Calibration Expiration Date: |  |  |  |
| ☐N/A I2C Snap TEMP/HUMIDITY |  |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |  |
|  | Sensor Type: |  |  |  |
|  | Serial #: |  |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta |  | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Equipment Association, e.g., FRI008, OVN001, FRE004, etc., if applicable. |  |  |  |
|  | Stimlabs Asset ID: |  |  |  |
|  | Calibration Date: |  |  |  |
|  | Calibration Expiration Date: |  |  |  |

| Comments: |  |  |  |  |
| --- | --- | --- | --- | --- |
| Verification |  | Signature |  | Date |
| Performed By: |  |  |  |  |
| Quality Reviewed By: |  |  |  |  |
| Equipment Verification- VPx 900MHz I2C Snap TEMP/HUMIDITY |  |  |  |  |
| Instructions: Record the specified equipment information in the table below. If a section is not applicable to the VPC system being installed, mark it as N/A. Record any additional information in the Comments section. |  |  |  |  |
| ☐N/A I2C Snap TEMP/HUMIDITY |  |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |  |
|  | Sensor Type: |  |  |  |
|  | Serial #: |  |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta |  | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Equipment Association, e.g., FRI008, OVN001, FRE004, etc., if applicable. |  |  |  |
|  | Stimlabs Asset ID: |  |  |  |
|  | Calibration Date: |  |  |  |
|  | Calibration Expiration Date: |  |  |  |
| ☐N/A I2C Snap TEMP/HUMIDITY |  |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |  |
|  | Sensor Type: |  |  |  |
|  | Serial #: |  |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta |  | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Equipment Association, e.g., FRI008, OVN001, FRE004, etc., if applicable. |  |  |  |
|  | Stimlabs Asset ID: |  |  |  |
|  | Calibration Date: |  |  |  |
|  | Calibration Expiration Date: |  |  |  |

| Comments: |  |  |  |  |
| --- | --- | --- | --- | --- |
| Verification |  | Signature |  | Date |
| Performed By: |  |  |  |  |
| Quality Reviewed By: |  |  |  |  |
| Equipment Verification- VPx 900MHz I2C Snap TEMP/HUMIDITY |  |  |  |  |
| Instructions: Record the specified equipment information in the table below. If a section is not applicable to the VPC system being installed, mark it as N/A. Record any additional information in the Comments section. |  |  |  |  |
| ☐N/A I2C Snap TEMP/HUMIDITY |  |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |  |
|  | Sensor Type: |  |  |  |
|  | Serial #: |  |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta |  | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Equipment Association, e.g., FRI008, OVN001, FRE004, etc., if applicable. |  |  |  |
|  | Stimlabs Asset ID: |  |  |  |
|  | Calibration Date: |  |  |  |
|  | Calibration Expiration Date: |  |  |  |
| ☐N/A I2C Snap TEMP/HUMIDITY |  |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |  |
|  | Sensor Type: |  |  |  |
|  | Serial #: |  |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta |  | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Equipment Association, e.g., FRI008, OVN001, FRE004, etc., if applicable. |  |  |  |
|  | Stimlabs Asset ID: |  |  |  |
|  | Calibration Date: |  |  |  |
|  | Calibration Expiration Date: |  |  |  |

| Comments: |  |  |  |  |
| --- | --- | --- | --- | --- |
| Verification |  | Signature |  | Date |
| Performed By: |  |  |  |  |
| Quality Reviewed By: |  |  |  |  |
| Equipment Verification- VPx 900MHz I2C Snap TEMP/HUMIDITY |  |  |  |  |
| Instructions: Record the specified equipment information in the table below. If a section is not applicable to the VPC system being installed, mark it as N/A. Record any additional information in the Comments section. |  |  |  |  |
| ☐N/A I2C Snap TEMP/HUMIDITY |  |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |  |
|  | Sensor Type: |  |  |  |
|  | Serial #: |  |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta |  | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Equipment Association, e.g., FRI008, OVN001, FRE004, etc., if applicable. |  |  |  |
|  | Stimlabs Asset ID: |  |  |  |
|  | Calibration Date: |  |  |  |
|  | Calibration Expiration Date: |  |  |  |
| ☐N/A I2C Snap TEMP/HUMIDITY |  |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |  |
|  | Sensor Type: |  |  |  |
|  | Serial #: |  |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta |  | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Equipment Association, e.g., FRI008, OVN001, FRE004, etc., if applicable. |  |  |  |
|  | Stimlabs Asset ID: |  |  |  |
|  | Calibration Date: |  |  |  |
|  | Calibration Expiration Date: |  |  |  |

| Comments: |  |  |  |  |
| --- | --- | --- | --- | --- |
| Verification |  | Signature |  | Date |
| Performed By: |  |  |  |  |
| Quality Reviewed By: |  |  |  |  |
| Equipment Verification- VPx 900MHz I2C Snap TEMP/HUMIDITY |  |  |  |  |
| Instructions: Record the specified equipment information in the table below. If a section is not applicable to the VPC system being installed, mark it as N/A. Record any additional information in the Comments section. |  |  |  |  |
| ☐N/A I2C Snap TEMP/HUMIDITY |  |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |  |
|  | Sensor Type: |  |  |  |
|  | Serial #: |  |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta |  | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Equipment Association, e.g., FRI008, OVN001, FRE004, etc., if applicable. |  |  |  |
|  | Stimlabs Asset ID: |  |  |  |
|  | Calibration Date: |  |  |  |
|  | Calibration Expiration Date: |  |  |  |
| ☐N/A I2C Snap TEMP/HUMIDITY |  |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |  |
|  | Sensor Type: |  |  |  |
|  | Serial #: |  |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta |  | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Equipment Association, e.g., FRI008, OVN001, FRE004, etc., if applicable. |  |  |  |
|  | Stimlabs Asset ID: |  |  |  |
|  | Calibration Date: |  |  |  |
|  | Calibration Expiration Date: |  |  |  |

| Comments: |  |  |  |  |
| --- | --- | --- | --- | --- |
| Verification |  | Signature |  | Date |
| Performed By: |  |  |  |  |
| Quality Reviewed By: |  |  |  |  |
| Equipment Verification- VPx 900MHz I2C Snap TEMP/HUMIDITY |  |  |  |  |
| Instructions: Record the specified equipment information in the table below. If a section is not applicable to the VPC system being installed, mark it as N/A. Record any additional information in the Comments section. |  |  |  |  |
| ☐N/A I2C Snap TEMP/HUMIDITY |  |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |  |
|  | Sensor Type: |  |  |  |
|  | Serial #: |  |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta |  | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Equipment Association, e.g., FRI008, OVN001, FRE004, etc., if applicable. |  |  |  |
|  | Stimlabs Asset ID: |  |  |  |
|  | Calibration Date: |  |  |  |
|  | Calibration Expiration Date: |  |  |  |
| ☐N/A I2C Snap TEMP/HUMIDITY |  |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |  |
|  | Sensor Type: |  |  |  |
|  | Serial #: |  |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta |  | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Equipment Association, e.g., FRI008, OVN001, FRE004, etc., if applicable. |  |  |  |
|  | Stimlabs Asset ID: |  |  |  |
|  | Calibration Date: |  |  |  |
|  | Calibration Expiration Date: |  |  |  |

| Comments: |  |  |  |  |
| --- | --- | --- | --- | --- |
| Verification |  | Signature |  | Date |
| Performed By: |  |  |  |  |
| Quality Reviewed By: |  |  |  |  |
| Equipment Verification- VPx 900MHz RTD |  |  |  |  |
| Instructions: Record the specified equipment information in the table below. If a section is not applicable to the VPC system being installed, mark it as N/A. Record any additional information in the Comments section. |  |  |  |  |
| ☐N/A RTD |  |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |  |
|  | Sensor Type: |  |  |  |
|  | Serial #: |  |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta |  | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Equipment Association, e.g., FRI008, OVN001, FRE004, etc., if applicable. |  |  |  |
|  | Stimlabs Asset ID: |  |  |  |
|  | Calibration Date: |  |  |  |
|  | Calibration Expiration Date: |  |  |  |
| ☐N/A RTD |  |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |  |
|  | Sensor Type: |  |  |  |
|  | Serial #: |  |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta |  | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Equipment Association, e.g., FRI008, OVN001, FRE004, etc., if applicable. |  |  |  |
|  | Stimlabs Asset ID: |  |  |  |
|  | Calibration Date: |  |  |  |
|  | Calibration Expiration Date: |  |  |  |

| Comments: |  |  |  |  |
| --- | --- | --- | --- | --- |
| Verification |  | Signature |  | Date |
| Performed By: |  |  |  |  |
| Quality Reviewed By: |  |  |  |  |
| Equipment Verification- VPx 900MHz RTD |  |  |  |  |
| Instructions: Record the specified equipment information in the table below. If a section is not applicable to the VPC system being installed, mark it as N/A. Record any additional information in the Comments section. |  |  |  |  |
| ☐N/A RTD |  |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |  |
|  | Sensor Type: |  |  |  |
|  | Serial #: |  |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta |  | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Equipment Association, e.g., FRI008, OVN001, FRE004, etc., if applicable. |  |  |  |
|  | Stimlabs Asset ID: |  |  |  |
|  | Calibration Date: |  |  |  |
|  | Calibration Expiration Date: |  |  |  |
| ☐N/A RTD |  |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |  |
|  | Sensor Type: |  |  |  |
|  | Serial #: |  |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta |  | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Equipment Association, e.g., FRI008, OVN001, FRE004, etc., if applicable. |  |  |  |
|  | Stimlabs Asset ID: |  |  |  |
|  | Calibration Date: |  |  |  |
|  | Calibration Expiration Date: |  |  |  |

| Comments: |  |  |  |  |
| --- | --- | --- | --- | --- |
| Verification |  | Signature |  | Date |
| Performed By: |  |  |  |  |
| Quality Reviewed By: |  |  |  |  |
| Equipment Verification- VPx 900MHz RTD |  |  |  |  |
| Instructions: Record the specified equipment information in the table below. If a section is not applicable to the VPC system being installed, mark it as N/A. Record any additional information in the Comments section. |  |  |  |  |
| ☐N/A RTD |  |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |  |
|  | Sensor Type: |  |  |  |
|  | Serial #: |  |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta |  | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Equipment Association, e.g., FRI008, OVN001, FRE004, etc., if applicable. |  |  |  |
|  | Stimlabs Asset ID: |  |  |  |
|  | Calibration Date: |  |  |  |
|  | Calibration Expiration Date: |  |  |  |
| ☐N/A RTD |  |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |  |
|  | Sensor Type: |  |  |  |
|  | Serial #: |  |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta |  | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Equipment Association, e.g., FRI008, OVN001, FRE004, etc., if applicable. |  |  |  |
|  | Stimlabs Asset ID: |  |  |  |
|  | Calibration Date: |  |  |  |
|  | Calibration Expiration Date: |  |  |  |

| Comments: |  |  |  |  |
| --- | --- | --- | --- | --- |
| Verification |  | Signature |  | Date |
| Performed By: |  |  |  |  |
| Quality Reviewed By: |  |  |  |  |
| Equipment Verification- VPx 900MHz RTD |  |  |  |  |
| Instructions: Record the specified equipment information in the table below. If a section is not applicable to the VPC system being installed, mark it as N/A. Record any additional information in the Comments section. |  |  |  |  |
| ☐N/A RTD |  |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |  |
|  | Sensor Type: |  |  |  |
|  | Serial #: |  |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta |  | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Equipment Association, e.g., FRI008, OVN001, FRE004, etc., if applicable. |  |  |  |
|  | Stimlabs Asset ID: |  |  |  |
|  | Calibration Date: |  |  |  |
|  | Calibration Expiration Date: |  |  |  |
| ☐N/A RTD |  |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |  |
|  | Sensor Type: |  |  |  |
|  | Serial #: |  |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta |  | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Equipment Association, e.g., FRI008, OVN001, FRE004, etc., if applicable. |  |  |  |
|  | Stimlabs Asset ID: |  |  |  |
|  | Calibration Date: |  |  |  |
|  | Calibration Expiration Date: |  |  |  |

| Comments: |  |  |  |  |
| --- | --- | --- | --- | --- |
| Verification |  | Signature |  | Date |
| Performed By: |  |  |  |  |
| Quality Reviewed By: |  |  |  |  |
| Equipment Verification- VPx 900MHz RTD |  |  |  |  |
| Instructions: Record the specified equipment information in the table below. If a section is not applicable to the VPC system being installed, mark it as N/A. Record any additional information in the Comments section. |  |  |  |  |
| ☐N/A RTD |  |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |  |
|  | Sensor Type: |  |  |  |
|  | Serial #: |  |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta |  | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Equipment Association, e.g., FRI008, OVN001, FRE004, etc., if applicable. |  |  |  |
|  | Stimlabs Asset ID: |  |  |  |
|  | Calibration Date: |  |  |  |
|  | Calibration Expiration Date: |  |  |  |
| ☐N/A RTD |  |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |  |
|  | Sensor Type: |  |  |  |
|  | Serial #: |  |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta |  | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Equipment Association, e.g., FRI008, OVN001, FRE004, etc., if applicable. |  |  |  |
|  | Stimlabs Asset ID: |  |  |  |
|  | Calibration Date: |  |  |  |
|  | Calibration Expiration Date: |  |  |  |

| Comments: |  |  |  |  |
| --- | --- | --- | --- | --- |
| Verification |  | Signature |  | Date |
| Performed By: |  |  |  |  |
| Quality Reviewed By: |  |  |  |  |
| Equipment Verification- VPx 900MHz RTD |  |  |  |  |
| Instructions: Record the specified equipment information in the table below. If a section is not applicable to the VPC system being installed, mark it as N/A. Record any additional information in the Comments section. |  |  |  |  |
| ☐N/A RTD |  |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |  |
|  | Sensor Type: |  |  |  |
|  | Serial #: |  |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta |  | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Equipment Association, e.g., FRI008, OVN001, FRE004, etc., if applicable. |  |  |  |
|  | Stimlabs Asset ID: |  |  |  |
|  | Calibration Date: |  |  |  |
|  | Calibration Expiration Date: |  |  |  |
| ☐N/A RTD |  |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |  |
|  | Sensor Type: |  |  |  |
|  | Serial #: |  |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta |  | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Equipment Association, e.g., FRI008, OVN001, FRE004, etc., if applicable. |  |  |  |
|  | Stimlabs Asset ID: |  |  |  |
|  | Calibration Date: |  |  |  |
|  | Calibration Expiration Date: |  |  |  |

| Comments: |  |  |  |  |
| --- | --- | --- | --- | --- |
| Verification |  | Signature |  | Date |
| Performed By: |  |  |  |  |
| Quality Reviewed By: |  |  |  |  |
| Equipment Verification- VPx 900MHz RTD |  |  |  |  |
| Instructions: Record the specified equipment information in the table below. If a section is not applicable to the VPC system being installed, mark it as N/A. Record any additional information in the Comments section. |  |  |  |  |
| ☐N/A RTD |  |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |  |
|  | Sensor Type: |  |  |  |
|  | Serial #: |  |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta |  | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Equipment Association, e.g., FRI008, OVN001, FRE004, etc., if applicable. |  |  |  |
|  | Stimlabs Asset ID: |  |  |  |
|  | Calibration Date: |  |  |  |
|  | Calibration Expiration Date: |  |  |  |
| ☐N/A RTD |  |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |  |
|  | Sensor Type: |  |  |  |
|  | Serial #: |  |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta |  | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Equipment Association, e.g., FRI008, OVN001, FRE004, etc., if applicable. |  |  |  |
|  | Stimlabs Asset ID: |  |  |  |
|  | Calibration Date: |  |  |  |
|  | Calibration Expiration Date: |  |  |  |

| Comments: |  |  |  |  |
| --- | --- | --- | --- | --- |
| Verification |  | Signature |  | Date |
| Performed By: |  |  |  |  |
| Quality Reviewed By: |  |  |  |  |
| Equipment Verification- VPx 900MHz RTD |  |  |  |  |
| Instructions: Record the specified equipment information in the table below. If a section is not applicable to the VPC system being installed, mark it as N/A. Record any additional information in the Comments section. |  |  |  |  |
| ☐N/A RTD |  |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |  |
|  | Sensor Type: |  |  |  |
|  | Serial #: |  |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta |  | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Equipment Association, e.g., FRI008, OVN001, FRE004, etc., if applicable. |  |  |  |
|  | Stimlabs Asset ID: |  |  |  |
|  | Calibration Date: |  |  |  |
|  | Calibration Expiration Date: |  |  |  |
| ☐N/A RTD |  |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |  |
|  | Sensor Type: |  |  |  |
|  | Serial #: |  |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta |  | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Equipment Association, e.g., FRI008, OVN001, FRE004, etc., if applicable. |  |  |  |
|  | Stimlabs Asset ID: |  |  |  |
|  | Calibration Date: |  |  |  |
|  | Calibration Expiration Date: |  |  |  |

| Comments: |  |  |  |  |
| --- | --- | --- | --- | --- |
| Verification |  | Signature |  | Date |
| Performed By: |  |  |  |  |
| Quality Reviewed By: |  |  |  |  |
| Equipment Verification- VPx 900MHz RTD |  |  |  |  |
| Instructions: Record the specified equipment information in the table below. If a section is not applicable to the VPC system being installed, mark it as N/A. Record any additional information in the Comments section. |  |  |  |  |
| ☐N/A RTD |  |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |  |
|  | Sensor Type: |  |  |  |
|  | Serial #: |  |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta |  | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Equipment Association, e.g., FRI008, OVN001, FRE004, etc., if applicable. |  |  |  |
|  | Stimlabs Asset ID: |  |  |  |
|  | Calibration Date: |  |  |  |
|  | Calibration Expiration Date: |  |  |  |
| ☐N/A RTD |  |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |  |
|  | Sensor Type: |  |  |  |
|  | Serial #: |  |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta |  | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Equipment Association, e.g., FRI008, OVN001, FRE004, etc., if applicable. |  |  |  |
|  | Stimlabs Asset ID: |  |  |  |
|  | Calibration Date: |  |  |  |
|  | Calibration Expiration Date: |  |  |  |

| Comments: |  |  |  |  |
| --- | --- | --- | --- | --- |
| Verification |  | Signature |  | Date |
| Performed By: |  |  |  |  |
| Quality Reviewed By: |  |  |  |  |
| Equipment Verification- VPx 900MHz RTD |  |  |  |  |
| Instructions: Record the specified equipment information in the table below. If a section is not applicable to the VPC system being installed, mark it as N/A. Record any additional information in the Comments section. |  |  |  |  |
| ☐N/A RTD |  |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |  |
|  | Sensor Type: |  |  |  |
|  | Serial #: |  |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta |  | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Equipment Association, e.g., FRI008, OVN001, FRE004, etc., if applicable. |  |  |  |
|  | Stimlabs Asset ID: |  |  |  |
|  | Calibration Date: |  |  |  |
|  | Calibration Expiration Date: |  |  |  |
| ☐N/A RTD |  |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |  |
|  | Sensor Type: |  |  |  |
|  | Serial #: |  |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta |  | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Equipment Association, e.g., FRI008, OVN001, FRE004, etc., if applicable. |  |  |  |
|  | Stimlabs Asset ID: |  |  |  |
|  | Calibration Date: |  |  |  |
|  | Calibration Expiration Date: |  |  |  |

| Comments: |  |  |  |  |
| --- | --- | --- | --- | --- |
| Verification |  | Signature |  | Date |
| Performed By: |  |  |  |  |
| Quality Reviewed By: |  |  |  |  |
| Equipment Verification- VPx 900MHz RTD |  |  |  |  |
| Instructions: Record the specified equipment information in the table below. If a section is not applicable to the VPC system being installed, mark it as N/A. Record any additional information in the Comments section. |  |  |  |  |
| ☐N/A RTD |  |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |  |
|  | Sensor Type: |  |  |  |
|  | Serial #: |  |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta |  | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Equipment Association, e.g., FRI008, OVN001, FRE004, etc., if applicable. |  |  |  |
|  | Stimlabs Asset ID: |  |  |  |
|  | Calibration Date: |  |  |  |
|  | Calibration Expiration Date: |  |  |  |
| ☐N/A RTD |  |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |  |
|  | Sensor Type: |  |  |  |
|  | Serial #: |  |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta |  | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Equipment Association, e.g., FRI008, OVN001, FRE004, etc., if applicable. |  |  |  |
|  | Stimlabs Asset ID: |  |  |  |
|  | Calibration Date: |  |  |  |
|  | Calibration Expiration Date: |  |  |  |

| Comments: |  |  |  |  |
| --- | --- | --- | --- | --- |
| Verification |  | Signature |  | Date |
| Performed By: |  |  |  |  |
| Quality Reviewed By: |  |  |  |  |
| Equipment Verification- VPx 900MHz RTD |  |  |  |  |
| Instructions: Record the specified equipment information in the table below. If a section is not applicable to the VPC system being installed, mark it as N/A. Record any additional information in the Comments section. |  |  |  |  |
| ☐N/A RTD |  |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |  |
|  | Sensor Type: |  |  |  |
|  | Serial #: |  |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta |  | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Equipment Association, e.g., FRI008, OVN001, FRE004, etc., if applicable. |  |  |  |
|  | Stimlabs Asset ID: |  |  |  |
|  | Calibration Date: |  |  |  |
|  | Calibration Expiration Date: |  |  |  |
| ☐N/A RTD |  |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |  |
|  | Sensor Type: |  |  |  |
|  | Serial #: |  |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta |  | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Equipment Association, e.g., FRI008, OVN001, FRE004, etc., if applicable. |  |  |  |
|  | Stimlabs Asset ID: |  |  |  |
|  | Calibration Date: |  |  |  |
|  | Calibration Expiration Date: |  |  |  |

| Comments: |  |  |  |  |
| --- | --- | --- | --- | --- |
| Verification |  | Signature |  | Date |
| Performed By: |  |  |  |  |
| Quality Reviewed By: |  |  |  |  |
| Equipment Verification- VPx 900MHz Snap Digital Differential Pressure |  |  |  |  |
| Instructions: Record the specified equipment information in the table below. If a section is not applicable to the VPC system being installed, mark it as N/A. Record any additional information in the Comments section. |  |  |  |  |
| ☐N/A Snap Digital Differential Pressure |  |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |  |
|  | Sensor Type: |  |  |  |
|  | Serial #: |  |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta |  | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Stimlabs Asset ID: |  |  |  |
|  | Calibration Date: |  |  |  |
|  | Calibration Expiration Date: |  |  |  |
| ☐N/A Snap Digital Differential Pressure |  |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |  |
|  | Sensor Type: |  |  |  |
|  | Serial #: |  |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta |  | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Stimlabs Asset ID: |  |  |  |
|  | Calibration Date: |  |  |  |
|  | Calibration Expiration Date: |  |  |  |

| Comments: |  |  |  |
| --- | --- | --- | --- |
| Verification |  | Signature | Date |
| Performed By: |  |  |  |
| Quality Reviewed By: |  |  |  |

| Equipment Verification- VPx 900MHz Snap Digital Differential Pressure |  |  |  |
| --- | --- | --- | --- |
| Instructions: Record the specified equipment information in the table below. If a section is not applicable to the VPC system being installed, mark it as N/A. Record any additional information in the Comments section. |  |  |  |
| ☐N/A Snap Digital Differential Pressure |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |
|  | Sensor Type: |  |  |
|  | Serial #: |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Stimlabs Asset ID: |  |  |
|  | Calibration Date: |  |  |
|  | Calibration Expiration Date: |  |  |
| ☐N/A Snap Digital Differential Pressure |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |
|  | Sensor Type: |  |  |
|  | Serial #: |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Stimlabs Asset ID: |  |  |
|  | Calibration Date: |  |  |
|  | Calibration Expiration Date: |  |  |

| Comments: |  |  |  |
| --- | --- | --- | --- |
| Verification |  | Signature | Date |
| Performed By: |  |  |  |
| Quality Reviewed By: |  |  |  |

| Equipment Verification- VPx 900MHz Snap Digital Differential Pressure |  |  |  |
| --- | --- | --- | --- |
| Instructions: Record the specified equipment information in the table below. If a section is not applicable to the VPC system being installed, mark it as N/A. Record any additional information in the Comments section. |  |  |  |
| ☐N/A Snap Digital Differential Pressure |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |
|  | Sensor Type: |  |  |
|  | Serial #: |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Stimlabs Asset ID: |  |  |
|  | Calibration Date: |  |  |
|  | Calibration Expiration Date: |  |  |
| ☐N/A Snap Digital Differential Pressure |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |
|  | Sensor Type: |  |  |
|  | Serial #: |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Stimlabs Asset ID: |  |  |
|  | Calibration Date: |  |  |
|  | Calibration Expiration Date: |  |  |

| Comments: |  |  |  |
| --- | --- | --- | --- |
| Verification |  | Signature | Date |
| Performed By: |  |  |  |
| Quality Reviewed By: |  |  |  |

| Equipment Verification- VPx 900MHz Snap Digital Differential Pressure |  |  |  |
| --- | --- | --- | --- |
| Instructions: Record the specified equipment information in the table below. If a section is not applicable to the VPC system being installed, mark it as N/A. Record any additional information in the Comments section. |  |  |  |
| ☐N/A Snap Digital Differential Pressure |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |
|  | Sensor Type: |  |  |
|  | Serial #: |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Stimlabs Asset ID: |  |  |
|  | Calibration Date: |  |  |
|  | Calibration Expiration Date: |  |  |
| ☐N/A Snap Digital Differential Pressure |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |
|  | Sensor Type: |  |  |
|  | Serial #: |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Stimlabs Asset ID: |  |  |
|  | Calibration Date: |  |  |
|  | Calibration Expiration Date: |  |  |

| Comments: |  |  |  |
| --- | --- | --- | --- |
| Verification |  | Signature | Date |
| Performed By: |  |  |  |
| Quality Reviewed By: |  |  |  |

| Equipment Verification- VPx 900MHz Snap Digital Differential Pressure |  |  |  |
| --- | --- | --- | --- |
| Instructions: Record the specified equipment information in the table below. If a section is not applicable to the VPC system being installed, mark it as N/A. Record any additional information in the Comments section. |  |  |  |
| ☐N/A Snap Digital Differential Pressure |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |
|  | Sensor Type: |  |  |
|  | Serial #: |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Stimlabs Asset ID: |  |  |
|  | Calibration Date: |  |  |
|  | Calibration Expiration Date: |  |  |
| ☐N/A Snap Digital Differential Pressure |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |
|  | Sensor Type: |  |  |
|  | Serial #: |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Stimlabs Asset ID: |  |  |
|  | Calibration Date: |  |  |
|  | Calibration Expiration Date: |  |  |

| Comments: |  |  |  |
| --- | --- | --- | --- |
| Verification |  | Signature | Date |
| Performed By: |  |  |  |
| Quality Reviewed By: |  |  |  |

| Equipment Verification- VPx 900MHz CRYO RTD |  |  |  |
| --- | --- | --- | --- |
| Instructions: Record the specified equipment information in the table below. If a section is not applicable to the VPC system being installed, mark it as N/A. Record any additional information in the Comments section. |  |  |  |
| ☐N/A CRYO RTD |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |
|  | Sensor Type: |  |  |
|  | Serial #: |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Equipment Association, e.g., FRI008, OVN001, FRE004, etc., if applicable. |  |  |
|  | Stimlabs Asset ID: |  |  |
|  | Calibration Date: |  |  |
|  | Calibration Expiration Date: |  |  |
| ☐N/A CRYO RTD |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |
|  | Sensor Type: |  |  |
|  | Serial #: |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Equipment Association, e.g., FRI008, OVN001, FRE004, etc., if applicable. |  |  |
|  | Stimlabs Asset ID: |  |  |
|  | Calibration Date: |  |  |
|  | Calibration Expiration Date: |  |  |

| Comments: |  |  |  |  |
| --- | --- | --- | --- | --- |
| Verification |  | Signature |  | Date |
| Performed By: |  |  |  |  |
| Quality Reviewed By: |  |  |  |  |
| Equipment Verification- VPx 900MHz CRYO RTD |  |  |  |  |
| Instructions: Record the specified equipment information in the table below. If a section is not applicable to the VPC system being installed, mark it as N/A. Record any additional information in the Comments section. |  |  |  |  |
| ☐N/A CRYO RTD |  |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |  |
|  | Sensor Type: |  |  |  |
|  | Serial #: |  |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta |  | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Equipment Association, e.g., FRI008, OVN001, FRE004, etc., if applicable. |  |  |  |
|  | Stimlabs Asset ID: |  |  |  |
|  | Calibration Date: |  |  |  |
|  | Calibration Expiration Date: |  |  |  |
| ☐N/A CRYO RTD |  |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |  |
|  | Sensor Type: |  |  |  |
|  | Serial #: |  |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta |  | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Equipment Association, e.g., FRI008, OVN001, FRE004, etc., if applicable. |  |  |  |
|  | Stimlabs Asset ID: |  |  |  |
|  | Calibration Date: |  |  |  |
|  | Calibration Expiration Date: |  |  |  |

| Comments: |  |  |  |  |
| --- | --- | --- | --- | --- |
| Verification |  | Signature |  | Date |
| Performed By: |  |  |  |  |
| Quality Reviewed By: |  |  |  |  |
| Equipment Verification- VPx 900MHz CRYO RTD |  |  |  |  |
| Instructions: Record the specified equipment information in the table below. If a section is not applicable to the VPC system being installed, mark it as N/A. Record any additional information in the Comments section. |  |  |  |  |
| ☐N/A CRYO RTD |  |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |  |
|  | Sensor Type: |  |  |  |
|  | Serial #: |  |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta |  | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Equipment Association, e.g., FRI008, OVN001, FRE004, etc., if applicable. |  |  |  |
|  | Stimlabs Asset ID: |  |  |  |
|  | Calibration Date: |  |  |  |
|  | Calibration Expiration Date: |  |  |  |
| ☐N/A CRYO RTD |  |  |  | Verified By/Date |
| Device Information | Manufacturer: |  |  |  |
|  | Sensor Type: |  |  |  |
|  | Serial #: |  |  |  |
|  | Component Location: ☐1125 Northmeadow Parkway ☐1225 Northmeadow Parkway ☐Northside Hospital Forsyth ☐Northside Hospital Cherokee ☐Northside Hospital Atlanta |  | Specify component location, e.g., Room 104, QC Lab, etc.: |  |
|  | Equipment Association, e.g., FRI008, OVN001, FRE004, etc., if applicable. |  |  |  |
|  | Stimlabs Asset ID: |  |  |  |
|  | Calibration Date: |  |  |  |
|  | Calibration Expiration Date: |  |  |  |

| Comments: |  |  |  |  |
| --- | --- | --- | --- | --- |
| Verification |  |  | Signature |  | Date |
| Performed By: |  |  |  |  |  |
| Quality Reviewed By: |  |  |  |  |  |
| Equipment Verification Cross Reference for Agreement |  |  |  |  |
| Instructions: Cross-reference the applicable equipment verification information documented in VP-00133, in the preceding data collection sheets, and available in the VPC software. Circle YES or NO if the applicable equipment verification information that is available agrees (YES) or disagrees (NO). Record any additional information in the Comments section. |  |  |  |  |
|  |  |  |  | Verified By/Date |
| Does the applicable equipment verification information recorded in the data collection sheets agree with the applicable equipment verification information recorded in VP-00133? |  |  | YES NO |  |
| Does the applicable equipment verification information recorded in the data collection sheets agree with the applicable equipment verification information available in the VPC application? |  |  | YES NO |  |
| Does the applicable equipment verification information recorded in VP-00133 agree with the applicable equipment verification information available in the VPC application? |  |  | YES NO |  |
| Comments: |  |  |  |  |
| Verification |  |  | Signature |  | Date |
| Performed By: |  |  |  |  |  |
| Quality Reviewed By: |  |  |  |  |  |