- PURPOSE

  - The purpose of this document is to outline the installation qualification protocol for Incucell 404-ECO Line incubator(s) according to the manufacturer’s (BMT Medical Technology s.r.o., or BMT) and the customer’s (Stimlabs, LLC or Stimlabs) specifications, procedures, and requirements.

- SCOPE

  - The scope of this validation protocol will verify Incucell 404-ECO Line incubator(s) (Figure 1) is/are appropriately located, properly installed, and correctly configured per the equipment manufacturer’s (BMT) and the customer’s (Stimlabs) specifications, procedures, and requirements. During processing of human-derived cells and tissues, to generate Human Cells, Tissues, and Cellular and Tissue-Based Products (HCT/P’s), Stimlabs utilizes conditioned solutions; the equipment subject to this validation package will be used to condition solutions prior to utilization.

  - Prior to utilizing solutions conditioned in the equipment, calibrated thermometers are utilized to measure the temperature of the solutions; only solutions that satisfy Stimlabs’ requirements are utilized in HCT/P production. Because this process exhibits verification of the performance of the equipment by measuring and documenting the temperature of all solutions utilized in HCT/P production, neither an Operational Qualification (OQ) nor a Performance Qualification (PQ) is required to ensure that the equipment subject to this validation is fit for the intended use.

Figure 1

- BACKGROUND

  - Stimlabs is a tissue bank involved in processing human-derived cells and tissues to generate Human Cells, Tissues, and Cellular and Tissue-Based Products (HCT/P’s). Processing human-derived cells and tissues involves utilization of conditioned solutions in Stimlabs’ methodologies to produce HCT/P’s. The equipment that is subject to this validation package will be utilized to enable Stimlabs to condition solutions according to the appropriate specifications, procedures, and requirements.

- REFERENCES

  - Food and Drug Administration 21 CFR Part 1271

  - American Association for Tissue Banks Standards for Tissue Banking

  - FDA Guidance for Industry – Process Validation: General Principles and Practices

  - GHTF/SG3/N99-10:2004 Quality Management Systems – Process Validation Guidance

  - INCUCELL 404 Operating Instructions

  - INCUCELL/INCUCELL V Laboratory Incubators with Natural/Forced Air Convention

  - Laboratory Drying Ovens and Incubators with innovated control automatics ECO line and EVO line

  - SOP013-001 Master Safety Plan

  - SOP015-001 Equipment Repair and Maintenance

  - SOP020-001 Master Validation Plan

  - QSP020-001 Process and Equipment Validation

  - VAP22-007-00 Incucell 404-ECO Line Process Validation Plan

  - VAR22-007-01 Incucell 404-ECO Line Installation Qualification Report

- DEFINITIONS

  - Acceptance Criteria: Predetermined requirements upon which pass/fail decisions are based. Acceptance criteria are specific and measurable.

  - Critical to Quality: Attribute of a part, assembly, sub-assembly, product, or process that has direct and significant impact on its actual or perceived quality to the customer or end user.

  - Controlled Environment: Defined zone in which sources of contamination are controlled by specified means.

  - Deviation: The action of departing from an established method, procedure, or acceptance criteria.

  - Environmental Control: Activities performed to control the environment for the purpose of minimizing the potential for contamination and cross-contamination of tissue.

  - HCT/P’s: Human Cells, Tissues, and Cellular and Tissue-Based Products.

  - In-process Controls: Any tests, samples, evaluations, monitoring, or measurements performed during processing or preservation that are designed to evaluate the processing or preservation procedure of the tissue subjected to processing or preservation for conformance to specification(s).

  - Installation Qualification (IQ): Establishing by objective evidence that all key aspects of the process equipment and ancillary system installation adhere to the manufacturer’s approved specification(s).

  - Management with Executive Responsibility (a.k.a. Management with Executive Responsibility/Management Representative/Head of Quality/Tissue Bank Director): Those senior employees of an establishment who have the authority to establish or make changes to the establishment’s quality policy and quality system.

  - Objective Evidence: Information, records, or statements of fact which are based on observation, measurement, or tests that can be verified.

  - Operational Qualification (OQ): Establishing by objective evidence the process control limits and action levels which result in a product that meets predetermined requirements and are within the operating limits required by the process.

  - Performance/Product Qualification (PQ): Establishing by objective evidence that the process, under anticipated conditions, consistently (repeatedly and reliably) produces a product which meets predetermined requirements.

  - Preventative Maintenance: Maintenance that is regularly performed on a piece of equipment to lessen the likelihood of it failing. Preventative maintenance is performed while the equipment is still working, so that it does not break down unexpectedly.

  - Processing Materials: Materials that are used in the manufacture or inspection of product and may come into direct contact with the product.

  - Process Validation: Establishing by objective evidence that the output of a process consistently meets predetermined specification.

  - Protocol: A pre-approved and prospective document that states the requirements and procedures for conducting a validation.

  - Validation/Study Leader: The individual who has the responsibility and authority over the operation of the equipment or process being validated.

  - Verification: Confirmation by examination with a provision of documented evidence that a specific requirement has been fulfilled.

- RESPONSIBILITIES

| Function | Name | Title | Author Protocols and Reports | Execute Stimlabs Protocols | Review and Approve Stimlabs Protocols and Reports |
| --- | --- | --- | --- | --- | --- |
| Validation/Study Leader | Stimlabs Personnel | Quality Engineer or Equivalent | X | X | X |
| Management with Executive Responsibility | Melissa O’ Connor | Chief Quality & Regulatory Officer |  |  | X |
| Quality Management | Stimlabs Personnel | Quality Engineering Manager or Equivalent |  |  | X |
| Operations Management | Stimlabs Personnel | Operations Management or Equivalent |  |  | X |

- SAFETY CONSIDERATIONS

  - Standard cleanroom and laboratory safety hazards that may be encountered during this validation will be addressed according to SOP013-001, Stimlabs Master Safety Plan.

  - 100 mm distance to the walls at the sides and back of the equipment must be sustained. The temperature of the air coming out of the exhaust may be up to 250 degrees Celsius, so the walls near the oven must be non-flammable. No flammable or volatile/combustible materials may be put into the equipment.

  - Goods are only to be put on trays in the equipment. Never place goods directly on the bottom of the equipment.

  - Any assembly/disassembly may only be done when the equipment is disconnected from the mains.

- EQUIPMENT

  - Manufacturing Equipment Description

    - Manufacturer: BMT Medical Technology s.r.o., or BMT

    - Serial Number: Serial number(s) will be verified and documented during execution of this protocol in the data collection sheets. See the Appendix for data collection sheets.

    - Model: INCUCELL 404-ECO Line

  - Manufacturing Equipment Parameters

    - Operating Range: See INCUCELL/INCUCELL V Laboratory Incubators with Natural/Forced Air Convention in the Appendix for operation temperature range.

    - Equipment Controls: See INCUCELL 404 Operating Instructions in the Appendix for equipment controls in the Control Panel section of the Equipment Description.

    - Connections and Utility Requirements: See INCUCELL/INCUCELL V Laboratory Incubators with Natural/Forced Air Convention in the Appendix for electrical data.

  - Test Equipment

    - Calibrated Thermometer Hygrometer (Data Logger, Chart Recorder, CMS, etc.)

      - One (1) calibrated data logger, chart recorder, or CMS will be utilized to verify the equipment Environmental Requirements.

      - Three (3) calibrated data loggers will be utilized to verify the equipment Operational Requirements.

    - A calibrated digital multimeter will be utilized to verify the mains connection for the equipment.

    - A tape meter (calibration not necessary) will be utilized to verify the external dimensions of the equipment.

- MATERIALS

  - Calibrated thermometer hygrometer(s)

  - Calibrated digital multimeter

  - Tape meter

- TEST METHODS

  - The following test methods will be utilized to satisfy Stimlabs’ specifications, procedures, and requirements during the INCUCELL 404-ECO Line Installation Qualification:

    - Pre-requisite Criteria Verification

    - Equipment Identification Verification

    - Equipment Installation Testing

    - Equipment Utilities Testing

    - Environmental Requirements Testing

    - Operational Requirements Testing

- PROCEDURE

  - Verify the pre-requisite criteria.

    - Verify the pre-requisite criteria by ensuring the process validation plan relevant to this validation package is approved and release, this installation qualification protocol is approved and released, and objective evidence of training on this protocol is present for all personnel participating in the qualification activities.

  - Verify Equipment Identification

    - Verify the identification of the equipment that is subject to this validation package by documenting the device trade name, device model, the manufacturer, the serial number, and the Stimlabs asset ID.

  - Verify Equipment Installation

    - Verify the installation of the equipment subject to this validation package by documenting the external dimensions, the installation clearances, equipment accessibility, critical component condition, and seating.

  - Verify Utility Connections

    - Verify equipment utilities by documenting the current type, phase, voltage, frequency, and current protection of the mains the equipment that is subject to this validation package will be connected to.

    - Document test equipment information for the digital multimeter that is required to perform this verification.

  - Verify Environmental Requirements

    - Verify the environmental requirements for the area of installation of the equipment that is subject to this validation package by recording the temperature and relative humidity of the area.

    - Document test equipment information for the thermometer hygrometer (data logger, chart recorder, or CMS) that is required to perform this verification

  - Verify operational requirements of the equipment that is subject to this validation package by:

    - Ensuring the device display functions properly when the equipment is powered on.

    - Ensuring the device “SET:” and internal chamber temperature on the display agree within 10% after allowing sufficient equilibration time following powering on the equipment.

    - Ensuring average temperatures recorded over the span of a sufficient time and internal chamber space that are indicative of normal operations are within 10% of the devices “SET:”.

      - Document test equipment information for the data loggers that are required to perform this verification.

  - Verify the results of the qualification activities by documenting, in summary, the outcome of all tests.

- LOT SIZE & SAMPLING PLAN

  - Statistical rationale/sampling plan is not applicable.

- ACCEPTANCE CRITERIA

  - Acceptance criteria and evaluation of all acceptance criteria for all testing conducted in this installation qualification is outlined in the respective data collection sheet for each test method and summarized in the Test Summary data collection sheet of the Appendix.

- MATERIAL DISPOSITION

  - There is no material disposition associated with this validation.

- DOCUMENTATION REQUIREMENTS

  - Results shall be summarized in an IQ Report.

  - Various Control Plans, including the Preventative Maintenance (PM) Schedule, safety/PPE requirements, environmental requirements, and re-validation criteria shall be documented in the final report.

  - Raw data sheets shall be included in the Appendix of the IQ report and/or a reference to the location of the raw data shall be included in the IQ report.

  - Calibration records (if applicable) shall be attached to the IQ report (in the Appendix).

  - Signed IQ Protocol Training forms will be attached to the IQ report (in the Appendix) if applicable.

- DEVIATIONS

  - Deviations will be documented using FRM020-012, Protocol Deviation Form

- PRE-REQUISITE STEPS

  - For each phase of validation, Stimlabs protocols must be approved and signed prior to the start of validation activities.

  - Per QSP020-001, Process and Equipment Validation, a validation report shall be completed and approved at the conclusion of each phase of the validation and a final summary report encompassing all phases of validation may be prepared in addition to the individual phase reports.

  - Objective evidence, via FRM011-001, Individual Training Form or FRM011-003, Group Training Form, of training is required for all persons participating in the qualification activities that are subject to this validation.

  - Calibration of the equipment subject to this validation package must occur prior to approval of the IQ report. Objective evidence of the calibration must be present in this validation package.

- APPENDIX

  - DATA COLLECTION SHEETS

  - INCUCELL 404-Operating Instructions

  - INCUCELL/INCUCELL V Laboratory Incubators with Natural/Forced Air Convention

  - Laboratory Drying Ovens and Incubators with innovated control automatics ECO line and EVO line

  - Test Equipment Certificates of Calibration

  - Equipment Certificate(s) of Calibration

  - Personnel Training

- REVISION HISTORY

| Revision Number | Description of Change | Changed By |
| --- | --- | --- |
| 01 | Original Issue | Alexander Waters |

  - APPENDIX 18.1: DATA COLLECTION SHEETS

| Pre-requisite Criteria |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| - Verify the pre-requisite criteria have been satisfied on this data collection sheet prior to testing. - Verification by the qualification technician must include the initials, signature and date verified. Additionally, a quality reviewer’s initials, signature, and date of quality review must be included. Verification by the qualification technician must occur concurrently with verification. Quality review may occur subsequently to verification. Hereafter, qualification technician and quality reviewer initials and date will suffice. - Record any additional information in the Comments section or indicate N/A if no comments are necessary. |  |  |  |  |  |  |
| Is the Process Validation Plan approved and released? |  |  |  | ☐YES ☐NO |  |  |
| Is this Installation Qualification Protocol (IQ) approved and released? |  |  |  | ☐YES ☐NO |  |  |
| Is objective evidence of training present for all persons participating in qualification activities? |  |  |  | ☐YES ☐NO |  |  |
| Evaluation of pre-requisite criteria: ☐PASSED ☐FAILED |  |  |  |  |  |  |
| Comments: | ☐N/A |  |  |  |
| Verification |  | Signature |  | Initials |  |  | Date |
| Verified By: |  |  |  |  |  |  |  |
| Quality Reviewed By: |  |  |  |  |  |  |  |

| Equipment Identification |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| - Verify and record the following equipment identification information. - Indicate if the information is acceptable. - Record any additional information in the Comments section or indicate N/A if no comments are necessary. |  |  |  |  |  |
| Parameter |  |  |  | Value |  | Accepted |  |
| Device Trade Name |  |  |  |  |  | ☐YES ☐NO |  |
| Device Model: |  |  |  |  |  | ☐YES ☐NO |  |
| Legal Manufacturer: |  |  |  |  |  | ☐YES ☐NO |  |
| Serial Number: |  |  |  |  |  | ☐YES ☐NO |  |
| Stimlabs Asset ID |  |  |  |  |  | ☐YES ☐NO |  |
| Evaluation of device identification information: ☐PASSED ☐FAILED |  |  |  |  |  |
| Comments: | ☐N/A |  |  |  |  |
| Verification |  | Initials |  | Date |  |
| Verified By: |  |  |  |  |  |
| Quality Reviewed By: |  |  |  |  |  |

| Equipment Installation |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| - Verify and record the following equipment installation information. - Indicate if the information is acceptable. - Record any additional information in the Comments section or indicate N/A if no comments are necessary. |  |  |  |  |  |  |
| Test Procedure Steps |  |  | Acceptance Criteria | Result |  | Accepted |
| Check the external dimensions of the equipment. |  |  | The actual external dimensions correspond to the external dimensions in the “Parameters of the apparatus” table of the operating instructions. | Width: Depth: Height: |  | ☐YES ☐NO |
| Check the clearance of the device in order to avoid overheating of surrounding walls/objects. |  |  | Distance from the device’s rear and side walls to surrounding walls/objects should be greater than or equal to 100 mm. | Distance from rear wall: Distance from left side wall: Distance from right side wall: |  | ☐YES ☐NO |
| Check the accessibility of the device from the front side for service and routine operations. |  |  | The device shall be accessible from the front side | ☐ The device is accessible from the front with the door being able to articulate for routine operations. ☐ Other (See Comments) |  | ☐YES ☐NO |
| Open the chamber door. Visually check the installation of the door seal in the chamber groove. |  |  | The door seal shall be properly installed in the chamber groove. The door seal is clean, in good condition and free of defects. | ☐ The door seal is free of defects and properly installed. ☐ Other (See Comments) |  | ☐YES ☐NO |
| Check the seating face. |  |  | Plane and horizontally balanced. | ☐ Plane and horizontally balanced. ☐Other (See Comments) |  | ☐YES ☐NO |
| Evaluation of equipment installation information: ☐PASSED ☐FAILED |  |  |  |  |  |  |
| Comments: | ☐N/A |  |  |  |  |  |
| Verification |  | Initials |  |  | Date |  |
| Verified By: |  |  |  |  |  |  |
| Quality Reviewed By: |  |  |  |  |  |  |

| Equipment Utilities |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| - Verify and record the following equipment utility information. - Indicate if the information is acceptable. - Record any additional information in the Comments section or indicate N/A if no comments are necessary. |  |  |  |  |  |  |  |
| Electric Parameters |  | Specified Value |  |  | Actual Value |  | Accepted |
| Current Type: |  |  | AC |  |  |  |  |  | ☐YES ☐NO |
| Phases: |  |  | ☐1 ☐3 |  |  |  |  |  | ☐YES ☐NO |
| Voltage ± fluctuation: |  |  | ☐115V ± 10% ☐230V ± 10% ☐400V ± 10% |  |  |  |  |  | ☐YES ☐NO |
| Frequency ± fluctuation |  |  | ☐50 Hz ± 5% ☐60 Hz ± 5% |  |  |  |  |  | ☐YES ☐NO |
| Current Protection [A]: |  |  | Should be sufficient to provide protection against the maximum device current stated in Laboratory Drying Ovens and Incubators with innovated control automatics ECO line and EVO line |  |  |  |  |  | ☐YES ☐NO |
| Test Equipment: |  |  |  |  | Name/Description: |  |  |  |  |
|  |  |  |  |  | Type/Manufacturer: |  |  |  |  |
|  |  |  |  |  | Serial Number: |  |  |  |  |
|  |  |  |  |  | Stimlabs Asset ID: |  |  |  |  |
|  |  |  |  |  | Calibration Due Date: |  |  |  |  |
| Evaluation of equipment utility information: ☐PASSED ☐FAILED |  |  |  |  |  |  |  |
| Comments: | ☐N/A |  |  |  |  |  |  |  |  |
| Verification |  |  | Initials |  |  | Date |  |
| Verified By: |  |  |  |  |  |  |  |
| Quality Reviewed By: |  |  |  |  |  |  |  |

| Environmental Requirements |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| - Verify and record the following environmental requirement information. - Indicate if the information is acceptable. - Record any additional information in the Comments section or indicate N/A if no comments are necessary. |  |  |  |  |  |  |  |
| Parameter |  |  | Acceptance Criteria |  | Result |  | Accepted |
| Temperature |  |  | 5 to 40 degrees Celsius |  |  |  | ☐YES ☐NO |
| Relative Humidity |  |  | Up to 80% at up to 31 degrees Celsius |  |  |  | ☐YES ☐NO |
| Test Equipment: |  |  |  | Name/Description: |  |  |  |  |
|  |  |  |  | Type/Manufacturer: |  |  |  |  |
|  |  |  |  | Serial Number: |  |  |  |  |
|  |  |  |  | Stimlabs Asset ID: |  |  |  |  |
|  |  |  |  | Calibration Due Date: |  |  |  |  |
| Evaluation of environmental requirement information: ☐PASSED ☐FAILED |  |  |  |  |  |  |  |
| Comments: | ☐N/A |  |  |  |  |  |  |  |
| Verification |  | Initials |  |  |  | Date |  |
| Verified By: |  |  |  |  |  |  |  |
| Quality Reviewed By: |  |  |  |  |  |  |  |

| Operational Requirements |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| - Verify and record the following equipment operational requirement information. - Indicate if the information is acceptable. - Record any additional information in the Comments section or indicate N/A if no comments are necessary. |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Test Procedure |  | Acceptance Criteria |  |  |  | Result |  |  |  |  | Accepted |  |  |
| Connect the device to the mains. Power on the device. Check the function of the display. |  |  |  | The display is functioning. The start screen displays. |  |  | ☐Functioning properly. ☐Other (See Comments |  |  |  |  |  | ☐YES ☐NO |  |  |
| Main Settings Verification |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| After powering the device on, ensure the main device settings are correctly set. |  |  |  | The device is set according to the “End User” actual requirements. |  |  | Device “SET:” |  |  |  |  |  | ☐YES ☐NO |  |  |
| After ensuring the device set point is satisfactory, allow, at a minimum, sixty (60) minutes for the chamber to reach temperature. |  |  |  |  | At least sixty (60) minutes must elapse for the chamber to reach temperature. |  |  | Start warm-up time: |  | Time elapsed: |  | ☐YES ☐NO |  |  |
|  |  |  |  |  |  |  |  | End warm-up time: |  |  |  |  |  |  |
| After allowing the appropriate time for the chamber to reach temperature, visually inspect the display and record the device “SET:” and the current chamber temperature on the display. |  |  |  | The device “SET:” and the device current chamber temperature on the display must agree within 10%. |  |  | DEVICE “SET:” Device chamber temperature: |  |  |  |  |  | ☐YES ☐NO |  |  |
| Evaluation of equipment operational requirement information: ☐PASSED ☐FAILED |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Comments: | ☐N/A |  |  |  |  |  |  |  |  |  |  |  |  |
| Verification |  |  | Initials |  |  |  |  |  | Date |  |  |  |  |  |  |
| Verified By: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Quality Reviewed By: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |

| Operational Requirements, continued. |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| - Verify and record the following equipment operational requirement information. - Indicate if the information is acceptable. - Record any additional information in the Comments section or indicate N/A if no comments are necessary. |  |  |  |  |  |  |
| Test Procedure |  |  | Acceptance Criteria | Result |  | Accepted |
| Place three (3) calibrated data loggers, which are programmed to record data for at least four (4) hours, inside the device chamber, placing them equidistant vertically from one another in the usable chamber volume and placing them centrally in their respective positions. Document test equipment identifiers for each chamber location. |  |  | Three (3) calibrated, programmed data loggers are vertically equidistant and centrally located in the usable chamber space and their identity is documented for each vertical position. | ☐Three calibrated, programmed data loggers are properly positioned in the chamber’s usable space: TOP: MIDDLE: BOTTOM: ☐Other (See Comments) |  | ☐YES ☐NO |
| Print the data logger reports and verify the elapsed time for each data logger is at least four (4) hours. |  |  | Each data logger recorded data for at least four (4) hours. | ☐Each data logger recorded data for at least four (4) hours. ☐Other (See Comments) |  | ☐YES ☐NO |
| Verify and document the report for each data logger presents an averaged recorded chamber temperature that is within 10% of the device “SET:” |  |  | All three (3) data loggers report an average temperature that is within 10% of the device “SET:” | TOP average temperature: MIDDLE average temperature: BOTTOM average temperature: ☐All three (3) data loggers’ average temperatures are within 10% of the device “SET:” ☐Other (See Comments) |  | ☐YES ☐NO |
| Evaluation of equipment operational requirement information: ☐PASSED ☐FAILED |  |  |  |  |  |  |
| Comments: | ☐N/A |  |  |  |  |  |
| Verification |  | Initials |  |  | Date |  |
| Verified By: |  |  |  |  |  |  |
| Quality Reviewed By: |  |  |  |  |  |  |

| Operational Requirements, continued. |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| - Verify and record the following equipment operational requirement information. - Record any additional information in the Comments section or indicate N/A if no comments are necessary. |  |  |  |  |  |
| Operational Requirements Test Equipment Information |  |  |  |  |  |
| Test Equipment: |  | Name/Description: |  |  |  |
|  |  | Type/Manufacturer: |  |  |  |
|  |  | Serial Number: |  |  |  |
|  |  | Stimlabs Asset ID: |  |  |  |
|  |  | Calibration Due Date: |  |  |  |
|  |  | Name/Description: |  |  |  |
|  |  | Type/Manufacturer: |  |  |  |
|  |  | Serial Number: |  |  |  |
|  |  | Stimlabs Asset ID: |  |  |  |
|  |  | Calibration Due Date: |  |  |  |
|  |  | Name/Description: |  |  |  |
|  |  | Type/Manufacturer: |  |  |  |
|  |  | Serial Number: |  |  |  |
|  |  | Stimlabs Asset ID: |  |  |  |
|  |  | Calibration Due Date: |  |  |  |
| Evaluation of equipment operational requirement information: ☐PASSED ☐FAILED |  |  |  |  |  |
| Comments: | ☐N/A |  |  |  |  |
| Verification |  |  | Initials |  | Date |
| Verified By: |  |  |  |  |  |
| Quality Reviewed By: |  |  |  |  |  |

| Data Collection Sheet Test Summary. |  |  |  |  |
| --- | --- | --- | --- | --- |
| - Indicate the results for each test performed. - Record any additional information in the Comments section or indicate N/A if no comments are necessary. |  |  |  |  |
| Name of Test |  |  | Result of Testing |  |
| Pre-requisite Criteria |  |  | ☐PASSED ☐FAILED ☐OTHER (See Comments) |  |
| Equipment Identification |  |  | ☐PASSED ☐FAILED ☐OTHER (See Comments) |  |
| Equipment Installation |  |  | ☐PASSED ☐FAILED ☐OTHER (See Comments) |  |
| Equipment Utilities |  |  |  | ☐PASSED ☐FAILED ☐OTHER (See Comments) |  |  |
| Environmental Requirements |  |  |  | ☐PASSED ☐FAILED ☐OTHER (See Comments) |  |  |
| Operational Requirements |  |  |  | ☐PASSED ☐FAILED ☐OTHER (See Comments) |  |  |
| Evaluation of all testing conducted as part of this installation qualification: ☐PASSED-The outputs of this installation qualification protocol have satisfied all acceptance criteria. ☐FAILED (See Comments) |  |  |  |  |
| Comments: | ☐N/A |  |  |  |  |  |
| Verification |  | Initials |  | Date |
| Verified By: |  |  |  |  |
| Quality Reviewed By: |  |  |  |  |