- PURPOSE

  - The purpose of this installation qualification is to define the qualification plan for the American BioTech Supply Refrigerator.  The refrigerator will be utilized for storing tissue and processing supplies that are required to be maintained in a refrigerated environment prior to processing.

  - This validation is a deliverable for PCR21-039 New Refrigerator for Tissue Storage.

  - The purpose of this revision to the original protocol was to remove a chart recorder as the data collection source for the refrigerator and use a temperature data logger in place.  A chart recorder will still be used for monitoring temperature and humidity.

- SCOPE

  - This validation applies to the American BioTech Supply Refrigerators for storing tissue and processing supplies that are required to be maintained in a refrigerated environment prior to processing.  This validation shall serve as the initial qualification of this equipment.

  - This equipment is located in Donor Services at the StimLabs Roswell, GA facility.

- BACKGROUND

  - StimLabs is a tissue bank involved in the processing of human derived cells and tissues into products for human use.  StimLabs receives tissue and other supplies that are required to be stored in a refrigerated environment prior to processing.  During the waiting period, the tissue or supplies must be held in a refrigerated environment to preserve their characteristics.

  - To maintain a refrigerated environment, a refrigerator must maintain its temperature between 1°C and 10°C.  The set point drift for the temperature controller must not exceed (-1°C /+ 4°C) in a given time span.  To ensure the refrigerator is fit for the intended purpose, StimLabs will perform an installation qualification and continue to monitor the performance of the refrigerator while it is being used to store human tissue and supplies through an operational qualification.

  - The American BioTech Supply  Refrigerator will be calibrated prior to the start of the validation activities.  The set point may be adjusted prior to the start of the validation if desired.

- REFERENCES

  - Food and Drug Administration 21 CFR Part 1271

  - American Association for Tissue Banks Standards for Tissue Banking

  - American BioTech Supply  Refrigerator Equipment Manual

  - FDA Guidance for Industry – Process Validation: General Principles and Practices

  - GHTF/SG3/N99-10:2004 Quality Management Systems – Process Validation Guidance

  - VAP21-023-00 New Refrigerators for Tissue Storage (PVP)

  - VAP21-023-01 New Refrigerators for Tissue Storage (IQ)

  - VAP21-023-02 New Refrigerators for Tissue Storage (OQ)

  - PCR21-039 New Refrigerators for Tissue Storage

- DEFINITIONS

  - Acceptance Criteria:  Pre-determined requirements upon which pass/fail decisions are based.  Acceptance criteria are specific and measurable.

  - Critical to Quality:  Attribute of a part, assembly, sub-assembly, product, or process that has direct and significant impact on its actual or perceived quality to the customer or end user.

  - Deviation:  The action of departing from an established method, procedure, or acceptance criteria.

  - Installation Qualification (IQ):  Establishing by objective evidence that all key aspects of the process equipment and ancillary system installation adhere to the manufacturer’s approved specification(s).

  - Process Validation:  A mechanism to ensure that a process whose output is not fully verifiable by inspection and test is capable of consistently (repeatedly and reliably) producing a result or product that meets pre-determined requirements.  Process validation proves that a process is under control and is used to determine the process variables, acceptable limits for those variables, and to establish appropriate in-process controls.

  - Protocol:  A pre-approved and prospective document that states the requirements and procedures for conducting a validation.

  - Validation/Study Leader: The individual who has the responsibility and authority over the operation of the equipment or process being validated.

  - Verification:  Confirmation by examination with a provision of documented evidence that a specific requirement has been fulfilled.

- RESPONSIBILITIES

| Function | Name | Author Protocol & Report | Review Protocol & Report | Execute Protocol | Update WI, SOP, Forms |
| --- | --- | --- | --- | --- | --- |
| Validation/Study Leader | Rob Frye | X |  |  |  |
| MwER/MR/HOQ/TBD | Melissa O’Connor |  | X |  |  |
| Head of Quality | Stacy Harmon |  | X |  |  |
| Operations Management | Michael Baird |  | X |  | X |
| Leads/Operators/Engineers | Personnel |  |  | X |  |

- SAFETY CONSIDERATIONS

  - Do not store volatile or flammable substances in this unit except in sealed containers.

  - Never use an acid-based cleaning solution

  - Before proceeding with maintenance or checking of the unit, set the power switch to off and disconnect the power

- EQUIPMENT

  - Manufacturing Equipment Description

    - Manufacturer: American BioTech Supply

    - Serial Number(s): will be identified on the data collection sets for the IQ

    - Model Number: MPR-1412R-PA

    - Location:  Support Lab

  - Manufacturing Equipment Parameters

    - Set point recommended is 3℃ for a specification range of  (-1°C/+4℃) to maintain the operation temperature between 3°C and 7°C, you would set “SET” = 3°C, and “Hy” = 4°C.  See equipment manual.

    - Electronic Temperature Control

    - 115V/60Hz/Single Phase

  - Test Equipment

    - Calibrated equipment shall be used for the execution of this protocol.  Equipment ID/serial numbers and calibration information will be documented on the data collection sheet.

    - Voltage Meter

    - Tape Measure

    - Data logger (temperature)

- PROCEDURE

  - Verify Equipment Identification

    - Confirm the model identified in this protocol.

    - Document serial number on the data collection sheet.

  - Verify Environmental Requirement

    - Temperature – Measure the temperature in the area where the unit is installed

    - Relative Humidity – Measure the relative humidity in the area where the unit is installed.

    - Positioning – Verify that the unit is not exposed to direct sunlight or other heat sources.

  - Verify Critical Dimensions

    - Verify the dimensional specification of the unit.

  - Verify Proper Installation

    - Leveling – make sure the unit is leveled properly so the doors operate correctly.

    - Stabilize – ensure the casters are properly adjusted to keep the cabinet level to make sure the unit will not rock, lean, or be unstable.

    - Glycol Filled Bottle – Verify the bottle used to hold the temperature probe contains a mix of 10% glycol and 90% water.

    - Door Gaskets – make sure the door gaskets make a proper seal to the cabinet.

    - Shelving – verify that shelves are securely attached.

    - Lamps – verify the lamps are properly installed, illuminated, and fully engaged in the lamp holders.

    - Verify wall space per manufacturer’s recommendations – verify that there is at least 2 to 3 inches of space between the cabinet and any adjacent wall or fixture.

  - Verify Utility Connections

    - Measure the voltage of the refrigerator power source.

    - Measure max amps being drawn by unit.

  - Verify Equipment Controls

    - Temperature – With the set-point set to 3℃, place one data logger with the probe positioned in the middle of the chamber to check the temperature of the chamber.  Allow the temperature to equilibrate for at least one hour after placing the probe.  Verify a set-point 3℃ (-1°C /+4℃) can be measured on the data logger.  Verify the temperature measured by the data logger agrees with the set point and display temperature on the temperature controller within (-1°C /+4℃).

    - Verify alarm with the door open – After reaching set point, open the door of the refrigerator and wait at least 5 minutes.  Verify the refrigerator displays visual and audible alarms when the temperature exceeds 9℃.

  - Verify Data Logger

    - Verify that data logger is programmed.

    - Verify data logger is placed in the center of the chamber.

    - After unit is turned on, verify the data logger is recording the temperature.  Monitor for at least one hour by recording the start time and temperature measurements the data logger recorded.  The temperature measured by the data logger should be in agreement with the temperature display reading within (-1°C /+4℃).

    - Remove data logger after the test, download data, initial and date print out with VAP21-023-01 and include a copy in the IQ final report.

- LOT SIZE & SAMPLING PLAN

  - None

- ACCEPTANCE CRITERIA

  - Equipment Verification

    - Must match Manufacturing Equipment Description section above.

  - Environment Verification

    - Temperature must be between 18℃ to 29℃ (65°F to 85°F)

    - Relative Humidity must be ≤ 80%

    - Verify unit is not installed in direct sunlight or other heat sources.

  - Critical Dimensions

    - The dimensions of the chamber should match the equipment specification within the tolerance described in the manufacturer’s specification.

  - Installation

    - Leveling

      - Level from front to back and side to side.

    - Stabilization

      - The unit should not rock, lean, or be otherwise unstable.

    - Bottle

      - Verify bottle used to hold the thermocouple probes is filled with a 90/10 mixture of water and glycol.

    - Door Gasket

      - Properly sealed

    - Lamps

      - Installed, illuminated, and properly locked in place

    - Wall Space

      - 2 – 3 inches of spacing between the unit and adjacent walls.

  - Utility Connections

    - Power

      - Verify power is 115V ±10%

    - Hertz

      - Verify Hertz is 60Hz ±10%

  - Equipment Controls

    - Set-point

      - Factory Setting of 3℃ with (-1°C /+4℃) must be measured by the data logger (average over at least one hour) in the empty chamber.

      - The data logger (average over at least one hour), set-point, and display temperature must agree within (-1°C /+4℃).

    - Alarm

      - Visible and audible alarm should sound after the door has been left open and the temperature reaches 9℃.  The test fails if the alarm does not sound when the display temperature exceeds 9℃.

  - Data Logger

    - Temperature recorded should match the temperature controller over a time span of at least one hour at 3°C (+4℃/-0°C).

- MATERIAL DISPOSITION

  - None

- DOCUMENTATION REQUIREMENTS

  - Results shall be summarized in an IQ Report.

  - Various Control Plans, including the Preventative Maintenance (PM) Schedule, safety/PPE requirements, environmental requirements, and re-validation criteria shall be documented in the final report.

  - Raw data sheets shall be included in the appendix of the IQ report and/or a reference to the location of the raw data shall be included in the IQ report.

  - Calibration records (if applicable) shall be attached to the IQ report (in the Appendix).

  - Signed IQ Protocol Training forms will be attached to the IQ report (in the Appendix).

- DEVIATIONS

  - Deviations to the IQ will be pre-approved using FRM020-012 Protocol Deviation Form and explained in the IQ report.

  - Copies of signed Protocol Deviation Forms will be attached to the IQ report (in the Appendix).

- PRE-REQUISITE STEPS

  - Approved Process Validation Plan (PVP) or justification for omission.

  - The IQ Protocol must be signed before validation activities may begin.

  - Completed Operator Training.

- APPENDIX

  - Blank data collection sheets, print additional copy for the other refrigerator

16.1 – DATA COLLECTION SHEETS

STOP!  Read Instructions.

Instructions:

- Complete one data collection sheet for each American BioTech Supply  Refrigerator.

- Verify the Pre-requisite criteria have been met on this data collection sheet prior to testing.  Note that the Pre-requisite criteria only need to be verified once for the validation.

- For attribute data: Circle “Pass” or “Fail” or write the words “Pass” or “Fail”

- For variable data: Record actual results in the units specified or add units if the units are not specified.

- Verification must include signature of the individual performing the step and the date performed.  Verification must be completed concurrently with performance.

- Follow Good Documentation Practice (GDP).

Pre-requisite Criteria (must be complete and verified prior to start of IQ activities):

| Pre-requisite | Complete (Circle One) | Verification Signature/Date |
| --- | --- | --- |
| Process Validation Plan is approved. | Yes No |  |
| IQ Protocol is signed and approved. | Yes No |  |
| Operator Training is complete and documented. | Yes No |  |

Pre-Requisites Verified By (signature & date): _____________________________________________________________

Quality Reviewed By (signature & date): ______________________________________________________________

Section A – Equipment Identification

| Equipment Name | Model Number | Manufacturer | Serial Number | Stimlabs ID Number | Verification Signature/Date |
| --- | --- | --- | --- | --- | --- |

| Specification or Acceptance Criteria | Result |  |  | Verification Signature/Date |  |
| --- | --- | --- | --- | --- | --- |
| External Dimensional Specifications ±1.0 ʺ | Depth: Actual: ____________ Manufacture: 34.75 ʺ | Pass / Fail |  |  |
|  | Height: Actual: ____________ Manufacture: 79.75 ʺ | Pass / Fail |  |  |
|  | Width: Actual: ____________ Manufacture: 54ʺ | Pass / Fail |  |  |
| Equipment is marked with StimLabs Equipment ID | Pass / Fail |  |  |  |  |
| Equipment is marked with a calibration sticker and is within current calibration | Pass / Fail |  |  |  |  |
| A logbook has been created for the equipment and is available for use | Pass / Fail |  |  |  |  |

Section A Completed By (signature & date): ________________________________________________________________

Quality Reviewed By (signature & date): ______________________________________________________________

Section B - Installation

| Specification or Acceptance Criteria | Result |  | Verification Signature/Date |
| --- | --- | --- | --- |
| level from front to back and side to side | Pass / Fail |  |  |
| Equipment does not rock, lean, or otherwise unstable. | Pass / Fail |  |  |
| Bottle contains 90/10 mixture of water to glycol. | Pass / Fail |  |  |
| Door gasket properly sealed. | Pass / Fail |  |  |
| Lamps installed, illuminated, and properly locked in place. | Pass / Fail |  |  |
| Wall space has at a minimum 2-3 inches between unit and adjacent walls | Left Side: | Pass / Fail |  |
|  | Right Side: | Pass / Fail |  |
|  | Back: | Pass / Fail |  |

Section B Completed By (signature & date): ________________________________________________________________

Quality Reviewed By (signature & date): ______________________________________________________________

Section C – Utility Connection

| StimLabs ID Number | Serial Number | Utility | Requirement | Result | Verification Signature/Date |
| --- | --- | --- | --- | --- | --- |
|  |  | Voltage | 115V ± 10% Actual: | Pass Fail |  |
|  |  | Hertz | 60Hz ±10% Actual: | Pass Fail |  |

Record Test/Measurement Equipment Used and Calibration Due Date (attach calibration certificate to final report): __________________________________________________________________________________________________________________

__________________________________________________________________________________________________________________

Section C Completed By (Signature & date): ________________________________________________________________

Quality Reviewed By (signature & date): ______________________________________________________________

Section D - Environmental Requirements

| StimLabs ID Number | Serial Number | Chart Recorder ID Number | Result |  | Verification Signature/Date |
| --- | --- | --- | --- | --- | --- |
|  |  |  | Humidity | Temperature |  |
|  |  |  | Range: 80% Actual: Pass / Fail | Range: 5°C to 40°C Actual: Pass / Fail |  |
|  |  | Cal Due Date: |  |  |  |

Record Test/Measurement Equipment Used and Calibration Due Date (attach calibration certificate to final report): __________________________________________________________________________________________________________________

__________________________________________________________________________________________________________________

__________________________________________________________________________________________________________________

| Specification and Acceptance Criteria | Result | Verification Signature/Date |
| --- | --- | --- |
| The equipment is located away from any source to of direct heat. | Pass / Fail |  |

Section D Completed By (signature & date): ________________________________________________________________

Quality Reviewed By (signature & date): ______________________________________________________________

Section E – Operational Requirements

| Specification and Acceptance Criteria | Reading/Results | Result | Verification Signature/Date |
| --- | --- | --- | --- |
| Set Point Verification – Unit turned on and set to 3°C. Data logger placed on shelf near the center of chamber. |  |  |  |
| Visual and audible alarm initiates after the door has been left open and the temperature reaches 9°C. | Audible alarm temp: ____________________°C Visual alarm temp: ____________________°C | Pass / Fail Pass / Fail |  |
| 3℃ (-1°C /+4℃) must be measured by the data logger (average over at least one hour) in the empty chamber. | Enter Set-Point _____________ ℃ | N/A |  |
|  | Time data logger placed: _____________ | N/A |  |
|  | Time data logger checked (at least one hour after time place in): ______________ | N/A |  |
|  | Set-Point reading after 1 hour _____________ ℃ | Pass / Fail |  |
| The data logger (average over at least one hour), set-point, and display temperature must agree within 3℃ (-1°C /+4℃) | Set-Point reading after 1 hour _____________ ℃ | Pass / Fail |  |
|  | Display reading after 1 hour _____________ ℃ | Pass / Fail |  |
| Temperature recorded should match the data logger temperature accurately over a time span of at least upon one hour 3℃ (-1°C /+4℃) | Results documented on data logger report | Pass / Fail |  |

Section E Completed By (signature & date): ________________________________________________________________

Quality Reviewed By (signature & date): ___________________________________________________________________