- PURPOSE

  - The purpose of this document is to outline the installation qualification protocol for PHCbi MPR-1412R Series Pharmaceutical Refrigerator(s) according to the manufacturer’s (PHCbi) and the customer’s (Stimlabs, LLC or Stimlabs) specifications, procedures, and requirements.

- SCOPE

  - This validation applies to the American BioTech Supply model ABT-HC-LS-49 Refrigerator that will be utilized for storing incoming supplies required to conduct operations activities at StimLabs. These incoming supplies must be stored in a refrigerated environment prior to utilization. This validation shall serve as the initial qualification of this equipment.

- BACKGROUND

  - StimLabs is a tissue bank involved in the processing of human derived cells and tissues into products for human use. StimLabs receives supplies that require storage in a refrigerated environment prior to utilization during operations activities. These incoming supplies must be stored in a refrigerated environment to preserve their characteristics until they are utilized.

  - To maintain a refrigerated environment, a refrigerator must maintain its temperature between 1 degree Celsius and 10 degrees Celsius. To ensure the refrigerator is fit for the intended purpose, StimLabs will perform an installation qualification and continue to monitor the performance of the refrigerator while it is being used to store supplies that require refrigeration through an operational qualification.

  - The American BioTech Supply model ABT-HC-LS-49 Refrigerator will be calibrated prior to operational requirement testing. The set point may be adjusted prior to operational requirement testing if necessary.

- REFERENCES

  - Food and Drug Administration 21 CFR Part 1271

  - American Association for Tissue Banks Standards for Tissue Banking

  - FDA Guidance for Industry–Process Validation: General Principles and Practices

  - GHTF/SG3/N99-10:2004 Quality Management Systems–Process Validation Guidance

  - PHCbi MPR-1412 Series Operating Instructions

  - VAP22-008-00 PHCbi MPR-1412R Series Process Validation Plan

  - VAP22-008-01 PHCbi MPR-1412R Series Installation Qualification Protocol

  - VAR22-008-01 PHCbi MPR-1412R Series Installation Qualification Report

  - VAR22-008-02 PHCbi MPR-1412R Series Operational Qualification Report

- DEFINITIONS

  - Acceptance Criteria: Predetermined requirements upon which pass/fail decisions are based. Acceptance criteria are specific and measurable.

  - Critical to Quality: Attribute of a part, assembly, sub-assembly, product, or process that has direct and significant impact on its actual or perceived quality to the customer or end user.

  - Controlled Environment: Defined zone in which sources of contamination are controlled by specified means.

  - Deviation: The action of departing from an established method, procedure, or acceptance criteria.

  - Environmental Control: Activities performed to control the environment for the purpose of minimizing the potential for contamination and cross-contamination of tissue.

  - HCT/P’s: Human Cells, Tissues, and Cellular and Tissue-Based Products.

  - In-process Controls: Any tests, samples, evaluations, monitoring, or measurements performed during processing or preservation that are designed to evaluate the processing or preservation procedure of the tissue subjected to processing or preservation for conformance to specification(s).

  - Installation Qualification (IQ): Establishing by objective evidence that all key aspects of the process equipment and ancillary system installation adhere to the manufacturer’s approved specification(s).

  - Management with Executive Responsibility (a.k.a. Management with Executive Responsibility/Management Representative/Head of Quality/Tissue Bank Director): Those senior employees of an establishment who have the authority to establish or make changes to the establishment’s quality policy and quality system.

  - Objective Evidence: Information, records, or statements of fact which are based on observation, measurement, or tests that can be verified.

  - Operational Qualification (OQ): Establishing by objective evidence the process control limits and action levels which result in a product that meets predetermined requirements and are within the operating limits required by the process.

  - Performance/Product Qualification (PQ): Establishing by objective evidence that the process, under anticipated conditions, consistently (repeatedly and reliably) produces a product which meets predetermined requirements.

  - Preventative Maintenance: Maintenance that is regularly performed on a piece of equipment to lessen the likelihood of it failing. Preventative maintenance is performed while the equipment is still working, so that it does not break down unexpectedly.

  - Processing Materials: Materials that are used in the manufacture or inspection of product and may come into direct contact with the product.

  - Process Validation: Establishing by objective evidence that the output of a process consistently meets predetermined specification.

  - Protocol: A pre-approved and prospective document that states the requirements and procedures for conducting a validation.

  - Validation/Study Leader: The individual who has the responsibility and authority over the operation of the equipment or process being validated.

  - Verification: Confirmation by examination with a provision of documented evidence that a specific requirement has been fulfilled.

- RESPONSIBILITIES

| Function | Name | Title | Author Protocols and Reports | Execute Stimlabs Protocols | Review and Approve Stimlabs Protocols and Reports |
| --- | --- | --- | --- | --- | --- |
| Validation/Study Leader | Stimlabs Personnel | Quality Engineer or Equivalent | X | X | X |
| Management with Executive Responsibility | Melissa O’ Connor | Chief Quality & Regulatory Officer |  |  | X |
| Quality Management | Stimlabs Personnel | Quality Engineering Manager or Equivalent |  |  | X |
| Operations Management | Stimlabs Personnel | Operations Management or Equivalent |  |  | X |

- ENVIRONMENTAL REQUIREMENTS

  - Temperature: Temperature must be between 5 degrees Celsius and 40 degrees Celsius.

  - Relative Humidity: Relative humidity must be ≤ 80%.

  - Positioning: The unit may not be installed in direct sunlight or near other sources of heat.

- PROCEDURE

  - Manufacturing Equipment Description

    - Manufacturer: PHC Corporation

    - Serial Number: Serial number(s) were verified during execution of IQ protocol, VAP22-008-01, and will be verified and documented during execution of this OQ protocol.

    - Model Number: MPR-1412R Series

  - Raw Materials

    - Data logger(s)

    - 9” x 12” bags

    - IAS16-006 containers

    - Water

  - Critical to Quality (CTQ) Attributes

    - Temperature accuracy and precision: It is critical that the equipment is able to exhibit accurate temperatures relative to the set point. Accuracy describes how close the measured temperature of the internal chamber is to the true value. It is also critical that the equipment is able to exhibit precise temperatures. Precision describes how close measured temperature values are, relative to one another. Temperature accuracy and precision of the internal chamber of the equipment demonstrate the process equipment is capable of consistently operating within the stated limits and tolerances or defined functional requirements of its intended use. Satisfactory temperature accuracy and precision should be demonstrable in both an empty test chamber and a loaded test chamber.

  - Process Variables

    - Following the IQ, no adjustments should need to be made to achieve the desired temperature set point. If adjustment is needed, the set point will be set to 3 degrees Celsius.

    - OQ Run 1–Empty Chamber

      - Verify that the data logger is programmed.

      - Verify the data logger is placed in the center of the chamber.

      - After unit is turned on, verify the data logger is recording the temperature. Monitor for a period of time that is ≥ twenty-four (24) hours by recording the start time and temperature measurements the data logger records. The average temperature measured by the data logger over a timespan that is ≥ twenty-four (24) hours should agree with the temperature display reading within (-1 degree Celsius /+4 degrees Celsius). For example, a setpoint of 3 degrees Celsius will have a specification range of (2 degrees Celsius to 7 degrees Celsius).

      - Remove the data logger after the test, download data, initial and date the data report with “VAP21-029-02 Run 1” included as a descriptor and include a copy in the OQ final report. See WI015-026 Operation of Cryopak iMini Data Logger

    - OQ Run 2–Full Chamber

      - Verify three (3) data loggers are programmed.

      - Start the data loggers after ensuring the unit is on and the set point is set to 3 degrees Celsius.

      - Load the chamber with representative materials that include appropriately partitioned aliquots of water that additively result in a total of 200L of water volume occupying the internal chamber of the equipment, and that are regularly distributed throughout the internal chamber of the equipment to simulate an internal chamber under maximal load. Representative materials may be partitioned into 9” x 12” bags or IAS16-006 containers or a normal and regularly distributable combination of 9” x 12” bags and IAS16-006 containers.

        - Filling 9” x 12” bags: Use a 1-liter beaker to achieve a total volume ≊2L of water.

          - Each 9” x 12” bag should be placed inside another 9” x 12” bag and sealed completely to prevent leaking/spilling.

        - Filling IAS16-006 containers: Use a 1-liter beaker to achieve a total volume ≊1L of water.

          - Each IAS16-006 container should have the lid securely affixed to prevent leaking/spilling.

        - If using a combination of 9” x 12” bags and IAS16-006 containers, exercise caution in ensuring the combination of contained materials results in a total volume ≊200L of water that is regularly distributed throughout the internal chamber of the equipment.

        - Open the door(s) of the equipment and distribute the representative materials regularly on the shelves to simulate typical loading conditions. When loading the samples, ensure there is a clearance ≥ 2-3” from the back wall and the ceiling of the chamber to allow for proper airflow and circulation.

      - Insert the 3 data logger thermocouples probes inside the loaded chamber. Measurement points will be taken from the bottom, middle, and top of the chamber at the center of each shelf.

      - Close the door and verify the set point is set to 3 degrees Celsius.

      - Begin recording temperature data for a period no less than 24 hours.

      - At the end of the run, remove the three data loggers, download, and review the data. Discard the representative samples after the average temperature data has been retrieved and deemed acceptable according to the specifications outlined in the acceptance criteria.

      - Initial and date the data reports and ensure “VAP21-029-02 Run 2” and the section of the chamber the data was recorded in are present in the data report descriptions.

  - Test Equipment

    - Data logger(s)

    - 9” x 12” bags

    - IAS16-006 containers

    - 1L beaker

    - 200L water

- LOT SIZE & SAMPLING PLAN

  - Run 1

    - The refrigerator shall be monitored for at least 24 hours at a set point of 3 degrees Celsius with an empty chamber to ensure it can operate as intended with no other variables (addition of samples) impacting the study.

  - Run 2

    - The refrigerator shall be monitored for 24 hours at a setpoint of 3 degrees Celsius with the chamber loaded with representative materials that total ≊200L of volume. 200L of volume is the volume which most closely represents a typical fully loaded chamber containing 100 donor samples, which is the maximum typical operating capacity. A total representative material sample size of ≊200L of water was chosen to simulate maximum capacity with samples that closely resemble the thermal load of real samples that will be stored during normal operations.

- ACCEPTANCE CRITERIA

  - Run 1 Temperature Uniformity

    - The temperature differences of the average temperatures recorded by the data logger thermocouple probes must not vary more than 2 degrees Celsius from each other throughout the run after the set-point is reached.

    - After the set point is achieved, the average temperatures recorded by the data logger thermocouple probes over 24 hours must not exceed 7 degrees Celsius nor fall below 2 degrees Celsius.

  - Run 2 Temperature Uniformity

    - The temperature differences of the average temperatures recorded by the data logger thermocouple probes must not vary more than 2 degrees Celsius from each other throughout the run after the set-point is reached.

    - After the set point is achieved, the average temperatures recorded by the data logger thermocouple probes over 24 hours must not exceed 7 degrees Celsius nor fall below 2 degrees Celsius.

- MATERIAL DISPOSITION

  - The representative materials shall be discarded at the conclusion of the study.

- DOCUMENTATION REQUIREMENTS

  - Results shall be summarized in an OQ report.

  - Various control plans, including validated process settings and the re-validation criteria shall be documented in the final report.

  - Raw data sheets shall be included in the appendix of the OQ report and/or a reference to the location of the raw data shall be included in the OQ report.

  - Calibration records (where applicable) shall be attached to the OQ report (in the appendix).

  - Signed OQ protocol training forms will be attached to the OQ report (in the appendix).

- DEVIATIONS

  - Deviations to the OQ will be reviewed for approval using FRM020-012 Protocol Deviation Form and described in the OQ report.

  - Copies of signed protocol deviation forms will be attached to the OQ report (in the appendix).

- PRE-REQUISITE STEPS

  - Approved Process Validation Plan (PVP) and IQ Report, or justification for omission.

  - Calibration certificates for equipment and measurement devices (gauges)

  - Approval and release of raw materials

  - Verification of environmental requirements

  - Completed Operator Training

  - Draft work instructions

  - The OQ Protocol must be signed before validation activities may begin.

- APPENDIX

  - Blank data collection sheets

APPENDIX 15.1–DATA COLLECTION SHEETS

STOP! Read Instructions.

Instructions:

- Complete one data collection sheet per run for each piece of equipment.

- Verify the pre-requisite criteria have been met on this data collection sheet prior to testing. Note that the pre-requisite criteria only need to be verified once for the validation.

- For attribute data: Circle “Pass” or “Fail” or write the words “Pass” or “Fail”

- For variable data: Record actual results in the units specified or add units if the units are not specified.

- Verification must include signature of the individual performing the step and the date performed. Verification must be completed concurrently with performance.

- Follow Good Documentation Practice (GDP).

Pre-requisite Criteria (must be complete and verified prior to start of OQ activities):

| Pre-requisite | Complete (Circle One) | Verification Signature/Date |
| --- | --- | --- |
| Process Validation Plan is approved. | Yes No |  |
| IQ Protocol is completed | Yes No |  |
| OQ Protocol approved | Yes No |  |
| Unit is Calibrated | Yes No |  |
| Operator Training is complete and documented. | Yes No |  |

Pre-requisites verified by (signature & date): _____________________________________________________________

Quality reviewed by (signature & date): ______________________________________________________________

Section A–OQ Run #1 Preparation

| Equipment Name | Model Number | Manufacturer | Serial Number |  | Stimlabs ID Number |  | Verification Signature/Date |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Specification and Acceptance Criteria |  | Reading/Results |  | Result |  | Verification Signature/Date |  |
| Thermocouple Setup with Data Loggers |  |  |  |  |  |  |  |
| Thermocouples are loaded on the bottom, middle, and top of chamber. |  | Data Logger ID Top_________________ Middle _____________ Bottom _____________ |  | Pass / Fail |  |  |  |

Section A completed by (signature & date): ________________________________________________________________

Quality reviewed by (signature & date): ______________________________________________________________

Section B-OQ Run #1

| Specification or Acceptance Criteria | Reading/Results | Result | Verification Signature/Date |
| --- | --- | --- | --- |
| Equipment should be on with the internal chamber empty. Close the unit door after thermocouples are loaded. Ensure set point is 3 degrees Celsius. Record temperature data for a period not less than 24 hours. Attach copies of data logger reports to the final OQ report. |  |  |  |
| Enter date and time empty chamber test was initiated. |  | N/A |  |
| Enter set point display (3 degrees Celsius). |  | Pass / Fail |  |
| Enter date and time set point display was checked (at least 24 hours after initiation). |  | N/A |  |
| Enter set point display after at least 24 hours has elapsed (3 degrees Celsius). |  | Pass / Fail |  |
| Equipment digital display reading after 24 hours. (2 degrees Celsius to 7 degrees Celsius) |  | Pass / Fail |  |
| Average temperature recorded over at least 24 hours for bottom data logger (2 degrees Celsius to 7 degrees Celsius). |  | Pass / Fail |  |
| Average temperature recorded over at least 24 hours for middle data logger (2 degrees Celsius to 7 degrees Celsius) |  | Pass / Fail |  |
| Average temperature recorded over at least 24 hours for top data logger (2 degrees Celsius to 7 degrees Celsius) |  | Pass / Fail |  |

Section B completed by (signature & date): ________________________________________________________________

Quality reviewed by (signature & date): ______________________________________________________________

Section C–Run #2 Preparation

| Equipment Name | Model Number | Manufacturer | Serial Number | StimLabs ID Number | Verification Signature/Date |
| --- | --- | --- | --- | --- | --- |

| Specification and Acceptance Criteria | Reading/Results | Result |  | Verification Signature/Date |
| --- | --- | --- | --- | --- |
| Data Logger Setup |  |  |  |  |
| Data Loggers Programmed | N/A | Pass / Fail |  |  |
| Dummy Load Assembly and Setup |  |  |  |  |
| 9”x12” bags filled with ≊2L of water. | Qty of 9” x 12” bags filled with ≊2L of water: ______________ | N/A |  |  |
|  | Verify each 9” x 12” bag is sealed and securely placed inside another 9” x 12” bag. | Pass / Fail |  |  |
|  | Total volume of water present in 9” x 12” bags: ______________ | N/A |  |  |
| IAS16-006 containers filled with ≊1L of water. | Qty of IAS16-006 containers filled with ≊1L of water: ______________ | N/A |  |  |
|  | Verify each IAS16-006 container has the lid securely affixed. | Pass / Fail |  |  |
|  | Total volume of water present in IAS16-006 containers: ______________ | N/A |  |  |
| Total Water Volume of Representative Materials (≊200L). | ______________ | Pass / Fail |  |  |
| Thermocouple Setup |  |  |  |  |
| Thermocouples are loaded on the bottom, middle, and top of chamber. | Data Logger ID Top_________________ Middle _____________ Bottom _____________ | Pass / Fail |  |  |

Section C completed by (signature & date): ________________________________________________________________

Quality reviewed by (signature & date): ______________________________________________________________

Section D-OQ Run #2

| Specification or Acceptance Criteria | Reading/Results | Result | Verification Signature/Date |
| --- | --- | --- | --- |
| Equipment should be on with the internal chamber LOADED. Thermocouples should be placed in the top, middle, and bottom of the chamber. Close the unit door after the thermocouples are loaded. Ensure set point is 3 degrees Celsius. Record temperature data for a period not less than 24 hours. Attach copies of data logger reports to the final OQ report. |  |  |  |
| Enter date and time loaded test was initiated. |  | N/A |  |
| Enter set point display (3 degrees Celsius) |  | Pass / Fail |  |
| Enter date and time set point was checked (at least 24 hours after initiation). |  | N/A |  |
| Enter set point display after at least 24 hours has elapsed (3 degrees Celsius). |  | Pass / Fail |  |
| Equipment digital display reading after 24 hours. (2 degrees Celsius to 7 degrees Celsius) |  | Pass / Fail |  |
| Average temperature recorded over at least 24 hours for bottom data logger (2 degrees Celsius to 7 degrees Celsius). |  | Pass / Fail |  |
| Average temperature recorded over at least 24 hours for middle data logger (2 degrees Celsius to 7 degrees Celsius) |  | Pass / Fail |  |
| Average temperature recorded over at least 24 hours for top data logger (2 degrees Celsius to 7 degrees Celsius) |  | Pass / Fail |  |

Section D completed by (signature & date): ________________________________________________________________

Quality reviewed by (signature & date): ______________________________________________________________