Approvals

This document has been reviewed and approved by the following individuals.  By approving this document, each individual acknowledges that it complies with the requirements and spirit of the Equipment, Process, and Test Method Validation Procedure (QP 09-002), and all associated procedures.  Modify as needed.

| Denean O’Brien Protocol Author |  | Date |  | Marissa Gardner Process Owner |  | Date |
| --- | --- | --- | --- | --- | --- | --- |
| Hetal Patel Quality Assurance |  | Date |  | Kenneth Broadley Executive Vice President |  | Date |
| Rob Frye Validation Reviewer |  | Date |  |  |  |

- Purpose

  - To demonstrate that the AN333.11 EOGas™ sterilizers and ancillary items at Chemence Medical, Inc. operates within required process parameters necessary for sterilization for new product packaging and sterilization configuration.

- Scope

  - The scope of this Operational Qualification includes verifying that the sterilizer operates within parameters for temperature, humidity, time and load configuration for Class II medical devices packaged within blister packaging with Tyvek lidding.

- Definitions

  - Acceptance Criteria – A set of pre-determined requirements upon which pass/fail decisions are based. Acceptance criteria must be specific and measurable.

  - Operational Qualification (OQ) – Establishes objective evidence process control limits and action levels which result in product that meets all predetermined requirements.

  - Process Validation – Establishing objective evidence that a process consistently produces a result which meets acceptance criteria.

- References

  - CHM0222815IQ & CHM0222815OQ – Installation & Operation Qualification of the AN333 EOGas Series 3 sterilizer s/n 141182.

  - AN333 EOGas Series 3 User Manual

- Responsibilities

  - The Process Owner is responsible for assisting to develop protocols, and then translating the protocols into executable procedures. The Process Owner is also responsible for addressing and documenting any deviations, along with reviewing and approving the validation protocol.

  - Quality Assurance oversees the validation to ensure that protocol is being followed and parameters are sufficient to validate the system, and approving any actions stemming from deviations. Quality Assurance will review and approve the data and validation protocol.

  - The Executive Vice President is responsible for overseeing the validation and assisting in developing protocols, and ensuring that parameters are sufficient to validate the system. The Executive Vice President will review and approve the validation protocol and any actions stemming from deviations.

  - The Validation Reviewer is responsible for ensuring parameters are sufficient to validate the system. The Validation Reviewer will review and approve the validation protocol.

- Safety Considerations

  - Ethylene oxide is a dangerous chemical and requires care when handling. New and used ethylene oxide canisters should be moved with caution, and not activated until sealed in bag and immediately prior to the sterilization chamber is closed. Suspected leaks or other issues should be reported immediately to a supervisor.

- OQ Documentation Requirements

  - Employees involved in the Operational Qualification will be trained on the protocol prior to beginning any sampling.

  - The protocol and any amendments or deviations will be approved and signed off on before beginning.

  - All records will be retained as per Chemence Medical Inc.’s policies. Records required under OSHA or other governmental agencies will be kept as per specific policies.

- Equipment Description

  - The AN333 EOGas Series 3 is a sterilization cabinet utilizing ethylene oxide to gas-sterilize heat-sensitive medical devices. Devices are loaded into bags and a canister of ethylene oxide is activated, which penetrates through permeable Tyvek® backing to sterilize the product. Ethylene oxide leaks out through the bag and is aerated through fans outside of the facility and away from any operators. After a sterilization cycle, bags are sliced open and a purge cycle is performed to remove any residual ethylene oxide.

- Procedure

  - Materials

    - Calibrated temperature/humidity logger

    - EOGas 3 Sterilization Liner Bag PN 34550053

    - Humiditube

    - Humidichips PN 34550049

    - Blister Packs filled with Dunnage material (Blister PN 34550001)

  - One temperature and humidity profile will be performed and documented in the appendix as outlined in the protocol.

  - As empty and partially loaded sterilizer profiles were generated in the original Operational Qualification, only a fully loaded sterilizer profile will be performed.

    - A fully loaded sterilizer consists of 10 totes, each containing 600 blister packs filled and sealed with dunnage, a humiditube with 2 humidichips and a data-logger.

      - The fully loaded sterilizer run will be documented in the appendix and summarized in the final report.

    - A run lasts 16 hours in INDUSTRIAL MODE.

    - Pre-Conditioning criteria must be met.

    - A calibrated temperature and humidity data-logger will be placed in each tote of the fully loaded sterilizer, for a total of ten (10) data-loggers.

    - Place the humiditube in the center of the load with 2 humidichips (document pre-weight in appendix).

    - After the temperature stabilizes, typically after around an hour, the sterilizer should remain within 50°C ± 3°C.

    - Hot and cold spots captured on the data-loggers will be recorded. Hot or cold spots are defined as a greater than 2°C deviation from the mean temperature recorded in the cabinet.

    - Humidity in each bag/tote must average between 30%-90% R/H.

- Acceptance Criteria

  - All test equipment requiring calibration has been properly calibrated. All certificates or calibration data will be provided.

  - The average preconditioning environment is ≥12hours at ≥68°F and ≥30%RH.

  - The average environment within the EO sterilization bag is between 47°C and 53°C for a 16-hour sterilization cycle. All hot and cold spots will be documented.

  - The average humidity per bag must be between 30%-90% R/H.

- Deviations

  - Any deviations will be listed in the final report of the addendum, along with any rectifying actions and an explanation of the rationale.

- Appendices

  - Appendix A: Training Verification

  - Appendix B: Pre-Conditioning

  - Appendix C: Data Logger Data

  - Appendix D: Transfer and Load

  - Appendix E: Consumable Information

Appendix A

Training Verification

| Printed Name | Signature | Date |
| --- | --- | --- |

Appendix B: Pre-Conditioning

| Staging for EO Sterilization |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Blister Packs Per Tote (Minimum Quantity and Maximum Quantity Are Both 600 Blister Packs per Tote |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Tote 1 |  |  |  | Tote 3 |  |  |  |  | Tote 5 |  |  | Tote 7 |  |  |  |  |  |  | Tote 9 |  |  |
| Tote 2 |  |  |  | Tote 4 |  |  |  |  | Tote 6 |  |  | Tote 8 |  |  |  |  |  |  | Tote 10 |  |  |
| Final Quantity of ALL Blister Packs |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Performed By/Date: |  |  |  |  |  |  |  |  |  |  | Checked By/Date: |  |  |  |  |  |  |  |  |  |  |
| Preconditioning *(Minimum 12 hours) |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Data Logger ID: |  |  |  |  |  |  |  |  |  | Calibration Due Date: |  |  |  |  |  |  |  |  |  |  |  |
| Data Logger Prepared as outlined in WI 09-014 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Prepared By/Date: |  |  |  |  |  |  |  |  |  | Verified By/Date: |  |  |  |  |  |  |  |  |  |  |  |
| Start Time: |  |  |  |  |  | Start Date: |  |  |  | End Time: |  |  |  |  |  |  |  | End Date: |  |  |  |
| Avg. Temperature (≥68ºF): |  |  |  |  |  |  |  |  |  | Avg. Relative Humidity (≥ 30%): |  |  |  |  |  |  |  |  |  |  |  |
| Operator Signature/Date: |  |  |  |  |  |  |  |  |  | Checked By/Date: |  |  |  |  |  |  |  |  |  |  |  |
| QA Signature/Date: |  |  |  |  | ☐ Pass ☐ Fail |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |

Appendix C: Data Logger Data

| Data Logger |  |  |  |  |  |  |  |  |  |  |  | Sterilization Run #: |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Start Time/Date: End Time/Date: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Data Logger ID: |  |  |  |  | Control Number (Location): |  |  |  | Average Temperature |  |  |  |  |  | AVG. RH% |  |  |  |
| Data Logger ID: |  |  |  |  | Control Number (Location): |  |  |  | Average Temperature |  |  |  |  |  | AVG. RH% |  |  |  |
| Data Logger ID: |  |  |  |  | Control Number (Location): |  |  |  | Average Temperature |  |  |  |  |  | AVG. RH% |  |  |  |
| Data Logger ID: |  |  |  |  | Control Number (Location): |  |  |  | Average Temperature |  |  |  |  |  | AVG. RH% |  |  |  |
| Data Logger ID: |  |  |  |  | Control Number (Location): |  |  |  | Average Temperature |  |  |  |  |  | AVG. RH% |  |  |  |
| Data Logger ID: |  |  |  |  | Control Number (Location): |  |  |  | Average Temperature |  |  |  |  |  | AVG. RH% |  |  |  |
| Data Logger ID: |  |  |  |  | Control Number (Location): |  |  |  | Average Temperature |  |  |  |  |  | AVG. RH% |  |  |  |
| Data Logger ID: |  |  |  |  | Control Number (Location): |  |  |  | Average Temperature |  |  |  |  |  | AVG. RH% |  |  |  |
| Data Logger ID: |  |  |  |  | Control Number (Location): |  |  |  | Average Temperature |  |  |  |  |  | AVG. RH% |  |  |  |
| Data Logger ID: |  |  |  |  | Control Number (Location): |  |  |  | Average Temperature |  |  |  |  |  | AVG. RH% |  |  |  |
| Prepared By/Date: |  |  |  |  |  |  |  |  | Verified By/Date: |  |  |  |  |  |  |  |  |  |
| Appendix D: Transfer and Load |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Operation 3 - Transfer & Load Time |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Time left White Room (b): |  |  |  |  |  | Sterilizer Temp (50ºC± 3o): |  |  |  | ºC |  | Vacuum is present: |  |  |  | ☐ Yes ☐ No |  |  |
| Printer Paper/Ink Checked: |  |  |  | ☐ Yes ☐ No |  | "LOAD" pressed on Display: |  |  |  | ☐ Yes ☐ No |  | Equipment Logbook filled out: |  |  |  |  |  | ☐ Yes ☐ No |
| 10 Totes Loaded into Chamber: |  |  |  |  |  | ☐ Yes ☐ No |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  |  | Total Transfer Time (c) (a-b) Not to exceed 17 minutes: |  |  |  |  |  | min |  |  |  |  |
| Chamber Doors Closed and Cycle Started: |  |  |  |  |  | ☐ Yes ☐ No |  | Time/Date that cycle started (countdown completed/doors locked): |  |  |  |  |  |  |  |  |  |  |
| “Sterilization In Process” Sign Visible: |  |  |  |  |  | ☐ Yes ☐ No |  |  |  |  |  |  |  |  |  |  |  |  |
| Operator Signature/Date: |  |  |  |  |  |  |  |  |  |  | Checked By/Date: |  |  |  |  |  |  |  |

Appendix E: Consumable information

| Sterilization Consumable Lot Information |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| EOGas bag lot number(s): |  |  |  |  |  |  |  |  |  |  |  |  |
| Humidichip lot number(s): |  |  |  | Exp. Date: |  |  |  | Scale/Balance ID: |  |  | CAL Due: |  |
| Control Number |  |  | Humidichip Pre Weight (g) |  | Humidichip Post Weight (g) |  | Net Weight Release (g) |  |  |  |  |  |
| Tote #1: |  |  |  |  |  |  |  |  |  |  |  |  |
| Tote #2: |  |  |  |  |  |  |  |  |  |  |  |  |
| Tote #3: |  |  |  |  |  |  |  |  |  |  |  |  |
| Tote #4: |  |  |  |  |  |  |  |  |  |  |  |  |
| Tote #5: |  |  |  |  |  |  |  |  |  |  |  |  |
| Tote #6: |  |  |  |  |  |  |  |  |  |  |  |  |
| Tote #7: |  |  |  |  |  |  |  |  |  |  |  |  |
| Tote #8: |  |  |  |  |  |  |  |  |  |  |  |  |
| Tote #9: |  |  |  |  |  |  |  |  |  |  |  |  |
| Tote #10: |  |  |  |  |  |  |  |  |  |  |  |  |

History Page

| Date | Revision | DCR # | Description |
| --- | --- | --- | --- |
| 08 Nov. 2012 | Original | 2188 | New Document |
| 24 Aug. 2015 | N/A | N/A | - Format change; Made into a fillable form - There was no content change |