- PURPOSE

  - The purpose of this operation qualification (OQ)  is to define the validation activities for the So-Low Cleanroom Refrigerator. The refrigerator will be utilized for storing tissue and processing supplies that are required to be maintained in a refrigerated environment prior to processing.

  - This validation is a deliverable for PCR21-025 New Cleanroom Construction and Support Lab Expansion.

- SCOPE

    - This validation applies to the So-Low Cleanroom Refrigerator for storing tissue and processing supplies that are required to be maintained in a refrigerated environment prior to processing. This validation shall serve as the initial qualification of this equipment.

    - This equipment is located in the newly constructed Cleanroom at the StimLabs Roswell, GA facility.

- BACKGROUND

  - StimLabs is a tissue bank involved in the processing of human derived tissues into products for human use.  Stimlabs has constructed a new ISO Class 7 cleanroom for the processing of human tissue at its facility in Roswell, GA.  During the tissue processing, it is necessary to refrigerate tissue and solutions at various stages in order to preserve the characteristics of the tissue over time.

  - To maintain a refrigerated environment, a refrigerator must maintain its temperature between 1°C and 10°C. The set point drift for the temperature controller must not exceed 2°C in a given time span.  To ensure the refrigerator is fit for the intended purpose, StimLabs will perform an installation qualification and continue to monitor the performance of the refrigerator while it is being used to store human tissue and supplies through an operational qualification.

  - The So-Low Cleanroom Refrigerator has been calibrated prior to the start of the validation activities.  The set point may be adjusted prior to the start of the validation if desired.  Loading the unit will cause the temperature to rise as the warmer product is introduced into the chamber.  If a large amount of product is to be introduced, the manufacturer recommends loading in stages, allowing several hours between stages to allow temperature to stabilize again before the introduction of additional warm product. The manufacture also recommends:

    - Minimum clearances must be maintained.

    - Avoid blocking the evaporator fans on the top shelf. The should be four inches of clearance around the fan to allow proper air flow.

    - Do not overload the unit. The maximum suggested load is about 75% of the chamber volume.

    - Minimize the time the door is left open.

- REFERENCES

  - FDA Guidance for Industry – Process Validation: General Principles and Practices

  - American Association for Tissue Banks Standards for Tissue Banking

  - So-Low DH Model Refrigerator Equipment Manual

  - Food and Drug Administration 21 CFR Part 1271

  - GHTF/SG3/N99-10:2004 Quality Management Systems – Process Validation Guidance

  - VAP21-010-00 So-Low Cleanroom Refrigerator (PVP)

  - VAP21-010-01 So-Low Cleanroom Refrigerator (IQ)

- DEFINITIONS

  - Acceptance Criteria:  Pre-determined requirements upon which pass/fail decisions are based.  Acceptance criteria are specific and measureable.

  - Critical to Quality:  Attribute of a part, assembly, sub-assembly, product or process that has direct and significant impact on its actual or perceived quality to the customer or end user.

  - Deviation:  The action of departing from an established method, procedure, or acceptance criteria.

  - LSL:  Lower Specification Limit; used to determine process capability.

  - Normal Distribution:  A function that represents the distribution of many random variables as a symmetrical, bell-shaped graph.

  - Operational Qualification:  Establishing by objective evidence the process control limits and action levels which result in a product that meets predetermined requirements.

  - PpK/CpK:  Process capability; an estimate of the process capability of a process during its initial set-up before it has been brought into a state of statistical control.

  - Process Validation:  A mechanism to ensure that a process whose output is not fully verifiable by inspection and test is capable of consistently (repeatedly and reliably) producing a result or product that meets pre-determined requirements.  Process validation proves that a process is under control and is used to determine the process variables, acceptable limits for those variables, and to establish appropriate in-process controls.

  - Protocol:  A pre-approved and prospective document that states the requirements and procedures for conducting a validation.

  - USL:  Upper specification limit; used in determining process capability.

  - Validation/Study Leader: The individual who has the responsibility and authority over the operation of the equipment or process being validated.

  - Verification:  Confirmation by examination with a provision of documented evidence that a specific requirement has been fulfilled.

- RESPONSIBILITIES

| Function | Name | Author Protocol | Review & Approve Protocol | Execute Protocol | Execute Testing |
| --- | --- | --- | --- | --- | --- |
| Validation/Study Leader | Michael Baird | X | X |  |  |
| MwER/MR/HOQ/TBD | Melissa O’Connor |  | X |  |  |
| Operations Management | Michael Baird |  | X |  |  |
| Quality Management | Stacy Harmon |  | X |  |  |
| Other | Rob Frye |  | X |  |  |
| Other | Trained Technician |  |  | X | X |

- ENVIRONMENTAL REQUIREMENTS

  - Temperature will be reference only

    - The equipment manual does not state a desired environment temperature range for proper operation.

  - Relative Humidity will be for reference only

    - The equipment manual does not state a desired environment humidity range for proper operation.

- PROCEDURE

  - Manufacturing Equipment Description

    - Manufacturer: So-Low

    - Serial Number(s): will be identified on the data collection sheets for the OQ

    - Model Number: DH4-24GD-TM-CR

    - Location:  New Cleanroom

  - Raw Materials

    - IAS16-006 Polypropylene Containers

    - Water

    - IAS17-042 Chart Recorder Paper

  - Critical to Quality (CTQ) Attributes

    - The critical attribute for this piece of equipment is its ability to maintain a constant set point temperature to achieve a refrigerated environment throughout the chamber when loaded and empty.

  - Process Variables

    - Following the IQ, no adjustments should need to be made to achieve the desired temperature set point.  If adjustment is needed, the set point should be set to 5°C.

    - OQ Run 1 – Empty Chamber

      - Load a new sheet of chart recorder paper in the chart recorder and ensure the pen is touching the paper.

      - Starting with the unit on, set at 5°C, and chamber empty, insert 3 thermocouples inside the chamber.  Measurement points will be taken from the bottom, middle, and top of the chamber at the centermost point.  The access port in the back of the unit is at the top, thus a temperature probe run through the access port cannot reach the bottom shelf.  Therefore, the thermocouple placed on the bottom shelf will be placed so that the wire runs through the front door to the data logger.

      - Close the door. Ensure the thermocouple wire at the bottom remain in the place after the door is shut.  Ensure the set point is set to 5℃.

      - Record the temperature for a period not less than 24 hours

      - At the end of the run, remove the thermocouples and analyze the data

      - Remove the chart recorder paper and write “VAP21-010-02 Run 1” on the paper, initial, and date.  The chart recorder paper should also be reviewed by QA.

    - OQ Run 2 – Full Chamber

      - Load a new sheet of chart reorder paper in the chart recorder.  Ensure the pen is touching the paper.

      - Starting with the unit on and at set point, load the chamber full of representative samples (20 polypropylene containers filler with 1000ml water on each shelf for a total of 80 containers. Distribute the load as evenly as possible throughout the chamber to represent operation conditions.

      - Insert 3 thermocouples inside the loaded chamber.  Measurement points will be taken from the bottom, middle, and top of the chamber at the centermost point.  The access port in the back of the unit is at the top, thus a temperature probe run through the access port cannot reach the bottom shelf.  Therefore, the thermocouple placed on the bottom shelf will be placed so that the wire runs through the front door to the data logger.

      - Close the door. Ensure the bottom thermocouple wire remain in the place after the door is shut.  Ensure the set point is set to 5°C.

      - Begin recording the temperature when the door is shut. Record the temperature for a period not less than 24 hours

      - At the end of the run, remove the thermocouples and analyze the data. Discard dummy materials.

      - Remove the chart recorder paper and write “VAP21-010-02 Run 2” on the paper, initial, and date.  The chart recorder paper should also be reviewed by QA.

  - Test Equipment

    - Data Logger with thermocouples (NIST Traceable)

- LOT SIZE & SAMPLING PLAN

  - Run 1

    - The refrigerator shall be monitored from 5°C  start with an empty chamber to determine if the chamber maintains a consistent temperature over 24-hour period.

  - Run 2

    - The refrigerator shall be monitored after being loaded with dummy product to confirm the ramp time and return to set point after loading and that the temperature set point and uniformity can be maintained while the chamber is fully loaded over a 24 hour period.

- ACCEPTANCE CRITERIA

  - Run 1 Temperature Uniformity

    - The temperature difference measured by the three thermocouples (overall average) must not vary more than 2°C  from each other throughout the entire run after the set-point is reached.

    - The overall average temperature measured by each thermocouple after set point is reached over 24 hours must not exceed more than 7°C  or fall below 3°C  throughout the entire run set point is reached. To pass the empty chamber OQ individual temperature reading as low as 2°C  or as high as 8°C  shall be accepted.

  - Run 2 Temperature Uniformity

    - The temperature difference measured by the three thermocouples (overall average) must not vary more than 2°C  from each other throughout the entire run after the set-point is reached.

    - The overall average temperature measured by each thermocouple after set point is reached over 24 hours must not exceed more than 7°C  or fall below 3°C  throughout the entire run set point is reached. To pass the empty chamber OQ individual temperature reading as low as 2°C  or as high as 8°C  shall be accepted.

- MATERIAL DISPOSITION

  - Dummy material shall be disposed at the completion of the validation or returned to inventory following proper decontamination, cleaning, and sterilization procedures.

- DOCUMENTATION REQUIREMENTS

  - Results shall be summarized in an OQ Report.

  - Various Control Plans, including validated process settings and the re-validation criteria shall be documented in the final report.

  - Raw data sheets shall be included in the appendix of the OQ report and/or a reference to the location of the raw data shall be included in the OQ report.

  - Calibration records (if applicable) shall be attached to the OQ report (in the Appendix).

  - Signed OQ Protocol Training forms will be attached to the OQ report (in the Appendix).

- DEVIATIONS

  - Deviations to the OQ will be pre-approved using FRM020-012 Protocol Deviation Form and explained in the OQ report.

  - Copies of signed Protocol Deviation Forms will be attached to the OQ report (in the Appendix).

- PRE-REQUISITE STEPS

  - Approved Process Validation Plan (PVP) and IQ Report (if required per PVP), or justification for omission.

  - Calibration certificates for equipment and measurement devices (gauges)

  - Approval and release of raw materials

  - Verification of environmental requirements

  - Completed Operator Training

  - Draft work instructions

  - The OQ Protocol must be signed before validation activities may begin.

- APPENDIX

  - Blank data collection sheets

APPENDIX 15.1 – DATA COLLECTION SHEETS

STOP!  Read Instructions.

Instructions:

- Complete one data collection sheet (pages 8-12) for each So-Low Cleanroom Refrigerator.

- Verify the Pre-requisite criteria have been met on this data collection sheet prior to testing.  Note that the Pre-requisite criteria only need to be verified once for the validation.

- For attribute data: Circle “Pass” or “Fail” or write the words “Pass” or “Fail”

- For variable data: Record actual results in the units specified or add units if the units are not specified.

- Verification must include signature of the individual performing the step and the date performed.  Verification must be completed concurrently with performance.

- Follow Good Documentation Practice (GDP).

Pre-requisite Criteria (must be complete and verified prior to start of OQ activities):

| Pre-requisite | Complete (Circle One) | Verification Signature/Date |
| --- | --- | --- |
| Process Validation Plan is approved. | Yes No |  |
| IQ Protocol is completed | Yes No |  |
| OQ Protocol approved | Yes No |  |
| Unit is Calibrated | Yes No |  |
| Operator Training is complete and documented. | Yes No |  |

Pre-Requisites Verified By (signature & date): _____________________________________________________________

Quality Reviewed By (signature & date): ______________________________________________________________

Section A – OQ Run #1 Preparation

| Equipment Name | Model Number | Manufacturer | Serial Number | Stimlabs ID Number | Verification Signature/Date |
| --- | --- | --- | --- | --- | --- |

| Specification and Acceptance Criteria | Reading/Results | Result | Verification Signature/Date |
| --- | --- | --- | --- |
| Chart Recorder Set Up |  |  |  |
| Paper is loaded | N/A | Pass / Fail |  |
| Pen is contacting the paper | N/A | Pass / Fail |  |
| Thermocouple Set Up |  |  |  |
| Thermocouples are loaded on the bottom, middle, and top of chamber. | Data Logger ID Top_________________ Middle _____________ Bottom _____________ | Pass / Fail |  |

Section A Completed By (signature & date): ________________________________________________________________

Quality Reviewed By (signature & date): ______________________________________________________________

Section B - OQ Run #1

| Specification or Acceptance Criteria | Reading/Results | Result | Verification Signature/Date |
| --- | --- | --- | --- |
| Chamber unit should be on and empty. Close the unit door after thermocouples are loaded. Ensure set point is set to 5°C. Record the temperature ramp for a period not less than 24 hours. Attach copies of data collected by data loggers to the final OQ report. |  |  |  |
| Enter date and time test began empty chamber. |  | N/A |  |
| Enter set point setting (5°C) |  | N/A |  |
| Enter date and time set point was checked (at least 24 hours of run time) |  | N/A |  |
| Display reading after 24 hours. (5°C ± 2°C) |  | Pass / Fail |  |
| Average temperature measure over 24 hours for bottom (5°C ± 2°C) |  | Pass / Fail |  |
| Average temperature measure over 24 hours for middle (5°C ± 2°C) |  | Pass / Fail |  |
| Average temperature measure over 24 hours for top (5°C ± 2°C) |  | Pass / Fail |  |

Section B Completed By (signature & date): ________________________________________________________________

Quality Reviewed By (signature & date): ______________________________________________________________

Section C – Run #2 Preparation

| Equipment Name | Model Number | Manufacturer | Serial Number | Stimlabs ID Number | Verification Signature/Date |
| --- | --- | --- | --- | --- | --- |

| Specification and Acceptance Criteria | Reading/Results |  | Result | Verification Signature/Date |  |
| --- | --- | --- | --- | --- | --- |
| Chart Recorder Set Up |  |  |  |  |  |
| Paper is loaded | N/A |  | Pass / Fail |  |  |
| Pen is contacting the paper | N/A |  | Pass / Fail |  |  |
| Dummy Load Assembly and Set Up |  |  |  |  |  |
| 80 polypropylene container filled with 1000L of water. | Qty of polypropylene containers filled with 2L of water assembled: ______________ | Pass / Fail |  |  |  |
|  | Verified 1000L of water in polypropylene container | Pass / Fail |  |  |  |
| 20 polypropylene containers filled with 1000L of water on each shelf for a total of 80 containers. | Qty per shelf Top Shelf 1: ________ Shelf 2: _______ Shelf 3: _______ Bottom Shelf 4: ____ | Pass / Fail |  |  |  |
| Thermocouple Set Up |  |  |  |  |  |
| Thermocouples are loaded on the bottom, middle, and top of chamber. | Data Logger ID Top_________________ Middle _____________ Bottom _____________ |  | Pass / Fail |  |  |

Section C Completed By (signature & date): ________________________________________________________________

Quality Reviewed By (signature & date): ______________________________________________________________

Section D - OQ Run #2

| Specification or Acceptance Criteria | Reading/Results | Result | Verification Signature/Date |
| --- | --- | --- | --- |
| Chamber unit should be LOADED and thermocouples placed around the items on top, middle, and bottom. Close the unit door after thermocouples are loaded. Ensure set point is set to 5°C. When the door is shut, begin recording the temperature and continue for a period not less than 24 hours. Attach copies of data collected by data loggers to the final OQ report. |  |  |  |
| Enter date and time door was shut on the loaded unit. |  | N/A |  |
| Enter set point setting (5°C) |  | N/A |  |
| Enter date and time set point was checked (at least 24 hours of run time) |  | N/A |  |
| Display reading after 24 hours. (5°C ± 2°C) |  | Pass / Fail |  |
| Average temperature measure over 24 hours for bottom (5°C ± 2°C) |  | Pass / Fail |  |
| Average temperature measure over 24 hours for middle (5°C ± 2°C) |  | Pass / Fail |  |
| Average temperature measure over 24 hours for top (5°C ± 2°C) |  | Pass / Fail |  |

Section D Completed By (signature & date): ________________________________________________________________

Quality Reviewed By (signature & date): ______________________________________________________________