- PURPOSE

  - The purpose of this validation is to provide individuals who’s responsibility is to approve the operational qualification protocol that did not approve REV 01 for the pouch sealer sealing process are effective and provides acceptable results when tested at the limits of the process.

  - This validation is a deliverable for PCR21-025 New Cleanroom Construction and Support Lab Expansion.

- SCOPE

  - This validation applies to the Van der Stahl heat sealer.  A heat sealer is used to seal inner and outer product pouches.  This validation includes the qualification of the heat sealer equipment and heat sealing process.  The inner pouch will maintain a sterile barrier.  Sterility maintenance was validated under VAP/VAR16-002 and the process settings are the same.

  - This equipment is located in the new cleanroom at the StimLabs Roswell, GA facility.

  - 4.5” x 6.75” Inner pouch, vial pouch, and outer pouch are constructed of different material, thus necessitating a different set of parameters to achieve an acceptable seal.  The sealers will be validated to define the process parameters for sealing both inner and outer pouches using the validated process setting.

- BACKGROUND

  - StimLabs is a tissue bank involved in the processing of human derived tissues into products for transplantation.  In order to provide the tissue products in a manner that complies with sterile transplantation procedures, StimLabs packages the tissue product in a dual pouch configuration, including both an inner and outer pouch.

  - To ensure the pouch sealer is fit for the intended purpose, StimLabs will perform an operational qualification and continue to monitor the performance of the pouch sealer while it is being used to seal product pouches.

- REFERENCES

  - Food and Drug Administration 21 CFR Part 1271

  - American Association for Tissue Banks Standards for Tissue Banking

  - Van Der Stahl Heat Sealer Equipment Manual

  - FDA Guidance for Industry – Process Validation: General Principles and Practices

  - GHTF/SG3/N99-10:2004 Quality Management Systems – Process Validation Guidance

  - FRM017-003 Seal Integrity Peel Test Form

  - SOP017-003 Physical Testing of Heat Seals Test Method

  - VAP21-014-01 Van Der Stahl Heat Sealer Installation Qualification

- DEFINITIONS

  - Acceptance Criteria:  Pre-determined requirements upon which pass/fail decisions are based.  Acceptance criteria are specific and measureable.

  - Critical to Quality:  Attribute of a part, assembly, sub-assembly, product or process that has direct and significant impact on its actual or perceived quality to the customer or end user.

  - Deviation:  The action of departing from an established method, procedure, or acceptance criteria.

  - LSL:  Lower Specification Limit; used to determine process capability.

  - Normal Distribution:  A function that represents the distribution of many random variables as a symmetrical, bell-shaped graph.

  - Operational Qualification:  Establishing by objective evidence the process control limits and action levels which result in a product that meets predetermined requirements.

  - PpK/CpK:  Process capability; an estimate of the process capability of a process during its initial set-up before it has been brought into a state of statistical control.

  - Process Validation:  A mechanism to ensure that a process whose output is not fully verifiable by inspection and test is capable of consistently (repeatedly and reliably) producing a result or product that meets pre-determined requirements.  Process validation proves that a process is under control and is used to determine the process variables, acceptable limits for those variables, and to establish appropriate in-process controls.

  - Protocol:  A pre-approved and prospective document that states the requirements and procedures for conducting a validation.

  - USL:  Upper specification limit; used in determining process capability.

  - Validation/Study Leader: The individual who has the responsibility and authority over the operation of the equipment or process being validated.

  - Verification:  Confirmation by examination with a provision of documented evidence that a specific requirement has been fulfilled.

| Function | Name | Author Protocol | Review & Approve Protocol | Execute Protocol | Update WI, SOP, Forms |
| --- | --- | --- | --- | --- | --- |
| Validation/Study Leader | Michael Baird | X | X |  |  |
| MwER/MR/HOQ/TBD | Melissa O’Connor |  | X |  |  |
| Operations Management | Michael Baird |  | X |  | X |
| Quality Management | Stacy Harmon |  | X |  |  |
| Validation Reviewer | Rob Frye | X | X |  |  |
| Sealing Pouches | Trained Personnel/ Technician(s) |  |  | X |  |
| Seal Testing & Bubble Emission | Trained QC Lab Technician (s) |  |  | X |  |

- RESPONSIBILITIES

- ENVIRONMENTAL REQUIREMENTS

  - N/A

- PROCEDURE

  - Manufacturing Equipment Description

    - Manufacturer: Van der Stahl

    - Serial number: will be identified in the OQ data sheets

    - Model Number: MS350

    - Location: New Cleanroom

  - Raw Materials

    - IAS19-009 – Primary Pouch (Inner Pouch)

      - 5.0" L x 3.5" W, Seal Width 0.25", Chevron Gap 0.5"

      - Steriflex-101 (48g PET/0.002LDPE)

      - Steriflex-112 (48g PET/0.002LDPE Peelable)

      - IAS16-015 – Primary Pouch (Inner Pouch)

      - 6.75" L x 4.5" W, Seal Width 0.25", Chevron Gap 0.5"

      - Steriflex-101 (48g PET/0.002LDPE)

      - Steriflex-112 (48g PET/0.002LDPE Peelable)

    - IAS16-016 - Outer Pouch

      - 9.25"L x 6.0" W, Seal Width 0.25", Chevron Gap 0.625"

      - Steriflex-301 (48g PET/98gWOPP/0.00035Foil/0.002LDPE)

      - Steriflex-301P (48g PET/98gWOPP/0.00035Foil/0.002Peelable)

    - Each pouch shall be empty when sealed for testing.  Product sealed in the pouches for production are two dimensional and would not challenge the seal.

  - Critical to Quality (CTQ) Attributes

    - Heating Temperature

      - Importance of factor

        - In the mechanism of producing a seal, the pouch material must be heated and thus result in molecular activity to a point of fluidity. Thus, the material flows between the mating surfaces to form the seal.

      - Effect on the response from varying factor

        - Insufficient temperature will not activate the adhesive, thus not having time to form a bond. Excessive temperature could burn/melt the mating surface, resulting in flange warpage/curling.

      - Control of factor

        - Temperature is controlled by a digital controller using a temperature sensor in direct contact to the heating element. This controller can be adjusted manually using specific output operation buttons located on the face of the control panel.

    - Cooling Temperature

      - Importance of Factor

        - In the mechanism of producing a seal a cooling cycle temperature must be established to allow film in the seal area to cool and reset under pressure providing consistent maximum seal properties. The bond that is created during the "cool down” period produces a very strong seal with excellent visual appearance.

      - Effect on the Response from Varying the Factor

        - Without cooling cycle, the sealed film/area could shrink or get cut off in worst case.  With the cooling cycle, impulse sealer produces a lot cleaner and more solid sealing results than constant heat sealer that has no cooling process.

      - Control of Factor

        - Cooling temperature is controlled by a digital temperature controller, which can be adjusted manually using specific output operation buttons located on the face of the control panel.

    - Seal Time

      - Importance of Factor

        - In the mechanism of producing a seal a sufficient seal time must be established to allow for the mating surface material to activate and bond.

      - Effect on the Response from Varying the Factor

        - If insufficient seal time is allotted, the mating surfaces will not have time to form a bond.  Excessive time could result in the burning of material, curling or warping of flange if the temperature is too high.

      - Control of Factor

        - Seal time is controlled by a digital timer, which can be adjusted manually using specific output operation buttons located on the face of the control panel.

  - Process Variables

    - Equipment Identification

      - Visually verify that the manufacturer model number, serial number and the Equipment Identification Number are affixed to each piece of equipment.  Record the results on the associated data sheet.

    - Process Parameters Challenge

      - Verify the following items are completed:

        - IQ activities completed.

        - Equipment is in a valid state of calibration.

        - Equipment is in a valid PM state.

    - Equipment Operation

      - Verify that the equipment powers up and down as expected per the associated datasheet.

      - Verify that operating control button’s function and equipment operates as expected per the associated datasheet.

    - Equipment Operation

      - Verify that the equipment powers up and down as expected per the associated datasheet.

| Factor / Condition | Range |  | Critical to Process (Y/N) | Test Factor (Y/N) |
| --- | --- | --- | --- | --- |
|  | Inner Pouch | Outer Pouch |  |  |
| Seal Heat Temperature | 285 ± 10°F | 325°F ± 10°F | Y | Y |
| Seal Time | 3.0 ± 2.0 sec. | 3.0 ± 2.0 sec. | Y | Y |
| Cooling Temperature | 150°F ± 10°F | 150°F ± 10°F | Y | Y |

    - The following will be used for the production of samples required for this protocol:

      - Before using the sealer, ensure the unit is connected to a power source.  Using the power switch, turn the unit on and make sure the power indicator light is on.

      - Carefully place open end of the pouch between the jaws of the sealer.  When the pouch is placed in the correct location, depress the foot switch to close the jaws and begin the seal sequence.

CAUTION

Never place hands between jaws!

    - Worse case combination of temperature and seal time will be evaluated.  Nominal setting will be evaluated to aid in finalizing process parameters. This evaluation will be conducted separately during Process Qualification (PQ) for the Pouch Sealer.

| Outer Pouch |  |  |  | Both Inner Pouches |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Factor | Units | Low | High | Factor | Units | Low | High |
| Sealing Temp. | F° | 315 | 335 | Sealing Temp. | F° | 275 | 295 |
| Dwell Time | Sec. | 1.0 | 5.0 | Dwell Time | Sec. | 1.0 | 5.0 |
| Cooling Temp. | F° | 140 | 160 | Cooling Temp. | F° | 140 | 160 |

    - The following describes the method to be used for producing and testing the test units.

      - One run at Low (L) and one run at High (H)  parameters will be performed.

      - Mark each Pouch with  Trial #, Test Method, and “VAP21-014-02”

      - Record pouch information and sealing process parameters in the associated datasheet.

| Outer Pouch 9.25ʺ x 6.0ʺ |  |  | Inner Pouch 6.75ʺx 4.5ʺ |  |  | Inner Pouch 5.0ʺx 3.5ʺ |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Trial # | Condition | # of Sample | Trial # | Condition | # of Sample | Trial # | Condition | # of Sample |
| 1 | L | 89 | 3 | L | 89 | 5 | L | 89 |
| 2 | H | 89 | 4 | H | 89 | 6 | H | 89 |

* Number of samples are derived from 30 for Peel Strength and 59 for Bubble Emission for a total # of samples to equal 89 samples per trial number.

    - Visual Inspection

      - Inspect all seals for any damage or other signs of improper seal integrity.  Seals which have burns, holes or other signs of improper seal integrity are rejected. Record results on the associated datasheet.

      - All seals must be a minimum of 0.25 inches in width. Seals which have widths less than 0.25 inches are rejected.  Record results on the associated datasheet.

  - Test Equipment

    - Van der Stahl PPT-100V Simple Peel Pouch Tester

    - Calibrated Ruler

  - Test Methods

    - Pouch Peel Strength Test

      - A peel test is performed by a trained StimLabs technician. (30) sealed pouches from each trial will be tested.  Note and record all results a Seal Integrity Peel Test Form.

      - During the peel test process, observe peeled samples for sign of fiber tear. Note and record all results on the associated datasheet.

    - Bubble Emission Test

      - A bubble emissions test is performed by outside vendor or internal technician per ASTM Standard F 2096-04 on (59) sealed pouches from each trial.

      - This test method applies to the detection of gross leaks in nonporous tray and pouch packages. Record results using the associated datasheet.

- LOT SIZE & SAMPLING PLAN

  - Sampling Plan & Statistical Rationale

    - Attribute Data (95/95): A sample size of at least 59 with zero non-conformances, provides a 95% confidence level that 95% of the population is within specification. This sample size is based on a nonparametric attribute plan derived from the binomial distribution.

    - Variable Data (95/95): A minimum of thirty (30) samples will be pulled from each lot for peel strength test. A sample size of thirty will make it possible to reliably determine if the data collected can be represented by the normal distribution. An accurate assessment of data normality is necessary to choose the proper statistical analysis method.

    - If normal – one sided lower limit x-ks methodology can be used to show that with 95% confidence, at least 95% of the population will exhibit a pouch peel strength greater than or equal to 1.41 lb/in.

      - x-ks ≥ 1.41

      - x – average (lb)

      - s – standard deviation

      - x-ks - The mean minus the k value multiplied by the standard deviation represents the lower statistical tolerance limit for the data set under consideration

      - k – 2.220 (based on 95% confidence / 95% conformance requirements.

    - If not normal – transform to determine if the data better fit any other known distribution and perform process capability analysis of the transformed data and verify that with 95% confidence, at least 95% of the population will exhibit pouch seal strength greater than or equal to 1.41 lb/in.

- ACCEPTANCE CRITERIA

  - Acceptance Criteria for Success

    - Visual Inspection

      - No damage, delamination, or other possible breaches such as channels, bubbles, tears, or punctures are permitted on the pouch seals.

      - No signs of fiber tear observed during the peel test process is permitted.

    - Dimensional Inspections

      - All seals must be a minimum of 0.25 inches in width.

    - Pouch Seal Strength Test

      - All pouches must exhibit a seal strength of ≥ 1.41 lb/in and the minimum test pressure per the peel strength method are reached.

        - Note: The limit for seal peel strength has been established based on the validation that was performed for the heal sealer equipment (see VAP/VAR16-005-03, VAP/VAR16-006-03) where identical pouch materials, equipment, and heat sealer settings were utilized during the execution of a performance qualification (PQ). The results of the heat sealer validation (3 runs) demonstrated a minimum seal peel strength of 1.41 lb/in. (this was the lowest value observed). The sterility maintenance of the identical pouch material sealed using the same materials, equipment and process settings has been demonstrated in VAP/VAR16-021.

    - Bubble Emission Test

      - Pouches must exhibit no gross leaks.

- MATERIAL DISPOSITION

  - All samples resulting from the execution of this OQ will be consumed for testing or discarded.

- DOCUMENTATION REQUIREMENTS

  - Results shall be summarized in an OQ Report.

  - Various Control Plans, including validated process settings and the re-validation criteria shall be documented in the final report.

  - Raw data sheets shall be included in the appendix of the OQ report and/or a reference to the location of the raw data shall be included in the OQ report.

  - Calibration records (if applicable) shall be attached to the OQ report (in the Appendix).

  - Signed OQ Protocol Training forms will be attached to the OQ report (in the Appendix).

- DEVIATIONS

  - Deviations to the OQ will be pre-approved using FRM020-012 Protocol Deviation Form and explained in the OQ report.

  - Copies of signed Protocol Deviation Forms will be attached to the OQ report (in the Appendix).

- PRE-REQUISITE STEPS

  - Approved Process Validation Plan (PVP) and IQ Report (if required per PVP), or justification for omission.

  - Calibration certificates for equipment and measurement devices (gauges)

  - Approval and release of raw materials

  - Verification of environmental requirements

  - Completed Operator Training

  - Draft work instructions

  - The OQ Protocol must be signed before validation activities may begin.

- APPENDIX

  - Blank data collection sheets

APPENDIX 15.1 – DATA COLLECTION SHEETS

STOP!  Read Instructions.

Instructions:

- Complete one data collection sheet (pages 12 – 18) for each Van der Stahl Heat Sealer.

- Verify the Pre-requisite criteria have been met on this data collection sheet prior to testing.  Note that the Pre-requisite criteria only need to be verified once for the validation.

- For attribute data: Circle “Pass” or “Fail” or write the words “Pass” or “Fail”

- For variable data: Record actual results in the units specified or add units if the units are not specified.

- Verification must include signature of the individual performing the step and the date performed.  Verification must be completed concurrently with performance.

- Follow Good Documentation Practice (GDP).

Pre-requisite Criteria (must be complete and verified prior to start of IQ activities):

| Pre-requisite | Complete (Circle One) | Verification Signature/Date |
| --- | --- | --- |
| Process Validation Plan is approved. | Yes No |  |
| IQ Protocol is completed | Yes No |  |
| OQ Protocol approved | Yes No |  |
| Unit is Calibrated | Yes No |  |
| Operator Training is complete and documented. | Yes No |  |

Pre-Requisites Verified By (signature & date): _____________________________________________________________

Quality Reviewed By (signature & date): ______________________________________________________________

NOTE: PRINT MULTIPLE DATA SHEETS FOR EACH SEAL MACHINE

Section A – Equipment Setup

| Equipment Name | Model Number | Manufacturer | Serial Number | Stimlabs ID Number | Verification Signature/Date |
| --- | --- | --- | --- | --- | --- |

| Operating Control | Requirement | Pass/Fail | Verification Signature/Date |
| --- | --- | --- | --- |
| On/Off Switch | Turn Switch to On- Digital Display lights |  |  |

Section A Completed By (signature & date): ________________________________________________________________

Quality Reviewed By (signature & date): ______________________________________________________________

Section B – Trial 1 Low Outer Pouch 9.25" x 6.0"

| Specification or Acceptance Criteria | Reading/Results | Result | Verification Signature/Date |
| --- | --- | --- | --- |
| FRM017-003 will be filled out peel strength test and trial number will be documented on the form |  |  |  |
| Lot number of outer pouch |  | N/A |  |
| Set Seal Temperature. 315°F | N/A | Pass / Fail |  |
| Set Dwell Time. 1 second | N/A | Pass / Fail |  |
| Set Cooling Temperature. 140°F | N/A | Pass / Fail |  |
| 89 Outer Pouches sealed and labeled for appropriate test. 30 – Seal Peel Strength 59 – Bubble Emission | Total Pouches Sealed | Pass / Fail |  |
|  | Total Pouches Sealed and labeled for Peel Strength ________________ | Pass / Fail |  |
|  | Total Pouches Sealed and labeled for Bubble Emission ________________ | Pass / Fail |  |
| All seals met minimum seal width of 0.25” | N/A | Pass / Fail |  |

Section B Completed By (signature & date): ________________________________________________________________

Quality Reviewed By (signature & date): ______________________________________________________________

Section C – Trial 2 High Outer Pouch 9.25" x 6.0"

| Specification or Acceptance Criteria | Reading/Results | Result | Verification Signature/Date |
| --- | --- | --- | --- |
| FRM017-003 will be filled out peel strength test and trial number will be documented on the form |  |  |  |
| Lot number of outer pouch |  | N/A |  |
| Set Seal Temperature. 335°F | N/A | Pass / Fail |  |
| Set Dwell Time. 5 second | N/A | Pass / Fail |  |
| Set Cooling Temperature. 160°F | N/A | Pass / Fail |  |
| 89 Outer Pouches sealed and labeled for appropriate test. 30 – Seal Peel Strength 59 – Bubble Emission | Total Pouches Sealed _________________ | Pass / Fail |  |
|  | Total Pouches Sealed and labeled for Peel Strength ________________ | Pass / Fail |  |
|  | Total Pouches Sealed and labeled for Bubble Emission ________________ | Pass / Fail |  |
| All seals met minimum seal width of 0.25” | N/A | Pass / Fail |  |

Section C Completed By (signature & date): ________________________________________________________________

Quality Reviewed By (signature & date): ______________________________________________________________

Section D – Trial 3 Low 4.5” x 6.75” inner pouches

| Specification or Acceptance Criteria | Reading/Results | Result | Verification Signature/Date |
| --- | --- | --- | --- |
| FRM017-003 will be filled out peel strength test and trial number will be documented on the form |  |  |  |
| Lot number of outer pouch |  | N/A |  |
| Set Seal Temperature. 275°F | N/A | Pass / Fail |  |
| Set Dwell Time. 1 second | N/A | Pass / Fail |  |
| Set Cooling Temperature. 140°F | N/A | Pass / Fail |  |
| 89 4.5” x 6.75” inner pouches sealed and labeled for appropriate test. 30 – Seal Peel Strength 59 – Bubble Emission | Total Pouches Sealed | Pass / Fail |  |
|  | Total Pouches Sealed and labeled for Peel Strength ________________ | Pass / Fail |  |
|  | Total Pouches Sealed and labeled for Bubble Emission ________________ | Pass / Fail |  |
| All seals met minimum seal width of 0.25” | N/A | Pass / Fail |  |

Section D Completed By (signature & date): ________________________________________________________________

Quality Reviewed By (signature & date): ______________________________________________________________

Section E – Trial 4 High 4.5” x 6.75” inner pouches

| Specification or Acceptance Criteria | Reading/Results | Result | Verification Signature/Date |
| --- | --- | --- | --- |
| FRM017-003 will be filled out peel strength test and trial number will be documented on the form |  |  |  |
| Lot number of outer pouch |  | N/A |  |
| Set Seal Temperature. 295°F | N/A | Pass / Fail |  |
| Set Dwell Time. 5 second | N/A | Pass / Fail |  |
| Set Cooling Temperature. 160°F | N/A | Pass / Fail |  |
| 89 4.5” x 6.75” inner pouches sealed and labeled for appropriate test. 30 – Seal Peel Strength 59 – Bubble Emission | Total Pouches Sealed: _________________ | Pass / Fail |  |
|  | Total Pouches Sealed and labeled for Peel Strength ________________ | Pass / Fail |  |
|  | Total Pouches Sealed and labeled for Bubble Emission ________________ | Pass / Fail |  |
| All seals met minimum seal width of 0.25” | N/A | Pass / Fail |  |

Section E Completed By (signature & date): ________________________________________________________________

Quality Reviewed By (signature & date): ______________________________________________________________

Section F – Trial 5 Low 3.5” x 5” inner pouches

| Specification or Acceptance Criteria | Reading/Results | Result | Verification Signature/Date |
| --- | --- | --- | --- |
| FRM017-003 will be filled out peel strength test and trial number will be documented on the form |  |  |  |
| Lot number of outer pouch |  | N/A |  |
| Set Seal Temperature. 275°F | N/A | Pass / Fail |  |
| Set Dwell Time. 1 second | N/A | Pass / Fail |  |
| Set Cooling Temperature. 140°F | N/A | Pass / Fail |  |
| 89 3.5” x 5” inner pouches sealed and labeled for appropriate test. 30 – Seal Peel Strength 59 – Bubble Emission | Total Pouches Sealed | Pass / Fail |  |
|  | Total Pouches Sealed and labeled for Peel Strength ________________ | Pass / Fail |  |
|  | Total Pouches Sealed and labeled for Bubble Emission ________________ | Pass / Fail |  |
| All seals met minimum seal width of 0.25” | N/A | Pass / Fail |  |

Section F Completed By (signature & date): ________________________________________________________________

Quality Reviewed By (signature & date): ______________________________________________________________

Section G – Trial 6 3.5” x 5” inner pouches

| Specification or Acceptance Criteria | Reading/Results | Result | Verification Signature/Date |
| --- | --- | --- | --- |
| FRM017-003 will be filled out peel strength test and trial number will be documented on the form |  |  |  |
| Lot number of outer pouch |  | N/A |  |
| Set Seal Temperature. 295°F | N/A | Pass / Fail |  |
| Set Dwell Time. 5 second | N/A | Pass / Fail |  |
| Set Cooling Temperature. 160°F | N/A | Pass / Fail |  |
| 89 3.5” x 5” inner pouches sealed and labeled for appropriate test. 30 – Seal Peel Strength 59 – Bubble Emission | Total Pouches Sealed | Pass / Fail |  |
|  | Total Pouches Sealed and labeled for Peel Strength ________________ | Pass / Fail |  |
|  | Total Pouches Sealed and labeled for Bubble Emission ________________ | Pass / Fail |  |
| All seals met minimum seal width of 0.25” | N/A | Pass / Fail |  |

Section G Completed By (signature & date): ________________________________________________________________

Quality Reviewed By (signature & date): ______________________________________________________________