- PURPOSE

  - The purpose of this validation document is to outline the Operational Qualification (OQ) protocol for the ViewPoint Cloud Continuous Monitoring System (VPC CMS or VPC) according to the manufacturer’s (Mesa Labs, Inc. or MLI) and the customer’s (Stimlabs, LLC or Stimlabs) specifications, procedures, and requirements.

- SCOPE

  - The scope of this validation protocol will verify the VPC CMS is fully operational, all sensors are functioning properly, data readings are valid and in range according to specifications outlined in QSP022-001 Storage of Tissue and Supplies, alarm functionality operates per MLI’s and Stimlabs’ specifications, alarm notifications are properly formatted and received by designated recipients via the expected method, and the reports created by the VPC system contain the correct information.

  - Validation of the software associated with the VPC CMS is considered out of scope. The software associated with the VPC CMS is defined as Off-the-Shelf Software (OTS) where the application was developed and validated by the manufacturer (MLI). Customizations requiring additional validation testing of the OTS are unnecessary because configurability functions, which are considered standard features of the validated OTS, will be utilized by the manufacturer in configuring the software to Stimlabs’ requirements. The VPC software is a cloud-hosted, web-based application and requires no installations on Stimlabs’ network infrastructure. A compliance certification attesting the software has been sufficiently tested and shown to possess all features and functionalities required for compliance with FDA 21CFR Part 11 will be attached as objective evidence of software validation by the manufacturer in this validation package.

- BACKGROUND

  - Stimlabs is a tissue bank involved in processing human-derived cells and tissues to generate Human Cells, Tissues, and Cellular and Tissue-Based Products (HCT/P’s). Stimlabs is responsible for human-derived cells and tissues that require specific environmental conditions throughout production of HCT/P’s. To facilitate the achievement, maintenance, and traceability of production environment requirements, Stimlabs will utilize the VPC CMS for continuous, local, and remote monitoring and management of environmentally controlled areas and equipment.

- REFERENCES

  - Food and Drug Administration 21 CFR Part 1271

  - American Association for Tissue Banks Standards for Tissue Banking

  - CAPA22-003 Data lag identified in Cobex Temperature chart recorders

  - FDA Guidance for Industry – Process Validation: General Principles and Practices

  - GHTF/SG3/N99-10:2004 Quality Management Systems – Process Validation Guidance

  - PCR22-007 Installation of Continuous Environmental Monitoring System

  - QSP014-001 Environmental Control and Monitoring

  - QSP020-001 Process and Equipment Validation

  - QSP022-001 Storage of Tissue and Supplies

  - SOP013-001 Master Safety Plan

  - VAP22-006-00 ViewPoint Cloud Continuous Monitoring System Process Validation Plan

  - VAP22-006-01 ViewPoint Cloud Continuous Monitoring System Installation Qualification Protocol

  - VAP22-006-03 ViewPoint Cloud Continuous Monitoring System Performance Qualification Protocol

  - VAR22-006-01 ViewPoint Cloud Continuous Monitoring System Installation Qualification Report

  - VAR22-006-02 ViewPoint Cloud Continuous Monitoring System Operational Qualification Report

  - VAR22-006-03 ViewPoint Cloud Continuous Monitoring System Performance Qualification Report

  - ViewPoint Cloud Software User Manual

  - ViewPoint Cloud v1.1.2 Compliance Certification

  - VP-00133 Installation and Operational Qualification Protocol for the ViewPoint Cloud Continuous Monitoring Systems

  - VP-00153 Performance Qualification Protocol for the ViewPoint Cloud Continuous Monitoring System

  - VPx Sensor User Manual

- DEFINITIONS

  - Acceptance Criteria: Predetermined requirements upon which pass/fail decisions are based. Acceptance criteria are specific and measurable.

  - Critical to Quality: Attribute of a part, assembly, sub-assembly, product, or process that has direct and significant impact on its actual or perceived quality to the customer or end user.

  - Chart recorder: Equipment used to monitor temperature and/or humidity in a specific area by recording information onto a paper disc or paper strip.

  - Controlled Environment: Defined zone in which sources of contamination are controlled by specified means.

  - Deviation:  The action of departing from an established method, procedure, or acceptance criteria.

  - Environmental Control: Activities performed to control the environment for the purpose of minimizing the potential for contamination and cross-contamination of tissue.

  - Environmental Monitoring: Activities performed to systematically observe and record data to characterize the environment to identify conditions under which the potential may exist for contamination or cross-contamination of tissue.

  - HCT/Ps: Human Cells, Tissues, and Cellular and Tissue-Based Products.

  - In-process Controls: Any tests, samples, evaluations, monitoring, or measurements performed during processing or preservation that are designed to evaluate the processing or preservation procedure of the tissue subjected to processing or preservation for conformance to specification(s).

  - Installation Qualification (IQ): Establishing by objective evidence that all key aspects of the process equipment and ancillary system installation adhere to the manufacturer’s approved specification(s).

  - Management with Executive Responsibility (a.k.a. Management with Executive Responsibility/Management Representative/Head of Quality/Tissue Bank Director): Those senior employees of an establishment who have the authority to establish or make changes to the establishment’s quality policy and quality system.

  - Objective Evidence: Information, records, or statements of fact which are based on observation, measurement, or tests that can be verified.

  - Operational Qualification (OQ): Establishing by objective evidence the process control limits and action levels which result in a product that meets predetermined requirements and are within the operating limits required by the process.

  - Off-the-Shelf Software (OTS): Applications developed by a vendor for general installations. These applications may or may not be validated by the vendor for release. At a minimum, a Risk Assessment should be done to determine if the program requires additional validation testing. However, any customization of these applications may require full validation and change control based on criticality and complexity of the customization/configuration (e.g., spreadsheet macros, changes to default parameters).

  - Performance/Product Qualification (PQ): Establishing by objective evidence that the process, under anticipated conditions, consistently (repeatedly and reliably) produces a product which meets predetermined requirements.

  - Preventative Maintenance: Maintenance that is regularly performed on a piece of equipment to lessen the likelihood of it failing. Preventative maintenance is performed while the equipment is still working, so that it does not break down unexpectedly.

  - Preventive Action: Action to eliminate the cause of a potential nonconformity or other undesirable situation in order to prevent occurrence; may be performed in conjunction with corrective action(s).

  - Process Validation: Establishing by objective evidence that the output of a process consistently meets predetermined specification.

  - Protocol: A pre-approved and prospective document that states the requirements and procedures for conducting a validation.

  - Subject Matter Expert (SME): These are individuals who provide authority in a particular area. These individuals will respond to investigator’s questions pertaining to specific areas when requested. The head of each department, or his/her designee, shall serve as the subject matter expert for the department. Each expert must understand the expectations associated with participation and availability during a regulatory inspection.

  - Validation/Study Leader: The individual who has the responsibility and authority over the operation of the equipment or process being validated.

  - Verification: Confirmation by examination with a provision of documented evidence that a specific requirement has been fulfilled.

- RESPONSIBILITIES

| Function | Name | Title | Develop, Author, and Pre-Approve VP-00133 and VP-00153 | Execute, Review and Approve VP-00133 and VP-00153 | Author Stimlabs Protocols and Reports | Execute Stimlabs Protocols | Review and Approve Stimlabs Protocols and Reports |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Validation/Study Leader | Alexander Waters | Quality Improvement Engineer |  | X* | X | X** | X |
| Subject Matter Expert | Mesa Labs | VPC CMS Manufacturer/Vendor Personnel | X | X |  |  |  |
| Management with Executive Responsibility | Melissa O’ Connor | Chief Quality & Regulatory Officer |  |  |  |  | X |
| Operations Management | Joel Cummings | Vice President of Operations |  |  |  |  | X |
| Quality Management | Justin Junnier | Continuous Improvement Manager |  |  |  |  | X |

*Stimlabs personnel will assist during the execution of the pre-written protocol, VP-00133, in resolution of observations and deviations, access to secure areas, and support for the overall completion of the qualification effort as detailed in the protocol. Objective evidence of training of utilized personnel will be attached.

**Stimlabs personnel will execute all elements of Stimlabs’ protocols. Objective evidence of training of utilized personnel will be attached.

- ENVIRONMENTAL REQUIREMENTS

  - No environmental requirements are necessary to conduct OQ protocol testing.

- TEST METHODS

  - The IQ and OQ portions of this validation were pre-written by MLI in one contiguous IOQ protocol (VP-00133 Installation and Operational Qualification Protocol for the ViewPoint Cloud Continuous Monitoring Systems). To follow Stimlabs’ validation procedure, QSP020-001 Process and Equipment Validation, IQ and OQ portions of the contiguous protocol will be evaluated individually for Stimlabs’ reporting purposes. Test attachments pertaining to specific phases of validation will be considered in the documentation for their respective portions and test attachments pertaining to both the IQ and OQ phases of validation will be considered in the documentation for both. Subsequently, all test attachments of the IOQ protocol will be executed sequentially and performed as written by the manufacturer. The following outlines the components of this validation that the OQ will be comprised of:

    - Operational Qualification (OQ): VP-00133 has been pre-written by MLI and will be executed by MLI with assistance from Stimlabs personnel. Refer to the Appendix section for protocol/data collection sheets that pertain to this OQ. The test methods performed by MLI pertinent to this OQ are listed below.

      - Protocol Execution Signature List

      - System Security – Account Access

      - Sensor Alarm Testing

      - Communication Failure Testing

      - Auto Report Testing

      - Documentation Generation

      - Additional Information

      - Protocol Execution Summary and Acceptance

      - Deviations and Corrective Action Form Log

      - Deviation and Corrective Action Form

    - No additional testing or data collection will be outlined or conducted according to this OQ protocol as sufficient testing to satisfy Stimlabs’ requirements will be performed during execution of the OQ portions of the IOQ outlined in VP-00133.

- PROCEDURE

  - Refer to VP-00133 in the Appendix. The procedure for this phase of the validation is outlined in VP-00133 and will be executed by the manufacturer as written by the manufacturer with assistance, as needed, from Stimlabs.

- LOT SIZE & SAMPLING PLAN

  - All functionalities of the VPC CMS required by Stimlabs will be subjected to test – a sampling plan or statistical rationale is not applicable.

- ACCEPTANCE CRITERIA

  - Refer to VP-00133 in the Appendix section. The acceptance criteria for this phase of the validation is outlined in VP-00133 and will be evaluated and documented by MLI and Stimlabs therein.

- MATERIAL DISPOSITION

  - There is no material disposition associated with this validation.

- DOCUMENTATION REQUIREMENTS

  - Results shall be summarized in an OQ report.

  - Various Control Plans, such as validated process settings and re-validation criteria shall be documented in an OQ report.

  - Raw data sheets shall be included in the Appendix of the OQ report and/or a reference to the location of the raw data shall be included in the OQ report.

  - Calibration records (if applicable) shall be attached to the OQ report (in the Appendix).

  - Signed OQ Protocol Training forms will be attached to the OQ report (in the Appendix) if applicable.

- DEVIATIONS

  - Deviations to the OQ will be pre-approved using FRM020-012, Protocol Deviation Form, in addition to being documented in the Deviations and Corrective Action Form Logs and Deviation and Corrective Action Forms of the protocol provided by the manufacturer.

  - Copies of signed Protocol Deviation Forms will be attached to the OQ report (in the Appendix).

- PRE-REQUISITE STEPS

  - Approved Process Validation Plan (VAP22-006-00).

  - The IQ Protocol, VAP22-006-01, must be approved before validation activities begin.

  - This OQ Protocol must be approved before validation activities may begin.

  - VP-00133 must be approved before validation activities may begin.

  - Completed operator training if applicable.

  - Work Instructions must be drafted.

- APPENDIX

  - ViewPoint Cloud Software User Manual

  - ViewPoint Cloud v1.1.2 Compliance Certification

  - VP-00133 Installation and Operational Qualification Protocol for the ViewPoint Cloud Continuous Monitoring Systems

  - VP-00153 Performance Qualification Protocol for the ViewPoint Cloud Continuous Monitoring System

  - VPx Sensor User Manual

  - Training Forms (if applicable)

- REVISION HISTORY

| Revision Number | Description of Change | Changed By |
| --- | --- | --- |
| 01 | Original Issue | Alexander Waters |