Pre-Execution Approvals

This document has been reviewed and approved by the following individuals.  By approving this document, each individual acknowledges that it complies with the requirements and spirit of the Stimlabs Process and Equipment Validation Procedure QSP020-001.

| Sean Miller Stimlabs Validation/Study Leader |  | Date |  | John Daniel Stimlabs Tissue Bank Director |  | Date |
| --- | --- | --- | --- | --- | --- | --- |
| Melissa Gaynor Stimlabs Head of Quality |  | Date |  | Matt Nystrom Stimlabs Operations Management |  | Date |
| Mike Corby WuXi AppTec Senior Technical Director |  | Date |  | Richard Nardini WuXi AppTec Quality Assurance Manager |  | Date |

1.0	PURPOSE

The purpose of this performance qualification (PQ) is to define the process of qualifying moist heat sterilization processes on items that are used in routine laboratory production, testing, and quality control activities to specific sterility assurance levels (SALs).

2.0	SCOPE

2.1	This PQ protocol applies to the following items, processes, and sterilization facility:

2.1.1	Items:	Laboratory production instruments, testing supplies, and quality control cleaning articles, which are separated into the following two (2) sterilizer load configurations:

2.1.1.1	Load Configuration MDL1 primarily consists of items such as metal instruments, silicon and metal testing supplies, quality control cleaning supplies, etc.  Load Configuration MDL1 is sterilized at a set point temperature of 132 °C.

2.1.1.2	Load Configuration PDL1 primarily consists of items such as polypropylene testing supplies, pouches that are constructed with polyester-based flexible film and spunbonded olefin backing, etc.  Load Configuration PDL1 is sterilized at a temperature set point of 121 °C.

2.1.2	Process:	The PQ consists of the following processes using a Steris steam sterilizer; model number LV250, serial number 0305716-26:

2.1.2.1	Three (3) empty chamber sterilizer profiles/mappings.

2.1.2.2	Three (3) heat penetration study cycle runs on a maximum MDL1 load configurations.

2.1.2.3	Three (3) heat penetration study cycle runs on a minimum MDL1 load configurations.

2.1.2.4	Three (3) heat penetration study cycle runs on a maximum PDL1 load configurations.

2.1.2.5	Three (3) heat penetration study cycle runs on a minimum PDL1 load configurations.

2.1.2.6	Three (3) microbiological challenge study cycle runs on the worst-case MDL1 load configuration.

2.0	SCOPE (Continued)

2.1.2.7	Three (3) microbiological challenge study cycle runs on the worst-case PDL1 load configuration.

Note:	In microbiological challenge study cycle runs, the worst-case load configuration is the one that achieved the least amount of lethality during heat penetration study cycle runs.

2.1.3	Sterilization Facility:	The validation is performed at Stimlabs LLC, which is located at 1225 Northmeadow Parkway, Suite 104, Roswell, GA 30076.

2.2	The execution of the PQ protocol is performed by WuXi AppTec (WXAT), which is an independent pharmaceutical, biopharmaceutical, and medical device outsourcing service company located at 1265 Kennestone Circle, Marietta, GA 30066.

3.0	OBJECTIVE

3.1	Sterilizer Empty Chamber Profile/Mapping

3.1.1	Empty chamber temperature profile/mapping study cycle runs are performed in accordance with British Standard EN 554:1994, “Sterilization of medical devices – Validation and routine control of sterilization by moist heat” and Parenteral Drug Association, Inc., Technical Report No. 1, “Validation of Moist Heat Sterilization Processes: Cycle Design, Development, Qualification and Ongoing Control, Revised 2007” to verify proper temperature uniformity/distribution of the heating medium across the sterilizer’s loading zone.  Empty chamber temperature profile/mapping study cycle runs are performed to check the acceptability of the following during the sterilization phase of the cycle run:

3.1.1.1	The variation in temperature of the sterilizer loading zone as measured by each temperature sensing probe.

3.1.1.2	The variation in temperature of the sterilizer loading zone as measured between all temperature sensing probes.

3.1.1.3	The difference in temperature of the sterilizer loading zone as compared between the sterilizer temperature set point and the temperature sensing probes

3.1.2	The empty chamber temperature profile/mapping uses a pre-defined number of calibrated temperature sensors at specified locations using a pre-vacuum sterilization cycle and operating temperature settings of 121 C and 132 C, which are routinely used for production sterilization processes.

3.0	OBJECTIVE (Continued)

3.1.3	The empty chamber temperature profile/mapping study cycle runs also check the sterilizer’s loading zone for any areas that consistently demonstrate cold temperatures relative to the rest of the sterilizer’s loading zone.

3.1.3.1	A consistently cold area is considered a “cold spot.”

3.1.3.2	If a cold spot is identified, the area is challenged in microbiological challenge study cycle runs.

3.1.3.3	The microbiological challenge study cycle runs confirm if the cold spot is getting the minimum required lethality to achieve the necessary sterility assurance level of 10-6.  If the minimum required lethality is achieved, then the cold spot is determined to have no negative impact to sterilization processes.

3.2	Physical and Microbiological PQ

3.2.1	A physical and microbiological PQ are performed using reference load configurations MDL1 and PDL1, which are processed using a Steris steam sterilizer, model number LV250, serial number 0305716-26.

3.2.2	The physical PQ includes four (4) series of at least three (3) exposures of items through the sterilization process.

3.2.2.1	Two (2) triplicate exposure series of maximum load configurations MDL1 and PDL1.

3.2.2.2	Two (2) triplicate exposure series of minimum load configurations MDL1 and PDL1.

3.2.3	The objective of the physical PQ is to demonstrate the following:

3.2.3.1	Through heat penetration study cycle runs, establish a lower boundary for the minimum number of load configuration MDL1 and PDL1 items that can be sterilized within the sterilizer during routine sterilization processes in order to achieve a known minimum SAL of 10-6.

3.2.3.2	Through heat penetration study cycle runs, establish an upper boundary for the maximum number of load configuration MDL1 and PDL1 items that can be sterilized within the sterilizer during routine sterilization processes in order to achieve a known minimum SAL of 10-6.

3.0	OBJECTIVE (Continued)

3.2.3.3	The demonstration of the acceptability of different loading configurations.

3.2.3.4	The determination of whether a minimum or maximum load configuration represents the worst-case load configuration for performing microbiological challenge study cycle runs.

3.2.3.5	Determination of the half-sterilization cycle time

3.2.3.6	Determination of the full-sterilization cycle time

3.2.3.7	Determination of adequate dry times

3.2.4	The results from the physical PQ are used in performing microbiological challenge study cycle runs.

3.2.5	The microbiological PQ includes two (2) series of at least three (3) exposures each of items through the sterilization process at half-sterilization cycle time.

3.2.5.1	One (1) triplicate exposure of the worst-case MDL1 load configuration that was determined through the heat penetration study cycle runs.

3.2.5.2	One (1) triplicate exposure of the worst-case PDL1 load configuration that was determined through the heat penetration study cycle runs.

Note:	The worst-case load configuration is one that accumulates the least amount of lethality from data shown from the heat penetration study cycle runs.

3.2.6	The objective of the microbiological PQ is to demonstrate the following:

3.2.6.1	Confirm achieving a minimum SAL of 10-3 at half-sterilization cycle times and a minimum of 10-6 at full-sterilization cycle times.

3.2.6.2	Confirm achieving a minimum total accumulated lethality or spore log reduction (SLR) of nine (9) logs at half-sterilization cycle times.

      - All microbiological challenge study cycle run biological indicators (BIs) test negative for growth.

3.2.6.4	Positive control BIs test positive for growth.

3.0	OBJECTIVE (Continued)

3.2.6.5	The sterilization process consistently delivers appropriate microbial inactivation over worst-case conditions not normally seen in routine sterilization processes.

3.2.6.6	Establish baseline temperature differences between the temperature measured at the reference measuring point (sterilizer drain) and the temperature values displayed by the sterilizer temperature controller.

  - Leak and Daily Air Removal Test (DART)

3.3.1	The sterilization process makes use of a vacuum.  Therefore, an air leakage test is carried out at specified intervals.

3.3.2	The sterilization process relies on the removal of air from the sterilizer chamber in order to achieve rapid and even penetration of steam into the sterilizer load.  Therefore, a steam penetration test is carried out each day before the sterilizer is used.

3.3.3	The steam penetration test is carried out using a device having a defined challenge to air removal and steam penetration for the process.

  - Any tests that do not have results which support the parameters defined in this approved PQ protocol are conclusively rationalized for their nonconformance and approved or the PQ is considered invalid.

- EXPERIMENTAL DESIGN

4.1	Validation Process Summary

4.1.1	All validation and qualification activities are performed in accordance with process validation plan VAP16-010-00, which is in harmony with international standard ISO 17665-1:2006.

4.1.2	An installation qualification (IQ) is performed to demonstrate that the sterilization system has been installed in accordance with established specifications.  The demonstration includes compliance with equipment manufacturer’s specifications after installation and conformance of services/utilities to specified requirements.

4.0	EXPERIMENTAL DESIGN (Continued)

4.1.3	An operational qualification (OQ) is performed to demonstrate that the installed sterilization system, including controls and recording systems, devices, and control programs operate according to defined manufacturer’s performance specifications and user requirements.

4.1.3.1	The confirmation of the relationship between set control values and measured process parameters.

4.1.3.2	The confirmation of recording systems and display devices meeting manufacturer’s performance specifications.

4.1.4	A physical and microbiological PQ are performed with reference load configurations within the sterilization system for the introduction of sterilized laboratory production instruments, testing supplies, and quality control cleaning articles utilized for routine testing and manufacturing processes.

4.2	Sterilization Method

4.2.1	The sterilization method used for the PQ is the “Conservative Process Definition Based on Inactivation of Reference Microorganisms (Overkill Method).”

4.2.2	This process is widely employed and is often used to sterilize re-usable items.  Qualifying a sterilization process for such items requires an approach different from that used for virgin items, because the challenge to the sterilization process is difficult to define and pre-sterilization treatments (such as cleaning) are difficult to validate and control.

4.2.3	A sterilization process in this situation is often conservative and designed to deliver a treatment exceeding that required to achieve the specified requirements for sterility.

4.2.4	This type of treatment can be determined either mathematically based on an empirical microorganism (full cycle approach) or from a reduced level of treatment delivered to reference microorganisms (partial cycle approach).

4.2.4.1	The approach used in this PQ is the partial cycle approach.

4.2.4.2	The sterilization load is exposed to the sterilizing agent under conditions designed to deliver a reduced level of treatment (Half-sterilization cycle time).

4.0	EXPERIMENTAL DESIGN (Continued)

4.2.4.3	The extent of treatment needed to inactivate the equivalent of 106 microorganisms on a biological indicator that complies with ISO 11138-3, “Sterilization of health care products — Biological indicators — Part 3: Biological indicators for moist heat sterilization processes” is confirmed.

4.2.4.4	The level of treatment identified is carried out in triplicate to demonstrate reproducibility.

4.2.4.5	If the inactivation of the equivalent of 106 viable microorganisms is confirmed, then the extent of treatment for the sterilization process is determined by extrapolation to a predicted probability of survival of 10−6 or better.  This is achieved by taking into account the nature of the inactivation kinetics of the sterilizing agent and the number and resistance of the microorganisms on the BI.

4.2.5	The overkill method is best suited to sterilization processes where linear inactivation kinetics is demonstrated.

4.2.6	The location within the load at which sterility is most difficult to achieve is established.

4.2.7	A challenge to the sterilization process is created by:

4.2.7.1	Placing reference microorganisms (BIs) within items a position or representative positions where sterilizing conditions are most difficult to achieve.

4.2.7.2	Inoculating moveable and porous items with reference microorganisms at positions where sterilizing conditions are most difficult to achieve.

4.2.8	The challenge is packaged in the same manner as routinely produced items and included within the sterilization load at a location where it is most difficult to achieve sterilizing conditions.

4.3	Sterilizing Agent and Delivery System

4.3.1	The sterilizing agent and delivery system used in this PQ is saturated steam with forced air removal by autoclave sterilization.

4.3.2	The sterilizing agent is primarily used for items where air is difficult to remove such as in porous materials, lumens and/or cavities, and packaged items.  It is also used for surface contact sterilization.

4.0	EXPERIMENTAL DESIGN (Continued)

4.3.3	Moisture in the steam has a major impact on its ability to denature or coagulate proteins of microorganisms; hence the decision to use saturated steam.

4.3.4	The sterilization process consists of six major phases:

4.3.4.1	An air removal phase where air is removed from the sterilizer chamber and the sterilization load by either a deep vacuum or a combination of vacuum and steam pulses above and/or below atmospheric pressure.

4.3.4.2	A charge phase where saturated steam enters the sterilizer chamber until the sterilization temperature and pressure are attained in the sterilizer chamber.

4.3.4.3	A plateau period phase (Sterilization phase) where the sterilization temperature and pressure are maintained in the sterilizer chamber by saturated steam for the specified time.

4.3.4.4	An exhaust phase where steam is exhausted from the sterilizer chamber, and a pre-determined level of vacuum is obtained.

4.3.4.5	A drying phase for items required to be dry where the temperature in the jacket surrounding the sterilizer chamber and the vacuum in the sterilizer chamber are both maintained for a pre-determined period.

4.3.4.6	A vacuum relief phase where air is admitted to the sterilizer chamber through a microbiologically-retentive filter until atmospheric pressure is reached.

- TESTING METHODOLOGY RATIONALE

  - Empty Chamber Sterilizer Profiles/Mappings

    - Empty chamber temperature profile/mapping study cycle runs are performed in accordance with the method specified in British Standard EN 554:1994 and ISO 17665-1:2006.

    - The rationale for the empty chamber temperature profile/mapping study cycle run temperature variation specifications is based on British Standard EN554:1994.

5.0	TESTING METHODOLOGY RATIONALE (Continued)

    - The rationale for the number of calibrated temperature sensors utilized in the empty chamber temperature profile/mapping study cycle runs is based on British Standard EN554:1994, which states that experience has shown that twelve (12) sensors per cubic meter are normally sufficient.

    - Standard ANSI/AAMI/ISO 20857:2010, “Sterilization of health care products – Dry heat – Requirements for the development, validation and routine control of a sterilization process for medical devices” also states that for small batch sterilizers (those having a chamber volume less than or equal to 10 cubic feet or 0.283 cubic meters), a minimum of ten (10) temperature sensors should be used.  For larger batch sterilizers (those having a chamber volume more than 10 cubic feet), a minimum of twenty (20) temperature sensors should be used.

    - The chamber volume of the Steris LV250 steam sterilizer is 0.25 cubic meters, Therefore, four (4) sensors for the empty chamber temperature profile/mapping study cycle runs could have been chosen as the minimum number of sensors for the study.  Taking a more conservative approach, Stimlabs uses a minimum of fourteen (14) calibrated temperature sensors to ensure that at least twelve (12) sensors are functional at all times.

Note:	British Standard EN554:1994 was superseded by ISO 17665-1:2006.  The rationale for the number of sensors used for empty chamber temperature profiles was excluded from ISO 17665-1:2006.  Therefore, the rationale for the number of temperature sensors utilized for empty chamber temperature profile studies will continue to reference British Standard EN554:1994.

5.2	Heat Penetration Study Cycle Runs

5.2.1	The rationale for heat penetration study cycle runs is in accordance with British Standard EN 554:1994 and ISO 17665-1:2006.

5.2.2	The number of temperature sensors used is not less than the number used for the empty chamber temperature profile/mapping study cycle runs with at least one sensor situated adjacent to the temperature sensor connected to the sterilizer’s recording/indicating instrumentation.

5.2.3	Reproducibility within acceptable limits is checked using a minimum of three (3) replicate cycles.

5.3	Microbiological Challenge Study Cycle Runs

5.3.1	The rationale for microbiological challenge study cycle runs is in accordance with British Standard EN 554:1994 and ISO 17665-1:2006.

5.0	TESTING METHODOLOGY RATIONALE (Continued)

5.3.2	In addition to the measurement of physical parameters established from heat penetration study cycle runs, the sterilization process is verified by microbiological methods or BIs that are positioned in and/or on items at specified locations and then exposed to an overkill process.

5.3.3	Items used in establishing the sterilization process is a representative of that to be produced routinely.

5.3.4	The equipment used is capable of reproducibly delivering a combination of process parameters with less lethality than the sterilization process, such that the level of inactivation of microorganisms results in a population that allows interpolation with regard to the total result.

5.3.5	The microbiological challenge study cycle runs use process challenge devices (PCDs) as worst-case challenge items within the sterilization process.

Note:	For the purposes of this PQ, the terms BI and PCD are synonymous.

5.3.6	PCDs/BIs represent the worst-case bioburden condition as compared to the items natural bioburden.

5.4	PCDs/BIs

5.4.1	The rationale for PCD/BI use in microbiological challenge study cycle runs is in accordance with ANSI/AAMI/ISO 20857:2010 and ISO 17665-1:2006.

5.4.2	The rationale for the number of BIs utilized for sterilization processes is based on ANSI/AAMI/ISO 20857:2010.  The standard states that for small batch sterilizers (those having a chamber volume less than or equal to 10 cubic feet or 0.283 cubic meters), a minimum of ten (10) BIs should be used.  For larger batch sterilizers (those having a chamber volume more than 10 cubic feet), a minimum of twenty (20) BIs should be used.  The chamber volume of the Steris LV250 steam sterilizer is 0.25 cubic meters.  Therefore, taking a more conservative approach, Stimlabs uses a minimum PCD/BI combination of twenty (20) for each microbiological challenge study cycle run.

5.4.3	ISO 17665-1:2006 suggests using microorganisms that demonstrate high resistance to moist heat.  The microorganisms which are suitable for use are Geobacillus stearothermophilus, Bacillus coagulans, Clostridium sporogenes, or Bacillus atrophaeus.

5.0	TESTING METHODOLOGY RATIONALE (Continued)

5.5	Triplicate Testing

5.5.1	The rationale for triplicate testing is in harmony with British Standard EN554:1994 and ISO 17665-1:2006

5.5.2	Empty chamber sterilizer profile/mapping study cycle runs, heat penetration study cycle runs, and microbiological challenge study cycle runs are performed in triplicate in order to show the following:

5.5.2.1	Repeatability and reproducibility in the testing process;

5.5.2.2	Rule out experimental bias or random error;

5.5.2.3	Establish the minimum number of tests needed to determine a standard deviation.

5.6	Leak and DART

The rationale for Leak and DART testing is in accordance with British Standard EN 554:1994 and ISO 17665-1:2006.

5.7	Sterilization Wrap

5.7.1	In this PQ, items requiring sterilization wrap are double wrapped as it provides a greater sterilization process challenge than single sterilization wrapped items as double sterilization wrapped items provide more opposition to steam penetration.

5.7.2	In accordance with a shelf life study(1) conducted on single and double sterilization wrapped items, single sterilization wrapped items did not carry any greater risk of bacterial contamination than double sterilization wrapped items and may lead to significant cost savings in both labor (time to wrap) and consumables (linen and recycling costs).

5.7.3	In routine sterilization processing, Stimlabs will have the option of sterilizing items using a single sheet of sterilization wrap.

(1)Webster, Joan and Radke, Elinor and George, Narelle and Faoagali, Joan and Harris, Mary (2005) Barrier properties and cost implications of a single versus a double wrap for storing sterile instrument packs.  American Journal of Infection Control 33(6): pp. 348-352.

6.0	DEFINITIONS

6.1	Aseptic technique:  Conditions and procedures used to exclude the introduction of microbial contamination.

6.2	Bioburden:  Populations of viable organisms on an item and/or package.

6.3	Bioburden estimate:  Value established for the bioburden by applying a factor to a pre-sterilization count to compensate for the efficiency of the defined technique used in recovery of microorganisms.

6.4	Biological Indicator (BI):  A carrier or suspension of a specific species of microorganism (bacteriologic spores) with a high resistance to a particular mode of sterilization being monitored.  It is contained in a package or delivery device designed to maintain the integrity of the inoculated carrier.

6.5	Cold Spot: Location consistently slowest to reach the sterilizer’s control set temperature or consistently accumulates the least amount of lethality.

6.6	Control Set Temperature:  Determined temperature that serves as the operating reference for the sterilizer control system so that the chamber temperature remains within the required range around the selected sterilization exposure temperature.

6.7	Critical Parameters:  Quantifiable parameters that influence the efficacy of a sterilization process.

6.8	Cycle:  Defined sequence of operational steps designed to achieve sterilization and carried out in a sealed chamber.

6.9	Deviation:  The action of departing from an established method, procedure, or acceptance criteria.

6.10	Discrepancy:  The state or quality of being discrepant; different; or inconsistent.

6.11	D-value:  The time, at a given temperature, needed to reduce the number of microorganisms by 90%, or one logarithm and is expressed in terms of time (minutes).

6.12	Exposure:  The period during a sterilization process (or cycle) in which an item is exposed to the sterilant at specified sterilization parameters.

6.13	Exposure or Holding Phase:  The sterilization cycle phase during which the measured quantities (e.g. temperature and time) are maintained within specified requirements that are needed to achieve the desired lethality.

6.0	DEFINITIONS (Continued)

6.14	Exposure Time:  Period for which the process parameters are maintained within their specified tolerances.

6.15	FBiological Value (FBIO):  The delivered lethality measured by the actual kill of microorganisms on or in a BI system.  FBIO = DT X LR, where:  DT = D-value of the BI at the reference temperature and LR = Log reduction of the BI population achieved during the cycle.

6.16	Final Summary Report:  A document summarizing the results derived from the execution of a protocol or procedure.  The final report includes a conclusion, which indicates protocol or procedure success or failure in accordance with established acceptable criteria for all critical process parameters as determined by Technical and Quality Management.

6.17	F0 Value:  Microbiological lethality of a sterilization process expressed in terms of the equivalent time, in minutes, at a temperature of 121.1 C with reference to microorganisms with a Z-value of 10 C.  One (1) minute is equivalent to one logarithm (base 10) reduction of microorganisms.

6.18	F-value (Lethality Factor):  A measure of sterilization effectiveness.  It is the calculated equivalent lethality in terms of minutes at a specific reference temperature using a Z-value other than 10 C.

6.19	FPhysical Value (FPHY):  F-values calculated from physical data (time and temperature).

6.20	Fractional Cycle:  Sterilization cycle in which exposure to the sterilizing agent is abbreviated.

6.21	Full-cycle:  Specified lethality exposure time that is at least double half-cycle times.

6.22	Half-cycle:  Specified minimum lethality time where a total of three (3) consecutive experiments result in total inactivation of BIs.

6.23	Heat Penetration:  Cycles performed to determine the actual temperature reached within the load.

6.24	Load Configuration:  Totality of attributes defining the presentation of the item through the sterilization process.  This configuration includes 1) the orientation of the item within the primary package; 2) the quantity and orientation of the primary packages within the secondary and tertiary package.

6.25	Log Reduction (LR):  Reduction in the number of viable microorganisms, expressed in log10 units, after exposure to a sterilization cycle.

6.0	DEFINITIONS (Continued)

6.26	Mapping:  A process of determining the differences and changes of conditions in a defined space under normal operating conditions.

6.27	Measurement Uncertainty:  An estimate of the range of values that the true value of measurement is within, with a specified level of confidence.

6.28	Microbiological Challenge:  Systems used in cycle development, validation, and routine monitoring, which demonstrate the microbial lethality of the sterilization process.

6.29	PCD:  Item designed to simulate a production item to be sterilized and to constitute a defined challenge to the sterilization process, and used to assess the effective performance of the process.

6.30	PCL:  Site that simulates “worst case” conditions as they are given for sterilizing agent(s) in the goods to be sterilized.

Note:	1) The site is so constituted that a biological indicator can be placed in the position that represents a rigorous challenge for the sterilizing agent to reach; 2) the site depends on the type of goods to be sterilized and the sterilization process parameters; 3) the biological indicator should not interfere with the function of the goods; and 4) in some sites an inoculated carrier may be used in place of a biological indicator.

6.31	Qualification:  Describes the testing and review of a piece of equipment, system, or sub-system of a process to assure its fitness for use.  Qualification deals with components or elements of a process, while validation deals with the entire process.

6.32	Spore Log Reduction:  Lethality observed in a full or fractional sterilization cycle.

6.33	Sterility Assurance Level (SAL):  The likelihood of something that has undergone a sterilization procedure being non-sterile (expressed as 1 x 10-x or probability of survival).

6.34	Sterility Testing:  Test performed to determine if viable microorganisms are present.

6.35	Sterilization:  Validated process used to render an item free from viable microorganisms.

Note:	In a sterilization process, the nature of microbial inactivation is described by an exponential function.  Therefore, the presence of a viable microorganism on any individual item can be expressed in terms of probability.  While this probability can be reduced to a very low number, it can never be reduced to zero.

6.36	Sterilizer:  Apparatus used to sterilize medical devices, equipment, and supplies by direct exposure to the sterilizing agent.

6.0	DEFINITIONS (Continued)

6.37	Temperature Coefficient (Z-value):  Temperature change required to cause a one logarithm decrease in the D-value of a microorganism and is expressed in degrees Celsius.

6.38	Validation:  The overall establishment of documented evidence through defined tests and challenges, that a system, manufacturing process, analytical method, and/or piece of equipment meets design criteria and that adequate provision have been established to keep it in a state of control.  Validation typically covers three (3) activities such as commissioning, verification of process specification, and performance qualification.

7.0	RESPONSIBILITIES

The following tasks are assigned with respect to the PQ process.

| Task | Stimlabs | WXAT |
| --- | --- | --- |
| Provide technical manuals for the PQ as needed | √ |  |
| Provide the operating performance specifications | √ |  |
| Write the PQ protocol |  | √ |
| Review and approve the PQ protocol | √ | √ |
| Coordinate the PQ activities |  | √ |
| Execute the PQ as specified within the protocol |  | √ |
| Collect all the PQ information/data |  | √ |
| Compile and analyze the PQ information/data |  | √ |
| Review and approve the executed PQ protocol | √ | √ |
| Write the PQ final summary report |  | √ |
| Review and approve the PQ final summary report | √ | √ |

8.0	EQUIPMENT AND MATERIAL REQUIREMENTS

8.1	One (1) calibrated Steris steam sterilizer, model number LV250, serial number 0305716-26

8.2	One (1) integral steam generator capable of producing high-pressure steam (90 psig maximum) for steam sterilizers utilizing sterilization temperatures up to 140 °C.

8.3	One (1) calibrated and validated temperature measurement system having fourteen (14) sensors with a system accuracy and measurement uncertainty of ≤ ±0.5 °C.

8.4	Four (4) sterilizer racks.

8.5	One (1) maximum MDL1 load configuration (See Section 15.0).

8.6	One (1) minimum MDL1 load configuration (See Section 15.0).

8.7	One (1) maximum PDL1 load configuration (See Section 15.0).

8.8	One (1) minimum PDL1 load configuration (See Section 15.0).

8.9	PCDs/BIs

8.9.1	A minimum of twenty-four (24) PCDs inoculated with Geobacillus stearothermophilus BIs having a certified microbial population and D-value meeting the requirements of Table 1.

8.9.1.1	PCDs consist of BI inoculated laboratory production instruments, testing supplies, and quality control cleaning articles that are constructed with porous material or are pivoted hand tools that allow movement by in-hand manipulation.  PCDs represent the most challenging items in the sterilization process.

Note:	A Bacteriostasis and Fungistasis (B/F) test is performed on single item type PCDs in order to assess whether or not the item type being sterilized inhibits the growth of microorganisms.  The performance of B/F testing is necessary in order to validate the sterility result and to confirm that there are no antimicrobial properties within the item that would inhibit the detection of microbial organisms during the sterility assay.  B/F test results are included in the final summary report.

8.9.1.2	Eighteen (18) PCDs are used in three (3) microbiological challenge study cycle runs.

8.0	EQUIPMENT AND MATERIAL REQUIREMENTS (Continued)

8.9.1.3	Six (6) PCDs are used as positive controls

8.9.2	A minimum of one hundred and four (104) conventional paper strip BIs inoculated with Geobacillus stearothermophilus having a certified microbial population and D-value meeting the requirements of Table 1.

8.9.2.1	One hundred and two (102) conventional paper strip BIs are used in six (6) sterilization process.

8.9.2.2	Two (2) conventional paper strip BIs are used as positive controls.

8.9.3	Confirmation of the BI population is required in accordance with WXAT standard operating procedure MA225SOP, “Total Viable Spore Count – Biological Indicators.”

8.9.4	The purpose of confirmation is to show that the BI population meets the manufacturer’s label claim according to United States Pharmacopeia (USP) current version, chapter <55>, “Biological Indicators – Resistance Performance Tests, Total Viable Spore Count.”

8.9.4.1	The BI population is confirmed to be within -0.3 log and +0.48 log of the manufacturer’s certified label spore count claim in accordance with USP monograph, “Biological Indicators for Moist Heat, Dry Heat, and Gaseous Modes of Sterilization, Nonpaper Carriers.”

8.9.4.2	D-values are accepted based on the manufacturer’s certificate of analysis (provided the supplier complies with the requirements of ANSI/AAMI/ISO 11138 guidelines and USP current version, chapter <1035>, “Biological Indicators for Sterilization,” as shown in Table 1.

Table 1 – Typical Characteristics for Commercially Supplied BI Systems

| Sterilization Mode | Example of a Typical D value (minutes) | Range of D values for Selecting a Suitable BI (minutes) | Limits for a Suitable Resistance (depending on the particular D value [minutes]) |  |
| --- | --- | --- | --- | --- |
|  |  |  | Survival Time | Kill Time |
| Moist Heat(1) at 121 °C or 132 °C | 1.9 | Min. 1.5 | Min. 4.5 | 13.5 |
|  |  | Max. 3.0 | Max. 14.0 | 32.0 |

(1) For 1.0 x 105 to 5.0 x 106 spores per carrier.

8.10	One (1) WXAT approved pen for documenting information required by the PQ protocol.

8.0	EQUIPMENT AND MATERIAL REQUIREMENTS (Continued)

8.11	Insulated gloves for placing and removing temperature sensors to and from the sterilizer.

8.12	100% polypropylene regular weight 1 ply sterilization wrap (i.e., Medline GEM2124S or equivalent).

8.13	Self-Seal 3.5" X 9" Sterilization Pouch (i.e., Medline MPP100520GS or equivalent).

8.14	Self-Seal 5.25" X 10" Sterilization Pouch (i.e., Medline MPP100525GS or equivalent).

8.15	Self-Seal 7.5" X 13" Sterilization Pouch (i.e., Medline MPP100545GS or equivalent).

8.16	Self-Seal 12" X 15" Sterilization Pouch (i.e., Medline MPP100555GS or equivalent).

8.17	Self-Seal 12" X 18" Sterilization Pouch (i.e., Medline MP100558GS or equivalent).

8.18	Surgical Instrument Sterilization Steam Autoclave Tape (i.e., Medline MDS200836 or equivalent).

8.19	Coolers with frozen gel packs for transporting processed and unprocessed BIs and PCDs.

8.20	Sterility testing medium of soybean-casein digest broth at volumes specified by B/F testing results.

8.21	Sterility testing equipment (ISO Class 5 Laminar Flow Hood or Bio-Safety Cabinet), cleanroom (ISO Class 6), test instruments (sterilized forceps, etc.), and supplies (disinfectants, low lint wipes, sterile body suits, etc.).

8.22	Validated incubator, 55 C to 60 C

8.23	Validated refrigerator, 2 C to 8 C

8.24	Steris DART Daily Air Removal Test Kit or equivalent

8.25	Calibrated balance with a range from 0 grams to 2000 grams and resolution of 1 gram.

9.0	STERILIZER MANUFACTURER’S DESIGN PARAMETERS AND SPECIFICATIONS

| Parameter | Specification |
| --- | --- |
| Chamber Size (W x H x D) | 20” x 20” x 38” |
| Overall Size (W x H x D) | 30” x 74½” x 45¼” (762mm x 1892mm x 1152 mm) |
| Electrical Requirements (with optional steam generator) | 120 VAC, 50/60 Hz, 2.5 Amps, 1Ф 208 VAC, 50/60 Hz, 83.2 Amps, 3Ф 240 VAC, 50/60 Hz, 72.2 Amps, 3Ф 440/480 VAC, 50/60Hz, 31/37 Amps, 3Ф |
| Steam Pressure | 50 to 80 psig (344.7 to 551.6 Kpa) |
| Steam Pressure Consumption | 116 lb/hr (53 kg/hr) peak |
| Utility Hot Water Pressure | 137.9 to 344.7 Kpa (20 to 50 psig) |
| Utility Hot Water Temperature | 140 F maximum |
| Utility Hot Water Consumption | Peak 1 gpm, per cycle 40 gal, Idle 1 gph |
| Environmental Temperature Operating Conditions | 50 F to 90 F (10 C to 32 C) |
| Environmental Relative Humidity Operating Conditions | 10% RH to 90% RH, non-condensing |
| Pollution Degree | 2 |
| A-Weighted Sound Power Level | 85 dBA maximum |

| MicroLogix 1200 PLC Item | MicroLogix 1200 PLC Specification |
| --- | --- |
| Environmental Temperature Operating Conditions | 32 F to 131 F (0 C to 55 C) |
| Environmental Relative Humidity Operating Conditions | 5% RH to 95% RH, non-condensing |

10.0	SAFETY PRECAUTIONS

10.1	All personnel working on or near sterilizers must observe all safety precautions and wear proper personal protective covering in order to avoid injury and/or equipment damage.

10.2	Do not use the sterilizer in any manner not specified in the operator’s manual as improper use may impair the safety features of the unit.

11.0	PROTOCOL DOCUMENTATION PRACTICES

11.1	Manual Entries in the Protocol and Appendices

11.1.1	All manual entries are written in indelible ink unless otherwise specified.

11.0	PROTOCOL DOCUMENTATION PRACTICES (Continued)

11.1.2	Errors are crossed out with a single line stroke, and the correct information is entered.  All corrections are initialed and dated with a justification code.  A justification code legend is included in the final summary report if errors are made.

11.1.3	Results are recorded on each page requiring data with the executer’s initials and date the data was recorded.

11.2	In the event that any of the sections do not adequately test a particular function, if additional or modified testing is required, if a section is determined not to apply, or if any general text corrections are required; the following acronyms and/or process is used:

11.2.1	The acronym AT denotes an additional test.  AT signifies that an additional test was deemed necessary during the protocol execution process, but was omitted from the protocol.  The AT is documented on the discrepancy log, which is included within this protocol as Appendix 1.

11.2.2	The acronym MT denotes a modified test.  MT signifies that a modification to an existing test was deemed necessary during the protocol execution process to more accurately reflect the conditions needed to satisfy the test requirement.  The MT is documented on the discrepancy log, which is included within this protocol as Appendix 1.

11.2.3	The acronym N/A or NA denotes not applicable.  N/A or NA is used when a step of the protocol is determined not to apply to the actual conditions.  The use of this acronym is not logged as a discrepancy.

11.2.4	General text corrections (i.e. typographical errors, spelling errors, erroneous entries by persons executing the PQ protocol, etc.) are rectified by lining out the error with a single line, entering the correct information, using a justification code, and then initialing and dating the entry.  These types of corrections are not logged as discrepancies, but are noted for future protocol revision.

11.3	After PQ protocol approval, the required data is gathered during the execution phase.

12.0	DISCREPANCIES

12.1	If there is a lack of compatibility or similarity between two or more facts shown within the PQ protocol, a section is not adequately tested, and/or an additional or modified testing is required; then this is considered a discrepancy and is documented using the discrepancy log, which is included within this protocol as Appendix 1.

12.0	DISCREPANCIES (Continued)

12.2	The discrepancy log contains the following:

12.2.1	Section:  Refers to the specific section where the discrepancy occurred.

12.2.2	Description/Comments:  Description or comments regarding the discrepancy.

12.2.3	Corrective Action:  Description of the steps taken to resolve the discrepancy.

12.2.4	Completed by:  Each discrepancy entry is initialed and dated by the person completing the discrepancy log.

13.0	DEVIATIONS

Any failure to comply with PQ protocol execution instructions, failure to meet defined PQ protocol specifications, and/or the specified PQ protocol acceptance criteria is considered a deviation and is documented on protocol deviation form FRM020-012.

14.0	EXECUTION LOG

14.1	The purpose of this execution log is to identify the trained persons participating in the PQ execution process.  PQ protocol training records are included in the final summary report.

| Printed Name | Title | Initials | Date |
| --- | --- | --- | --- |

14.2	All tests are performed by trained and experienced technical personnel and documented in a scientific manner using this established format.

15.0	LOAD CONFIGURATIONS SPECIFICATIONS

15.1	A maximum load configuration represents the upper boundary for the maximum number of items that can be processed routinely.

15.2	The following attributes substantiate a maximum load configuration:

15.2.1	Peak sterilizer load volume and density

15.0	LOAD CONFIGURATION SPECIFICATIONS (Continued)

15.2.2	Various residual air concentrations (air traps)

15.2.3	Complex article composition (varying materials and material)

15.2.4	High heterogeneousness of load

15.2.5	Four (4) sterilizer racks utilized

15.3	A minimum load configuration represents the lower boundary for the minimum number of items that can be processed routinely.

15.4	The following attributes substantiate a minimum load configuration:

15.4.1	Minimal sterilizer load volume and density

15.4.2	Minimal residual air concentrations (air traps)

15.4.3	Simple article composition

15.4.4	Low heterogeneousness of load

15.4.5	One (1) sterilizer rack utilized

Note:	In qualifying load configurations that represent the upper and lower boundaries of representative laboratory production instruments, testing supplies, and quality control items, the end user is assured that sterilizing any combination of like or unlike items specified in Section 15.5 and 15.6 is assured a SAL of at least a 10-6.

15.5	The following items comprise the maximum MDL1 load configuration.

15.5.1	Four (4) sterilizer loading racks.

15.5.2	Two (2) Artisan Metal Works Silicone Baking Mats, part number T3602MATA

15.5.2.1	Each item is wrapped in a double layer of sterilization wrap.

15.5.2.2	Wrapping is performed in accordance with Stimlabs procedures.

15.5.3	Six (6) Integra Stainless Steel Scissors, part number V95-26.

15.5.3.1	Each item is wrapped in a double layer of sterilization wrap.

15.0	LOAD CONFIGURATION SPECIFICATIONS (Continued)

15.5.3.2	Wrapping is performed in accordance with Stimlabs procedures.

15.5.4	Six (6) Vetmed USA Mayo Scissors, part number 26-611.

15.5.4.1	Each item is wrapped in a double layer of sterilization wrap.

15.5.4.2	Wrapping is performed in accordance with Stimlabs procedures.

15.5.5	Two (2) MABIS Precision Forceps, part number 25-736-000.

15.5.5.1	Each item is wrapped in a double layer of sterilization wrap.

15.5.5.2	Wrapping is performed in accordance with Stimlabs procedures.

15.5.6	One (1) Valve Solutions Stainless Steel Perforating Roller, part number N/A.

15.5.6.1	The item is double wrapped with sterilization pouches.

15.5.6.2	The wrapped item is then placed inside of a sterilization container that has the bottom side of the container lined with two (2) layers of sterilization wrap.

15.5.6.3	The item is placed on top of the double layer of sterilization wrap.

15.5.6.4	Wrapping is performed in accordance with Stimlabs procedures.

15.5.7	Two (2) Valve Solutions Stainless Steel Cutting Templates, part number N/A

15.5.7.1	Each item is double wrapped in sterilization pouches.

15.5.7.2	Wrapping is performed in accordance with Stimlabs procedures.

15.5.8	Two (2) Grafco Stainless Steel Trays, UNSPSC code 85100000.

15.5.8.1	Each item is wrapped in a double layer of sterilization wrap.

15.5.8.2	Wrapping is performed in accordance with Stimlabs procedures.

15.5.9	One (1) Grafco Stainless Steel Tray, UNSPSC code 56101700.

15.5.9.1	The item is wrapped in a double layer of sterilization wrap.

15.5.9.2	Wrapping is performed in accordance with Stimlabs procedures.

15.0	LOAD CONFIGURATION SPECIFICATIONS (Continued)

15.5.10	One (1) VirTis Stainless Steel Bulk Lyophilization Tray, part number 333948.

15.5.10.1	The item is wrapped in a double layer of sterilization wrap.

15.5.10.2	Wrapping is performed in accordance with Stimlabs procedures.

15.5.11	One (1) VirTis Stainless Steel Vial Lyophilization Tray, part number 399473 with one (1) metal insert and two (2) Lyophilizer Shelf Spacers, part numbers 100003743.

15.5.11.1	The lyophilizer shelf spacers are wrapped together in a double layer of sterilization wrap and placed within the vial lyophilization tray.

15.5.11.2	The metal insert is then placed within the vial lyophilization tray and the entire tray with its contents is wrapped in a double layer of sterilization wrap.

15.5.11.3	Wrapping is performed in accordance with Stimlabs procedures.

15.5.12	Five (5) Contec Polyester EasyReach Quilted Cleaning Pads, part number MEQT0001.

15.5.12.1	Each item is double wrapped in sterilization pouches.

15.5.12.2	Wrapping is performed in accordance with Stimlabs procedures.

15.5.13	Two (2) Valve Solutions Stainless Steel No-Go Gauges, part number N/A.

15.5.13.1	Each item is double wrapped in sterilization pouches.

15.5.13.2	Wrapping is performed in accordance with Stimlabs procedures.

15.5.14	One (1) Medline Sterilization Container, part number WAG1V13ASA.

15.5.14.1	The Valve Solutions Stainless Steel Perforating Roller is placed within the container during the sterilization process.  The container allows easy transport of the stainless steel roller without the risk of contamination.

15.5.14.2	The container is designed to maintain a sterile barrier.  Therefore, the item will not need to be covered with sterilization wrap.

15.0	LOAD CONFIGURATION SPECIFICATIONS (Continued)

15.5.15	The physical layout of the sterilizer shelving for maximum load configuration MDL1 is shown in Appendices 2 through 5, which is located at the end of this protocol.

15.6	The following items comprise the maximum PDL1 load configuration.

15.6.1	Four (4) sterilizer loading racks.

15.6.2	Thirty-two (32) Nalgene™ Square Wide-Mouth Polypropylene Copolymer 1000 mL Bottles with Closures, part number 2110-0032.

15.6.2.1	Each bottle is prepared by removing the bottle closure from the top of the bottle and placing it on the bottom side of the container.

15.6.2.2	Each bottle, with accompanying closure, is wrapped in a double layer of sterilization wrap.

15.6.2.3	Wrapping is performed in accordance with Stimlabs procedures.

15.6.3	One (1) GEPCO/General Econopak 16” x 26” Tyvek Sterilization Pouch, part number 01626TWBBSSSD.

15.6.3.1	The item is wrapped in a double layer of sterilization wrap.

15.6.3.2	Wrapping is performed in accordance with Stimlabs procedures.

15.6.4	Two (2) Small Parts Polypropylene Cutting Boards (12” x 24” x 0.25”), part number PPST-0250-F.

15.6.4.1	Each item is wrapped in a double layer of sterilization wrap.

15.6.4.2	Wrapping is performed in accordance with Stimlabs procedures.

15.6.5	The physical layout of the sterilizer shelving for maximum load configuration PDL1 is shown in Appendices 6 through 9, which is located at the end of this protocol.

15.7	The following items comprise the minimum MDL1 load configuration.

15.7.1	One (1) sterilizer loading rack at sterilizer shelf location #4.

15.7.2	One (1) Integra Stainless Steel Scissor, part number V95-26.

15.0	LOAD CONFIGURATION SPECIFICATIONS (Continued)

15.7.2.1	The item is wrapped in a double layer of sterilization wrap.

15.7.2.2	Wrapping is performed in accordance with Stimlabs procedures.

15.7.3	One (1) Vetmed USA Mayo Scissor, part number 26-611.

15.7.3.1	The item is wrapped in a double layer of sterilization wrap.

15.7.3.2	Wrapping is performed in accordance with Stimlabs procedures.

15.7.4	One (1) Valve Solutions Stainless Steel Perforating Roller, part number N/A.

15.7.4.1	The spiked head of the item is wrapped in a double layer of sterilization wrap.

15.7.4.2	The sterilization wrapped item is then double wrapped in sterilization pouches.

15.7.4.3	Wrapping is performed in accordance with Stimlabs procedures.

15.7.5	One (1) Contec Polyester EasyReach Quilted Cleaning Pad, part number MEQT0001.

15.7.5.1	The item is double wrapped in sterilization pouches.

15.7.5.2	Wrapping is performed in accordance with Stimlabs procedures.

15.7.6	The physical layout of the sterilizer shelving for minimum load configuration MDL1 is shown in Appendix 10, which is located at the end of this protocol.

15.8	The following items comprise the minimum PDL1 load configuration.

15.8.1	One (1) sterilizer loading rack at sterilizer shelf location #4.

15.8.2	Three (3) Nalgene™ Square Wide-Mouth Polypropylene Copolymer 1000 mL Bottles with Closures, part number 2110-0032.

15.8.2.1	Each bottle is prepared by removing the bottle closure from the top of the bottle and placing it on the bottom side of the container.

15.8.2.2	Each bottle with accompanying closure is wrapped in a double layer of sterilization wrap.

15.0	LOAD CONFIGURATION SPECIFICATIONS (Continued)

15.8.2.3	Wrapping is performed in accordance with Stimlabs procedures.

15.8.3	One (1) 2” x 4” x 0.25” section from the Small Parts Polypropylene Cutting Board, part number PPST-0250-F.

15.8.3.1	The item is wrapped in a double layer of sterilization wrap.

15.8.3.2	Wrapping is performed in accordance with Stimlabs procedures.

15.8.4	The physical layout of the sterilizer shelving for minimum load configuration PDL1 is shown in Appendix 11, which is located at the end of this protocol.

16.0	PCD PREPARATION

16.1	In accordance with AAMI TIR12:2010, “Designing, testing, and labeling reusable medical devices for reprocessing in health care facilities: A guide for medical device manufacturers,” microbiological challenges are placed in the most difficult-to-sterilize, accessible locations of an item.

16.2	The same principle from AAMI TIR12:2010 is used for Stimlabs laboratory production instruments, testing supplies, and quality control cleaning articles that are constructed of porous material or are pivoted hand tools that allow movement by in-hand manipulation.

16.3	If it is not possible to reach the most difficult-to-sterilize, accessible locations of an item with conventional biological indicators (i.e., Bio-threads, paper discs, spore strips, etc.) then an item may be inoculated with the specified microbiological challenge by using a liquid suspension BI.

16.4	For the purposes of this PQ, when an item is inoculated with a BI, then the item is deemed a PCD.  The following PCD items are considered the most difficult-to-sterilize due to porous material construction or are pivoted hand tools that allow movement by in-hand manipulation:

16.4.1	Artisan Metal Works Silicone Baking Mat, part number T3602MATA, which is wrapped in a double layer of sterilization wrap.

16.4.2	Integra Stainless Steel Scissor, part number V95-26, which is wrapped in a double layer of sterilization wrap.

16.4.3	Vetmed USA Mayo Scissor, part number 26-611, which is wrapped in a double layer of sterilization wrap.

16.0	PCD PREPARATION (Continued)

16.4.4	Valve Solutions Stainless Steel Perforating Rollers, part number N/A, which is wrapped in a double layer of sterilization wrap.

16.4.5	Contec Polyester EasyReach Quilted Cleaning Pads, part number MEQT0001, which is double wrapped in sterilization pouches.

16.4.6	One (1) 2” x 4” x 0.25” section from the Small Parts Polypropylene Cutting Board, part number PPST-0250-F, which is wrapped in a double layer of sterilization wrap.

Note:	The 2” x 4” x 0.25” section of cutting board is manually grooved several times with a knife to simulate excessive cutting on the board as a worst-case sterilization challenge when inoculated with BIs.

16.5	Appendices 12 through 17 shows where the BI inoculates are applied to each PCD.

16.6	The PCDs are prepared and used during the microbiological challenge study cycle run phase of the PQ.

16.7	WXAT personnel confirm the microbial population of the bacterial spore suspension in accordance with applicable procedures prior to conducting the microbiological challenge study cycle runs.

16.8	WXAT personnel inoculate a minimum of four (4) each of the items specified in Section 16.4 for a minimum total of twenty-four (24) PCDs.

16.8.1	Eighteen (18) are used for processing in the microbiological challenge study cycle runs.

16.8.2	Six (6) are used as positive controls as a part of the microbiological challenge study cycle run process.

16.9	WXAT personnel perform the inoculation process using the BI material specified in Section 8.9.1 and by following WXAT standard operating procedure MA229SOP, Product Inoculation.”

16.10	WXAT personnel wrap each PCD as specified in Section 16.4.

17.0	TESTING PROCEDURE

17.1	Leak and DART

17.0	TESTING PROCEDURE (Continued)

17.1.1	A leak and DART are performed each day of testing.

17.1.2	The DART is performed before the leak test.

17.1.2.1	The DART is performed to determine the efficiency of air removal from loads.  DARTs detect problems at a very early stage because they are performed daily and are very sensitive to small changes in the sterilizer’s performance.

17.1.2.2	Use the DART specified in Section 8.24.

17.1.2.3	Place the DART on the lowest shelf over the drain, in an otherwise empty chamber as shown in Figure 1.  During the vacuum phase of the cycle, air is removed through the tortuous challenge chamber of the DART and then steam enters the chemical indicator chamber during the exposure phase.

Figure 1 – DART Placement in Empty Chamber

17.1.2.4	If passing, the DART indicator bars will turn black, which indicates that residual air was completely removed by allowing complete steam penetration as shown in Figure 2.

Figure 2 – Passing DART Results

17.1.2.5	A failure indicates air is trapped in the chamber and all the bars will not turn black as shown in Figure 3.

17.0	TESTING PROCEDURE (Continued)

Figure 3 – Typical Failing Result

17.1.3	The leak test determines whether there are any leaks in the chamber and/or pipework which could permit ingress of air into the chamber during the vacuum stages of a sterilization cycle.

17.1.4	The leak test is performed in an otherwise empty chamber.

      - A leak rate of greater than 1.0 mmHg/minute indicates that the sterilizer is in need of service and should be addressed.

17.1.4.2	A leak test does not definitively determine whether the unit is fit for use as it is intended as a maintenance indicator only.  The DART is the primary indicator of whether the sterilizer is fit for use.

17.2	Empty Chamber Sterilizer Profiles/Mappings

17.2.1	A total of six (6) empty chamber sterilizer profiles/mappings are performed.

17.2.2	The sterilizer settings are specified in Tables 2 and 3.

17.0	TESTING PROCEDURE (Continued)

Table 2 – Study Parameters and Sterilizer Settings

| Run | SterilizationCycle Type | Sterilizer Temp. Setting | Temp. Range | Cycle Time* | Print Interval | Overdrive Setting | Over Temp. Setting | Under Temp. Setting | Jacket Pressure Range |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1st | Pre-Vacuum | 121 C | 121 C to 124 C | 15 Mins. | 1.0 Min. | 1.0 C | 1.0 C | 1.0 C | ~16.0 to 20.0 PSIG |
| 2nd | Pre-Vacuum | 121 C | 121 C to 124 C | 15 Mins. | 1.0 Min. | 1.0 C | 1.0 C | 1.0 C | ~16.0 to 20.0 PSIG |
| 3rd | Pre-Vacuum | 121 C | 121 C to 124 C | 15 Mins. | 1.0 Min. | 1.0 C | 1.0 C | 1.0 C | ~16.0 to 20.0 PSIG |
| 4th | Pre-Vacuum | 132 C | 132 C to 135 C | 15 Mins. | 1.0 Min. | 1.0 C | 1.0 C | 1.0 C | ~27.0 to 31.0 PSIG |
| 5th | Pre-Vacuum | 132 C | 132 C to 135 C | 15 Mins. | 1.0 Min. | 1.0 C | 1.0 C | 1.0 C | ~27.0 to 31.0 PSIG |
| 6th | Pre-Vacuum | 132 C | 132 C to 135 C | 15 Mins. | 1.0 Min. | 1.0 C | 1.0 C | 1.0 C | ~27.0 to 31.0 PSIG |

Table 3 – Study Parameters and Sterilizer Settings

| Run | SterilizationCycle Type | Pulse Pressure Charge | Dry Cont. | Pulse Pressure Vacuum | Purge Time Setting | Vacuum Pulses | STRL Control | Start Dry Vacuum Time | Dry Time | Exhaust |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1st | Pre-Vacuum | 26.0 PSIG | 0 | 10.0 inHG | 1.5 Minutes | 4 | Drain | 10.0 inHG | 0 Mins. | 0 |
| 2nd | Pre-Vacuum | 26.0 PSIG | 0 | 10.0 inHG | 1.5 Minutes | 4 | Drain | 10.0 inHG | 0 Mins. | 0 |
| 3rd | Pre-Vacuum | 26.0 PSIG | 0 | 10.0 inHG | 1.5 Minutes | 4 | Drain | 10.0 inHG | 0 Mins. | 0 |
| 4th | Pre-Vacuum | 26.0 PSIG | 0 | 10.0 inHG | 1.5 Minutes | 4 | Drain | 10.0 inHG | 0 Mins. | 0 |
| 5th | Pre-Vacuum | 26.0 PSIG | 0 | 10.0 inHG | 1.5 Minutes | 4 | Drain | 10.0 inHG | 0 Mins. | 0 |
| 6th | Pre-Vacuum | 26.0 PSIG | 0 | 10.0 inHG | 1.5 Minutes | 4 | Drain | 10.0 inHG | 0 Mins. | 0 |

*The minimum sterilization cycle time is in accordance with British Standard EN554:1994.

17.0	TESTING PROCEDURE (Continued)

17.2.3	A total of fourteen (14) calibrated temperature system sensors are used during each empty chamber sterilizer profile/mapping study cycle run.

17.2.4	A total of four (4) sterilizer racks are used during each empty chamber sterilizer profile/mapping study cycle run.

17.2.5	The calibrated temperature sensors are placed in accordance with Appendices 18 through 21 for each empty chamber sterilizer profile/mapping study cycle run.

17.2.6	One calibrated temperature sensor is placed near the drain to compare its reading with the sterilizer control temperature.

17.2.7	The results from each empty chamber sterilizer profile/mapping study cycle run are verified in meeting the following:

17.2.7.1	The variation in temperature of a temperature sensing probe does not exceed ±1 C or a 2 C range during the sterilization phase of the cycle run within the sterilizer loading zone.

17.2.7.2	The variation between all the temperature sensing probes (sensor-to-sensor readings) do not differ from each other by more than +2 °C during the sterilization phase of the cycle run within the sterilizer loading zone.

17.2.7.3	The difference in temperature of the sterilizer loading zone as compared between the set point temperature of the sterilizer and the temperature sensing probes is no greater than +3 °C during the sterilization phase of the cycle run.

17.2.8	The results from all empty chamber sterilizer profile/mapping study cycle runs are analyzed to determine if a consistent cold spot within the sterilizer loading zone is present.

17.2.9	If a cold spot is identified, the area is challenged in microbiological challenge study cycle runs.

17.3	Heat Penetration Study Cycle Runs (FPHY)

17.3.1	Heat penetration study cycle runs are performed to gather thermometric data in order to establish and determine the following:

17.0	TESTING PROCEDURE (Continued)

17.3.1.1	Establish the boundaries for the minimum and maximum number of load configuration MDL1 and PDL1 items that can be sterilized within the sterilizer during routine sterilization processes.

17.3.1.2	Determine the half-sterilization cycle time at the conclusion of each minimum and maximum load configuration that achieves a calculated minimum SAL of 10-3.

17.3.1.3	Determine the full-sterilization cycle time at the conclusion of each minimum and maximum load configuration that achieves a calculated minimum SAL of 10-6.

17.3.1.4	Determine if the maximum or minimum load configuration is the worst-case load configuration by analyzing which load configuration accumulated the least amount of accumulated lethality at the end of all triplicate heat penetration study cycle runs.

- Accumulated lethality is determined by the following equation using the validated temperature measuring system’s software:

F=DΣL

Where:

D = Time increment between each temperature reading

L = Lethal rate calculated for each temperature reading.  The lethal rate equation is shown by the following:

L=10(T-Tref)Z

Where:

T = Sensor Temperature

Tref = Reference Temperature

Z = Z-value of 10 C

- Table 4 lists the required temperature measurement system program settings for each temperature sensor used during the heat penetration study cycle runs:

17.0	TESTING PROCEDURE (Continued)

Table 4 – Temperature Measurement System Sensor Program Settings for FPHY Study Runs

| Temperature Sensor Settings |  |  |  |
| --- | --- | --- | --- |
| D121°C-value & D132°C-value | Z-value | Accumulation Temperature | Logging Interval |
| 3.0 Minutes | 10 C | 100 C | 1.0 Minutes |

17.3.1.5	Determination of adequate dry time.

- Absorbent material wicks condensate away from items and disperses it over a greater surface area for more efficient drying.

- However, an excessive amount of absorbent material or the incorrect type of absorbent material can impede air removal and sterilant penetration and interfere with proper drying.

- The Initial dry time used for the first heat penetration study cycle run is 20 minutes in accordance with AAMI TIR12:2010 and is adjusted if needed by five (5) minute intervals.

- In this PQ, adequate dry time is evaluated by the following:

- A sampling of any absorbent materials used (e.g., surgical towels, tray liners, sterilization wrapping materials) are removed and weighed, and the post-processing weight is compared with the pre-processing weight. The difference in weight must not exceed +3%.

Note:	A calibrated balance is used to weigh a sampling of absorbent material wrapped items from each sterilizer shelf before and after sterilization using Appendices 83 and 84.

- No evidence of visible moisture (condensation) on or within sterilized packages, containers, or trays.

- If the first heat penetration study cycle run, which is performed on a maximum load configuration, does not meet the adequate dry time requirements, then five (5) minutes is added to the sterilizer’s dry time program parameter.

- If the second heat penetration study cycle run on a maximum load configuration does not meet the adequate dry time requirements, then an additional five (5) minutes is added to the sterilizer’s dry time program parameter.

17.0	TESTING PROCEDURE (Continued)

- If the third heat penetration study cycle run on a maximum load configuration does not meet the adequate dry time requirements, then an additional five (5) minutes is added to the sterilizer’s dry time program parameter.

Note:	Dry time is increased by five (5) minute intervals using a maximum load configuration until successful results are achieved.  Once adequate dry time has been reached, repeatability in achieving proper drying must be demonstrated by repeating at least three (3) heat penetration study cycle runs on a maximum load configuration.

- Documenting the dry time verification/calibration process is done using Appendix 46.

- Dry times determined acceptable for maximum load configurations are deemed acceptable for minimum load configurations.

17.3.2	Maximum Load Configuration Heat Penetration Study Cycle Runs

17.3.2.1	The maximum load configuration heat penetration study cycle runs are performed in accordance with the settings specified in Tables 5 and 6.

17.3.2.2	Fourteen (14) validated and calibrated temperature measuring system sensors are used for each maximum load configuration heat penetration study cycle run.

17.3.2.3	The temperature measuring system sensors are placed within each load as specified by Appendices 22 through 45.

17.3.2.4	Each sensor is placed within each sterilization wrapped or pouch item specified by the attachments.

17.3.2.5	The dry time is verified/calibrated using Attachment 46.  The initial dry time for the first maximum load configuration heat penetration study cycle run is 20 minutes.

17.3.2.6	The minimum initial sterilization time is 15 minutes, which may require adjustment in order to achieve a minimum SAL of 10-3.

17.3.2.7	The data from the maximum load configuration heat penetration study cycle runs are included in the final summary report.

Heat Penetration Study Cycle Parameters and Settings for Max and Min Load Configurations MDL1 and PDL1

Table 5 – Study Parameters and Sterilizer Settings

| Run | Cycle Type | Load Configuration | Sterilizer Temp. Setting | Min. Cycle Time(1) | Print Interval | Overdrive Setting | Over Temp. Setting | Under Temp. Setting | Jacket Pressure Range |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | Pre-Vacuum | PDL1 | 121 C | 15 Mins. | 1.0 Min. | 1.0 C | 1.0 C | 1.0 C | ~16.0 to 20.0 PSIG |
| 2 | Pre-Vacuum | PDL1 | 121 C | 15 Mins. | 1.0 Min. | 1.0 C | 1.0 C | 1.0 C | ~16.0 to 20.0 PSIG |
| 3 | Pre-Vacuum | PDL1 | 121 C | 15 Mins. | 1.0 Min. | 1.0 C | 1.0 C | 1.0 C | ~16.0 to 20.0 PSIG |
| 1 | Pre-Vacuum | MDL1 | 132 C | 15 Mins. | 1.0 Min. | 1.0 C | 1.0 C | 1.0 C | ~27.0 to 31.0 PSIG |
| 2 | Pre-Vacuum | MDL1 | 132 C | 15 Mins. | 1.0 Min. | 1.0 C | 1.0 C | 1.0 C | ~27.0 to 31.0 PSIG |
| 3 | Pre-Vacuum | MDL1 | 132 C | 15 Mins. | 1.0 Min. | 1.0 C | 1.0 C | 1.0 C | ~27.0 to 31.0 PSIG |

Table 6 – Study Parameters and Sterilizer Settings

| Run | Load Configuration | Pulse Pressure Charge | Dry Cont. | Pulse Pressure Vacuum | Purge Time Setting | Vacuum Pulses | STRL Control | Start Dry Vacuum Time | Dry Time(2) | Exhaust |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | PDL1 | 26.0 PSIG | 0 | 10.0 inHG | 1.5 Minutes | 4 | Drain | 10.0 inHG | See Attachment 46 | 0 |
| 2 | PDL1 | 26.0 PSIG | 0 | 10.0 inHG | 1.5 Minutes | 4 | Drain | 10.0 inHG | See Attachment 46 | 0 |
| 3 | PDL1 | 26.0 PSIG | 0 | 10.0 inHG | 1.5 Minutes | 4 | Drain | 10.0 inHG | See Attachment 46 | 0 |
| 1 | MDL1 | 26.0 PSIG | 0 | 10.0 inHG | 1.5 Minutes | 4 | Drain | 10.0 inHG | See Attachment 46 | 0 |
| 2 | MDL1 | 26.0 PSIG | 0 | 10.0 inHG | 1.5 Minutes | 4 | Drain | 10.0 inHG | See Attachment 46 | 0 |
| 3 | MDL1 | 26.0 PSIG | 0 | 10.0 inHG | 1.5 Minutes | 4 | Drain | 10.0 inHG | See Attachment 46 | 0 |

(1) Minimum cycle time in accordance with British Standard EN554:1994

17.0	TESTING PROCEDURE (Continued)

17.3.3	Minimum Load Configuration Heat Penetration Study Cycle Runs

17.3.3.1	The minimum load configuration heat penetration study cycle runs are performed in accordance with the settings specified in Tables 5 and 6.

17.3.2.2	Fourteen (14) validated and calibrated temperature measuring system sensors are used for each minimum load configuration heat penetration study cycle run.

17.3.2.3	The temperature measuring system sensors are placed within each load as specified by Appendices 47 through 52.

17.3.2.3	Each sensor is placed within each sterilization wrapped or pouch item specified by the attachments.

17.3.2.4	The sterilization time may require adjustment in order to achieve a minimum SAL of 10-3.

17.3.2.5	The data from the minimum load configuration heat penetration study cycle runs are included in the final summary report.

17.4	Microbiological Challenge Study Cycle Runs

17.4.1	Three (3) microbiological challenge study cycle runs are performed on the worst-case MDL1 and PDL1 load configurations.

Note:	The worst-case load configuration is the one that achieved the least amount of accumulated lethality during the heat penetration study cycle runs.

17.4.2	The microbiological challenge study runs are performed in order to confirm that half-sterilization cycle times established during the heat penetration study cycle runs are achieving a minimum SAL of 10-3, thus demonstrating that full-sterilization cycle times are providing a minimum SAL of 10-6.

17.4.3	The microbiological challenge study cycle runs are performed in accordance with the settings specified in Tables 7 and 8.

Microbiological Challenge Study Cycle Parameters and Settings for

Max and Min Load Configurations MDL1 and PDL1

Table 7 – Study Parameters and Sterilizer Settings

| Run | Cycle Type | Load Configuration | Sterilizer Temp. Setting | Cycle Time(1) | Print Interval | Overdrive Setting | Over Temp. Setting | Under Temp. Setting | Jacket Pressure Range |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | Pre-Vacuum | PDL1 | 121 C | See note | 1.0 Min. | 1.0 C | 1.0 C | 1.0 C | ~16.0 to 20.0 PSIG |
| 2 | Pre-Vacuum | PDL1 | 121 C | See note | 1.0 Min. | 1.0 C | 1.0 C | 1.0 C | ~16.0 to 20.0 PSIG |
| 3 | Pre-Vacuum | PDL1 | 121 C | See note | 1.0 Min. | 1.0 C | 1.0 C | 1.0 C | ~16.0 to 20.0 PSIG |
| 1 | Pre-Vacuum | MDL1 | 132 C | See note | 1.0 Min. | 1.0 C | 1.0 C | 1.0 C | ~27.0 to 31.0 PSIG |
| 2 | Pre-Vacuum | MDL1 | 132 C | See note | 1.0 Min. | 1.0 C | 1.0 C | 1.0 C | ~27.0 to 31.0 PSIG |
| 3 | Pre-Vacuum | MDL1 | 132 C | See note | 1.0 Min. | 1.0 C | 1.0 C | 1.0 C | ~27.0 to 31.0 PSIG |

Table 8 – Study Parameters and Sterilizer Settings

| Run | Load Configuration | Pulse Pressure Charge | Dry Cont. | Pulse Pressure Vacuum | Purge Time Setting | Vacuum Pulses | STRL Control | Start Dry Vacuum Time | Dry Time(2) | Exhaust |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | PDL1 | 26.0 PSIG | 0 | 10.0 inHG | 1.5 Minutes | 4 | Drain | 10.0 inHG | See Attachment 46 | 0 |
| 2 | PDL1 | 26.0 PSIG | 0 | 10.0 inHG | 1.5 Minutes | 4 | Drain | 10.0 inHG | See Attachment 46 | 0 |
| 3 | PDL1 | 26.0 PSIG | 0 | 10.0 inHG | 1.5 Minutes | 4 | Drain | 10.0 inHG | See Attachment 46 | 0 |
| 1 | MDL1 | 26.0 PSIG | 0 | 10.0 inHG | 1.5 Minutes | 4 | Drain | 10.0 inHG | See Attachment 46 | 0 |
| 2 | MDL1 | 26.0 PSIG | 0 | 10.0 inHG | 1.5 Minutes | 4 | Drain | 10.0 inHG | See Attachment 46 | 0 |
| 3 | MDL1 | 26.0 PSIG | 0 | 10.0 inHG | 1.5 Minutes | 4 | Drain | 10.0 inHG | See Attachment 46 | 0 |

(1) Time determined by data from the he heat penetration study cycle runs.

17.0	TESTING PROCEDURE (Continued)

17.4.4	The temperature measurement system sensors for each microbiological challenge cycle run are programmed in accordance with Table 9:

Table 9 – Temperature Measurement System Sensor Program Settings for FBIO Study Runs

| Sensor Settings |  |  |  |
| --- | --- | --- | --- |
| D-value | Z-value | Lethality Accumulation Temperature | Logging Interval |
| BI Manufacturer’s D-value Listed on the Certificate of Analysis | Manufacturer’s Z-value Listed on the Certificate of Analysis | 100 C | 1.0 Minutes |

17.4.5	If a consistent “Cold Spot” was identified during the heat penetration study cycle runs, then the “Cold Spot” is challenged.

17.4.5.1	The “Cold Spot” is manually added to the applicable microbiological challenge study run appendices, if the location was not already identified by a BI or PCD, and challenged with BIs and measured with a temperature measuring system sensor.

17.4.5.2	A predefined sensor location shown on the microbiological challenge study run appendices may require movement to accommodate the “Cold Spot.”

17.4.6	If the maximum load configuration for MDL1 or PDL1 is determined to be the worst-case load configuration, then temperature measurement system sensors, PCDs, and BIs are placed in the locations specified by Appendices 53 to 76.

17.4.6.1	To preserve the integrity of a prepared PCD, the temperature measurement system sensors are placed outside of the sterilization wrap or pouch the PCD is contained in.

17.4.6.2	The following is needed for each microbiological challenge study run on the maximum MDL1 load configuration:

- Five (5) PCDs

- Fifteen (15) BIs

- Fourteen (14) temperature measurement system sensors

17.4.6.3	The following is needed for each microbiological challenge study run on the maximum PDL1 load configuration:

- One (1) PCD

17.0	TESTING PROCEDURE (Continued)

- Nineteen (19) BIs

- Fourteen (14) temperature measurement system sensors

17.4.7	If the minimum load configuration for MDL1 or PDL1 is determined to be the worst-case load configuration, then temperature measurement system sensors, PCDs, and BIs are placed in the locations specified by Appendices 77 to 82.

17.4.7.1	To preserve the integrity of a prepared PCD, the temperature measurement system sensors are placed outside of the sterilization wrap or pouch the PCD is contained in.

17.4.7.2	The following is needed for each microbiological challenge study run on the maximum MDL1 load configuration:

- Five (5) PCDs

- Fifteen (15) BIs

- Fourteen (14) temperature measurement system sensors

17.4.7.3	The following is needed for each microbiological challenge study run on the maximum PDL1 load configuration:

- One (1) PCD

- Nineteen (19) BIs

- Fourteen (14) temperature measurement system sensors

17.5	PCD/BI Post Sterilization Disposition

17.5.1	Upon completion of the microbiological challenge study cycle runs, the processed PCDs/BIs are removed by WXAT personnel and placed into a transport cooler with frozen gel packs along with the positive controls.

17.5.2	The PCDs/BIs are then transferred to WXAT’s validated refrigeration units until submitted to the Sterility Department for testing.

17.5.3	The PCDs/BIs are submitted to WXAT’s Sterility Department for testing as quickly as feasible.

17.0	TESTING PROCEDURE (Continued)

17.5.4	Copies of completed appendices, reports, and raw data records are included in the final summary report.

17.6	Sterility Testing

17.6.1	The PCDs and BIs from the microbiological challenge study cycle runs, plus the positive controls, are individually tested for growth according to WXAT’s standard operating procedures.

17.6.2	Traceability to specific locations is maintained.

17.6.3	Test samples are checked for growth according to established WXAT sterility department inspection schedules.

17.6.4	Positive samples may be removed from incubation prior to seven (7) days of incubation.

17.6.5	Sterility test results are recorded for each BI as positive or negative for growth of Geobacillus stearothermophilus.

17.6.6	For a positive test that does not display the characteristics of Geobacillus stearothermophilus, a microbial identification is performed to confirm the organism’s genus and species.

17.6.7	Sterility test results are included in the final summary report.

17.7	Microbiological Challenge Study Cycle Runs Lethality (FBIO)

17.7.1	The delivered lethality measured by the actual kill of microorganisms from the PCDs/BIs is calculated by the following:

FBIO=Process D-value x SLR

Where:

SLR=Spore log reduction of the BI population

17.7.2	The FBIO from the microbiological challenge study cycle runs is expected to demonstrate a lethality value of ≥ 12 logarithms.

17.0	TESTING PROCEDURE (Continued)

17.8	SLR

17.8.1	SLR calculations for half-sterilization cycle times are determined by the following equation:

SLR =Process Equivalent Time FTREF,   ZAProcess D-value

Where:

FTREF,   ZA= TA x 10-TRef-TAZA

FTREF,   ZA= Process equivalent time

TA= Process clock time interval

TREF = Process reference temperature

ZA = Actual Z-value of the challenge organism

17.8.2	The expectation for the microbiological challenge cycle is to show a process SLR value of ≥ 9 logarithms.

17.9	SLR calculations for full-sterilization cycle times are determined by the following equation:

SLR =Process Equivalent Time x 2 FTREF,   ZAProcess D-value

17.10	SAL

17.10.1	The following equation is used to calculate the minimum SAL for the microbiological challenge cycles at the half-sterilization cycle time:

SALHalf-Cycle=10(LogNo- Half Cycle SLR)

17.0	TESTING PROCEDURE (Continued)

17.10.2	The following equation is used to calculate the minimum SAL for the microbiological challenge cycles at the full-sterilization cycle time:

SALFull-Cycle=10(LogNo-(Full Cycle SLR)

    - The microbiological challenge study cycle runs must calculate to show a final minimum SAL of 10-6.

- ROUNDING

18.1	In accordance with USP current version, Section 7 of General Notices, the following rounding rules apply when reporting test results:

    - The observed or calculated values shall be rounded off to the number of decimal places that is in agreement with the limit expression.

    - Numbers should not be rounded until the final calculations for the reportable value have been completed.  Intermediate calculations (e.g., slope for linearity) may be rounded for reporting purposes, but the original non-rounded value should be used for any additional required calculations.

    - Acceptance criteria are fixed numbers and are not rounded.

    - When experimentally meaningless digits are rejected, treat the excess digits as a decimal fraction.  If the decimal fraction is equal to or greater than 0.5, increase the last digit retained by 1.  If the decimal fraction is less than 0.5, the last digit retained is left unchanged.

18.2	Table 10 shows examples of rounding numerical values for comparison with requirements.

Table 10 – Rounding Numerical Values for Comparison with Requirements

| Compendial Requirement | Unrounded Value | Rounded Result | Conforms |
| --- | --- | --- | --- |
| Assay Limit ≥ 98.0% | 97.96% | 98.0% | Yes |
|  | 97.92% | 97.9% | No |
|  | 97.95% | 98.0% | Yes |
| Assay Limit ≤ 101.5% | 101.55% | 101.6% | No |
|  | 101.46% | 101.5% | Yes |
|  | 101.45% | 101.5% | Yes |

18.0	ROUNDING (Continued)

Table 10 – Rounding Numerical Values for Comparison with Requirements (Continued)

| Compendial Requirement | Unrounded Value | Rounded Result | Conforms |
| --- | --- | --- | --- |
| Limit Test ≤ 0.02% | 0.025% | 0.03% | No |
|  | 0.015% | 0.02% | Yes |
|  | 0.027% | 0.03% | No |
| Limit Test ≤ 3 ppm | 3.5 ppm | 4 ppm | No |
|  | 3.4 ppm | 3 ppm | Yes |
|  | 2.5 ppm | 3 ppm | Yes |

19.0	ACCEPTANCE CRITERIA

19.1	PCDs used in the PQ must pass B/F testing.

19.2	The BIs used in the PQ must pass TVSC testing where the certified BI population is confirmed to be within -0.3 log and +0.48 log of the manufacturer’s label spore count claim.

19.3	The empty chamber sterilizer profile/mapping study cycle runs must demonstrate the following:

19.3.1	The variation in temperature of a temperature sensing probe does not exceed ±1 C or a 2 C range during the sterilization phase of the cycle run within the sterilizer loading zone.

19.3.2	The variation between all the temperature sensing probes (sensor-to-sensor readings) do not differ from each other by more than +2 °C during the sterilization phase of the cycle run within the sterilizer loading zone.

19.3.3	The difference in temperature of the sterilizer loading zone as compared between the set point temperature of the sterilizer and the temperature sensing probes is no greater than +3 °C during the sterilization phase of the cycle run.

19.4	The Leak and DART must pass before conducting cycle runs for the day.

19.5	SLR at Half-Sterilization Cycle Time

The microbiological challenge cycle is to show a process SLR value of ≥ 9 logarithms.

19.6	SLR at Half-Sterilization Cycle Time

The microbiological challenge cycle is to show a process SLR value of ≥ 12 logarithms.

19.0	ACCEPTANCE CRITERIA (Continued)

19.7	SAL at Half-Sterilization Cycle Time

The microbiological challenge study cycle runs must calculate a minimum SAL of 10-3.

19.8	SAL at Full-Sterilization Cycle Time

The microbiological challenge study cycle runs must calculate a minimum SAL of 10-6.

19.9	Sterility Test Results

19.9.1	Negative test results for growth of Geobacillus stearothermophilus for all PCDs from the microbiological challenge study cycle runs.

19.9.2	Negative test results for growth of Geobacillus stearothermophilus for all conventional BIs from the microbiological challenge study cycle runs.

19.9.3	Positive test results for growth of Geobacillus stearothermophilus for all positive control PCDs.

19.9.4	Positive test results for growth of Geobacillus stearothermophilus for all positive control conventional BIs.

19.10	A minimum of thirteen (13) temperature measurement system sensors must be functional at all times for all runs conducted within the PQ.

20.0	INTERPRETATION OF DATA/CONCLUSION

If the acceptance criteria of this PQ protocol are met, then the Load configuration MDL1 and PDL1 using the Steris steam sterilizer; model number LV250, serial number 0305716-26 is considered qualified.

21.0	FINAL SUMMARY REPORT

The final summary report is written following the execution of the PQ protocol and includes, but is not limited to, the sections listed below:

| Section | Description |
| --- | --- |
| Scope/Objective | Description of the objectives of the PQ |
| Summary of Results | Summary of the results for each major section of the protocol including a summary of deviation(s) |
| Conclusion | General conclusion for all testing and observations indicating the acceptability of the sterilizer |
| Attachments | Deviation Reports and the executed PQ protocol is included within the final summary report |

22.0	REVISIONS

22.1	If it becomes necessary to make changes to the approved PQ protocol, the revisions and reasons for changes are documented, reported to all respective parties, signed, dated, and maintained with the PQ protocol.

22.2	If the change affects the validity of the PQ, Stimlabs assess the change and determines if any or all phases of the PQ execution process needs repeating.

23.0	AMENDMENTS

If an amendment is required to improve the clarity, make corrections to, or revise the original content of an already executed protocol, then the amendment is issued, documented, and retained with the final summary report.

24.0	RETENTION OF RECORDS

24.1	WXAT retains the master electronic data files generated from the PQ process.

24.2	Stimlabs retains all paper records generated from the PQ process.

25.0	INFORMATIVE REFERENCES

25.1	British Standard EN 554:1994, “Sterilization of medical devices – Validation and routine control of sterilization by moist heat”

25.2	Parenteral Drug Association, Inc., Technical Report No. 1, “Validation of Moist Heat Sterilization Processes: Cycle Design, Development, Qualification and Ongoing Control, Revised 2007”

25.3	Process Validation Plan VAP16-010-00

25.4	ISO 11138-3, “Sterilization of health care products — Biological indicators — Part 3: Biological indicators for moist heat sterilization processes”

25.5	ANSI/AAMI/ISO 20857:2010, “Sterilization of health care products – Dry heat – Requirements for the development, validation and routine control of a sterilization process for medical devices”

25.6	WXAT Standard Operating Procedure MA225SOP, “Total Viable Spore Count – Biological Indicators”

25.7	USP current version, chapter <55>, “Biological Indicators – Resistance Performance Tests, Total Viable Spore Count”

25.8	USP monograph, “Biological Indicators for Moist Heat, Dry Heat, and Gaseous Modes of Sterilization, Nonpaper Carriers”

25.9	ANSI/AAMI/ISO 11138 guidelines

25.10	USP current version, chapter <1035>, “Biological Indicators for Sterilization”

25.11	AAMI TIR12:2010, “Designing, testing, and labeling reusable medical devices for reprocessing in health care facilities: A guide for medical device manufacturers”

25.12	WXAT Standard Operating Procedure MA229SOP, Product Inoculation”

25.13	USP current version, Section 7 of General Notices

25.14	Webster, Joan and Radke, Elinor and George, Narelle and Faoagali, Joan and Harris, Mary (2005) Barrier properties and cost implications of a single versus a double wrap for storing sterile instrument packs.  American Journal of Infection Control 33(6): pp. 348-352.

25.15	Form

Protocol Deviation Form FRM020-012

26.0	Post-Execution Approvals

The execution of this document has been reviewed and approved by the following individuals.  By approving this document, each individual acknowledges that the execution process complied with the protocol requirements and spirit of the Stimlabs Process and Equipment Validation Procedure QSP020-001.

| Sean Miller Stimlabs Validation/Study Leader |  | Date |  | John Daniel Stimlabs Tissue Bank Director |  | Date |
| --- | --- | --- | --- | --- | --- | --- |
| Melissa Gaynor Stimlabs Head of Quality |  | Date |  | Matt Nystrom Stimlabs Operations Management |  | Date |
| Mike Corby WuXi AppTec Senior Technical Director |  | Date |  | Richard Nardini WuXi AppTec Quality Assurance Manager |  | Date |

Appendix 1 – Discrepancy Log*

| Section | Description/Comments | Corrective Action | Completed By: Initial/Date |
| --- | --- | --- | --- |

*This logsheet can be copied in the event the number of discrepancies exceeds the fields provided on this page.

Appendix 2 – Maximum Load Configuration MDL1 (Sterilizer Shelf #1)

Note:	All items are laid flat upon the sterilizer rack

Appendix 3 – Maximum Load Configuration MDL1 (Sterilizer Shelf #2)

Note:	All items are laid flat upon the sterilizer rack

Appendix 4 – Maximum Load Configuration MDL1 (Sterilizer Shelf #3)

Note:	All items are laid flat upon the sterilizer rack

Appendix 5 – Maximum Load Configuration MDL1 (Sterilizer Shelf #4)

Note:	All items are laid flat upon the sterilizer rack

Appendix 6 – Maximum Load Configuration PDL1 (Sterilizer Shelf #1)

Note:	All items are laid on their side flat upon the sterilizer rack

Appendix 7 – Maximum Load Configuration PDL1 (Sterilizer Shelf #2)

Note:	All items are laid flat upon the sterilizer rack

Appendix 8 – Maximum Load Configuration PDL1 (Sterilizer Shelf #3)

Note:	All items are laid on their side flat upon the sterilizer rack

Appendix 9 – Maximum Load Configuration PDL1 (Sterilizer Shelf #4)

Note:	All items are laid flat upon the sterilizer rack

Appendix 10 – Minimum Load Configuration MDL1 (1 Rack, Sterilizer Shelf #4 Position)

Note:	All items are laid flat upon the sterilizer rack

Appendix 11 – Minimum Load Configuration PDL1 (1 Rack, Sterilizer Shelf #4 Position)

Note:	All items are laid on their side flat upon the sterilizer rack

Appendix 12 – Laboratory Item Showing BI Inoculum Locations (PCD)

Note:	BI inoculate surface locations on outer edge and center of silicone mat.

Artisan Metal Works Silicone Baking Mat, Part Number T3602MATA

Appendix 13 – Laboratory Item Showing BI Inoculum Locations (PCD)

Note:	BI inoculate location between the scissor’s pivoted blades.

Integra Stainless Steel Scissor, Part Number V95-26

Appendix 14 – Laboratory Item Showing BI Inoculum Locations (PCD)

Note:	BI inoculate location between the scissor’s pivoted blades.

Vetmed USA Mayo Scissor, Part Number 26-611

Appendix 15 – Laboratory Item Showing BI Inoculum Locations (PCD)

Note:	BI inoculate locations between the perforating head and head shaft assembly.

Valve Solutions Stainless Steel Perforating Rollers, Part Number N/A

Appendix 16 – Laboratory Item Showing BI Inoculum Locations (PCD)

Note:	BI inoculate locations are within the inside pockets of the quilted cleaning pad.

Contec Polyester EasyReach Quilted Cleaning Pads, Part Number MEQT0001

Appendix 17 – Laboratory Item Showing BI Inoculum Locations (PCD)

Note:	Surface BI inoculate locations on the piece of polypropylene cutting board with the knife grooved side.

2” x 4” x 0.25” section from a Small Parts Polypropylene Cutting Board, Part Number PPST-0250-F

Appendix 18 – Empty Chamber Temperature Sensors Locations (Sterilizer Shelf #1)

Appendix 19 – Empty Chamber Temperature Sensors Locations (Sterilizer Shelf #2)

Appendix 20 – Empty Chamber Temperature Sensors Locations (Sterilizer Shelf #3)

Appendix 21 – Empty Chamber Temperature Sensors Locations (Sterilizer Shelf #4)

Appendix 22 – Load Configuration MDL1 Maximum Heat Penetration Run #1 Temperature Sensor Locations (Sterilizer Shelf #1)

Appendix 23 – Load Configuration MDL1 Maximum Heat Penetration Run #1 Temperature Sensor Locations (Sterilizer Shelf #2)

Appendix 24 – Load Configuration MDL1 Maximum Heat Penetration Run #1 Temperature Sensor Locations (Sterilizer Shelf #3)

Appendix 25 – Load Configuration MDL1 Maximum Heat Penetration Run #1 Temperature Sensor Locations (Sterilizer Shelf #4)

Appendix 26 – Load Configuration MDL1 Maximum Heat Penetration Run #2 Temperature Sensor Locations (Sterilizer Shelf #1)

Appendix 27 – Load Configuration MDL1 Maximum Heat Penetration Run #2 Temperature Sensor Locations (Sterilizer Shelf #2)

Appendix 28 – Load Configuration MDL1 Maximum Heat Penetration Run #2 Temperature Sensor Locations (Sterilizer Shelf #3)

Appendix 29 – Load Configuration MDL1 Maximum Heat Penetration Run #2 Temperature Sensor Locations (Sterilizer Shelf #4)

Appendix 30 – Load Configuration MDL1 Maximum Heat Penetration Run #3 Temperature Sensor Locations (Sterilizer Shelf #1)

Appendix 31 – Load Configuration MDL1 Maximum Heat Penetration Run #3 Temperature Sensor Locations (Sterilizer Shelf #2)

Appendix 32 – Load Configuration MDL1 Maximum Heat Penetration Run #3 Temperature Sensor Locations (Sterilizer Shelf #3)

Appendix 33 – Load Configuration MDL1 Maximum Heat Penetration Run #3 Temperature Sensor Locations (Sterilizer Shelf #4)

Appendix 34 – Load Configuration PDL1 Maximum Heat Penetration Run #1 Temperature Sensor Locations (Sterilizer Shelf #1)

Appendix 35 – Load Configuration PDL1 Maximum Heat Penetration Run #1 Temperature Sensor Locations (Sterilizer Shelf #2)

Appendix 36 – Load Configuration PDL1 Maximum Heat Penetration Run #1 Temperature Sensor Locations (Sterilizer Shelf #3)

Appendix 37 – Load Configuration PDL1 Maximum Heat Penetration Run #1 Temperature Sensor Locations (Sterilizer Shelf #4)

Appendix 38 – Load Configuration PDL1 Maximum Heat Penetration Run #2 Temperature Sensor Locations (Sterilizer Shelf #1)

Appendix 39 – Load Configuration PDL1 Maximum Heat Penetration Run #2 Temperature Sensor Locations (Sterilizer Shelf #2)

Appendix 40 – Load Configuration PDL1 Maximum Heat Penetration Run #2 Temperature Sensor Locations (Sterilizer Shelf #3)

Appendix 41 – Load Configuration PDL1 Maximum Heat Penetration Run #2 Temperature Sensor Locations (Sterilizer Shelf #4)

Appendix 42 – Load Configuration PDL1 Maximum Heat Penetration Run #3 Temperature Sensor Locations (Sterilizer Shelf #1)

Appendix 43 – Load Configuration PDL1 Maximum Heat Penetration Run #3 Temperature Sensor Locations (Sterilizer Shelf #2)

Appendix 44 – Load Configuration PDL1 Maximum Heat Penetration Run #3 Temperature Sensor Locations (Sterilizer Shelf #3)

Appendix 45 – Load Configuration PDL1 Maximum Heat Penetration Run #3 Temperature Sensor Locations (Sterilizer Shelf #4)

Appendix 46 – Dry Time Verification/Calibration

Maximum Load Configuration MDL1

| Run | Cycle Type | Sterilizer Temp. Setting | Document Program Dry Time* | Did Dry Time Meet Weight and Visual Inspection Requirements? | Add 5 Minutes to Dry Time Program? |
| --- | --- | --- | --- | --- | --- |
| 1 | Pre-Vacuum | 132 C |  | Yes No | Yes No |
| 2 | Pre-Vacuum | 132 C |  | Yes No | Yes No |
| 3 | Pre-Vacuum | 132 C |  | Yes No | Yes No |
| PQ Executor’s Initials: |  |  |  | Date: |  |

Maximum Load Configuration PDL1

| Run | Cycle Type | Sterilizer Temp. Setting | Document Program Dry Time* | Did Dry Time Meet Weight and Visual Inspection Requirements? | Add 5 Minutes to Dry Time Program? |
| --- | --- | --- | --- | --- | --- |
| 1 | Pre-Vacuum | 121 C |  | Yes No | Yes No |
| 2 | Pre-Vacuum | 121 C |  | Yes No | Yes No. |
| 3 | Pre-Vacuum | 121 C |  | Yes No | Yes No |
| PQ Executor’s Initials: |  |  |  | Date: |  |

*The initial dry time is 20 minutes.

Note:	This appendix can be copied in the event the number of three (3) minimum successful heat penetration study cycle runs does not demonstrate adequate dry time.

Appendix 47 – Load Configuration MDL1 Minimum Heat Penetration Run #1 Temperature Sensor Locations (Sterilizer Shelf #4)

Appendix 48 – Load Configuration MDL1 Minimum Heat Penetration Run #2 Temperature Sensor Locations (Sterilizer Shelf #4)

Appendix 49 – Load Configuration MDL1 Minimum Heat Penetration Run #3 Temperature Sensor Locations (Sterilizer Shelf #4)

Appendix 50 – Load Configuration PDL1 Minimum Heat Penetration Run #1 Temperature Sensor Locations (Sterilizer Shelf #4)

Appendix 51 – Load Configuration PDL1 Minimum Heat Penetration Run #2 Temperature Sensor Locations (Sterilizer Shelf #4)

Appendix 52 – Load Configuration PDL1 Minimum Heat Penetration Run #3 Temperature Sensor Locations (Sterilizer Shelf #4)

Appendix 53 – Load Configuration MDL1 Maximum Micro Challenge Run #1 Temperature Sensor, PCD, and BI Locations (Sterilizer Shelf #1)

Appendix 54 – Load Configuration MDL1 Maximum Micro Challenge Run #1 Temperature Sensor, PCD, and BI Locations (Sterilizer Shelf #2)

Appendix 55 – Load Configuration MDL1 Maximum Micro Challenge Run #1 Temperature Sensor, PCD, and BI Locations (Sterilizer Shelf #3)

Appendix 56 – Load Configuration MDL1 Maximum Micro Challenge Run #1 Temperature Sensor, PCD, and BI Locations (Sterilizer Shelf #4)

Appendix 57 – Load Configuration MDL1 Maximum Micro Challenge Run #2 Temperature Sensor, PCD, and BI Locations (Sterilizer Shelf #1)

Appendix 58 – Load Configuration MDL1 Maximum Micro Challenge Run #2 Temperature Sensor, PCD, and BI Locations (Sterilizer Shelf #2)

Appendix 59 – Load Configuration MDL1 Maximum Micro Challenge Run #2 Temperature Sensor, PCD, and BI Locations (Sterilizer Shelf #3)

Appendix 60 – Load Configuration MDL1 Maximum Micro Challenge Run #2 Temperature Sensor, PCD, and BI Locations (Sterilizer Shelf #4)

Appendix 61 – Load Configuration MDL1 Maximum Micro Challenge Run #3 Temperature Sensor, PCD, and BI Locations (Sterilizer Shelf #1)

Appendix 62 – Load Configuration MDL1 Maximum Micro Challenge Run #3 Temperature Sensor, PCD, and BI Locations (Sterilizer Shelf #2)

Appendix 63 – Load Configuration MDL1 Maximum Micro Challenge Run #3 Temperature Sensor, PCD, and BI Locations (Sterilizer Shelf #3)

Appendix 64 – Load Configuration MDL1 Maximum Micro Challenge Run #3 Temperature Sensor, PCD, and BI Locations (Sterilizer Shelf #4)

Appendix 65 – Load Configuration PDL1 Maximum Micro Challenge Run #1 Temperature Sensor, PCD, and BI Locations (Sterilizer Shelf #1)

Appendix 66 – Load Configuration PDL1 Maximum Micro Challenge Run #1 Temperature Sensor, PCD, and BI Locations (Sterilizer Shelf #2)

Appendix 67 – Load Configuration PDL1 Maximum Micro Challenge Run #1 Temperature Sensor, PCD, and BI Locations (Sterilizer Shelf #3)

Appendix 68 – Load Configuration PDL1 Maximum Micro Challenge Run #1 Temperature Sensor, PCD, and BI Locations (Sterilizer Shelf #4)

Appendix 69 – Load Configuration PDL1 Maximum Micro Challenge Run #2 Temperature Sensor, PCD, and BI Locations (Sterilizer Shelf #1)

Appendix 70 – Load Configuration PDL1 Maximum Micro Challenge Run #2 Temperature Sensor, PCD, and BI Locations (Sterilizer Shelf #2)

Appendix 71 – Load Configuration PDL1 Maximum Micro Challenge Run #2 Temperature Sensor, PCD, and BI Locations (Sterilizer Shelf #3)

Appendix 72 – Load Configuration PDL1 Maximum Micro Challenge Run #2 Temperature Sensor, PCD, and BI Locations (Sterilizer Shelf #4)

Appendix 73 – Load Configuration PDL1 Maximum Micro Challenge Run #3 Temperature Sensor, PCD, and BI Locations (Sterilizer Shelf #1)

Appendix 74 – Load Configuration PDL1 Maximum Micro Challenge Run #3 Temperature Sensor, PCD, and BI Locations (Sterilizer Shelf #2)

Appendix 75 – Load Configuration PDL1 Maximum Micro Challenge Run #3 Temperature Sensor, PCD, and BI Locations (Sterilizer Shelf #3)

Appendix 76 – Load Configuration PDL1 Maximum Micro Challenge Run #3 Temperature Sensor, PCD, and BI Locations (Sterilizer Shelf #4)

Appendix 77 – Load Configuration MDL1 Minimum Micro Challenge Run #1 Temperature Sensor, PCD, and BI Locations (Sterilizer Shelf #4)

Appendix 78 – Load Configuration MDL1 Minimum Micro Challenge Run #2 Temperature Sensor, PCD, and BI Locations (Sterilizer Shelf #4)

Appendix 79 – Load Configuration MDL1 Minimum Micro Challenge Run #3 Temperature Sensor, PCD, and BI Locations (Sterilizer Shelf #4)

Appendix 80 – Load Configuration PDL1 Minimum Micro Challenge Run #1 Temperature Sensor, PCD, and BI Locations (Sterilizer Shelf #4)

Appendix 81 – Load Configuration PDL1 Minimum Micro Challenge Run #2 Temperature Sensor, PCD, and BI Locations (Sterilizer Shelf #4)

Appendix 82 – Load Configuration PDL1 Minimum Micro Challenge Run #3 Temperature Sensor, PCD, and BI Locations (Sterilizer Shelf #4)

Appendix 83 – Weight Confirmation Check Sheet (Max MDL1 Load Configuration)

| Shelf # | Item Description / Cycle Run Number | Gram Weight Before Sterilization | Gram Weight after Sterilization | ≤ +3% Weight Difference? |
| --- | --- | --- | --- | --- |
| 1 |  |  |  | Yes No |
| 2 |  |  |  | Yes No |
| 3 |  |  |  | Yes No |
| 4 |  |  |  | Yes No |
| 1 |  |  |  | Yes No |
| 2 |  |  |  | Yes No |
| 3 |  |  |  | Yes No |
| 4 |  |  |  | Yes No |
| 1 |  |  |  | Yes No |
| 2 |  |  |  | Yes No |
| 3 |  |  |  | Yes No |
| 4 |  |  |  | Yes No |
| 1 |  |  |  | Yes No |
| 2 |  |  |  | Yes No |
| 3 |  |  |  | Yes No |
| 4 |  |  |  | Yes No |
| Balance Identification Number: |  |  | Due Date: |  |
| PQ Executor’s Initials: |  |  | Date: |  |

Note:	This appendix can be copied in the event the number of three (3) minimum successful heat penetration study cycle runs does not demonstrate adequate dry time.

Appendix 84 – Weight Confirmation Check Sheet (Max PDL1 Load Configuration)

| Shelf # | Item Description / Cycle Run Number | Gram Weight Before Sterilization | Gram Weight after Sterilization | ≤ +3% Weight Difference? |
| --- | --- | --- | --- | --- |
| 1 |  |  |  | Yes No |
| 2 |  |  |  | Yes No |
| 3 |  |  |  | Yes No |
| 4 |  |  |  | Yes No |
| 1 |  |  |  | Yes No |
| 2 |  |  |  | Yes No |
| 3 |  |  |  | Yes No |
| 4 |  |  |  | Yes No |
| 1 |  |  |  | Yes No |
| 2 |  |  |  | Yes No |
| 3 |  |  |  | Yes No |
| 4 |  |  |  | Yes No |
| 1 |  |  |  | Yes No |
| 2 |  |  |  | Yes No |
| 3 |  |  |  | Yes No |
| 4 |  |  |  | Yes No |
| Balance Identification Number: |  |  | Due Date: |  |
| PQ Executor’s Initials: |  |  | Date: |  |

Note:	This appendix can be copied in the event the number of three (3) minimum successful heat penetration study cycle runs does not demonstrate adequate dry time.

- PURPOSE

- SCOPE

- BACKGROUND

- REFERENCES

  - Food and Drug Administration 21 CFR Part 1271

  - American Association for Tissue Banks Standards for Tissue Banking

  - Equipment Manual

  - FDA Guidance for Industry – Process Validation: General Principles and Practices

  - GHTF/SG3/N99-10:2004 Quality Management Systems – Process Validation Guidance

  - Process Validation Plan (PVP _________________)

- DEFINITIONS

  - Acceptance Criteria:  Pre-determined requirements upon which pass/fail decisions are based.  Acceptance criteria are specific and measureable.

  - Critical to Quality:  Attribute of a part, assembly, sub-assembly, product or process that has direct and significant impact on its actual or perceived quality to the customer or end user.

  - Deviation:  The action of departing from an established method, procedure, or acceptance criteria.

  - LSL:  Lower Specification Limit; used to determine process capability.

  - Normal Distribution:  A function that represents the distribution of many random variables as a symmetrical, bell-shaped graph.

  - Performance Qualification:  Establishing by objective evidence the process control limits and action levels which result in a product that meets predetermined requirements.

  - PpK/CpK:  Process capability; an estimate of the process capability of a process during its initial set-up before it has been brought into a state of statistical control.

  - Process Validation:  A mechanism to ensure that a process whose output is not fully verifiable by inspection and test is capable of consistently (repeatedly and reliably) producing a result or product that meets pre-determined requirements.  Process validation proves that a process is under control and is used to determine the process variables, acceptable limits for those variables, and to establish appropriate in-process controls.

  - Protocol:  A pre-approved and prospective document that states the requirements and procedures for conducting a validation.

  - USL:  Upper specification limit; used in determining process capability.

  - Validation/Study Leader: The individual who has the responsibility and authority over the operation of the equipment or process being validated.

  - Verification:  Confirmation by examination with a provision of documented evidence that a specific requirement has been fulfilled.

- RESPONSIBILITIES

| Function | Author Protocol | Review Protocol | Execute Protocol | Execute Testing |
| --- | --- | --- | --- | --- |
| Validation/Study Leader |  |  |  |  |
| Tissue Bank Director |  |  |  |  |
| Head of Quality |  |  |  |  |
| Operations Management |  |  |  |  |
| Other |  |  |  |  |
| Other |  |  |  |  |
| Other |  |  |  |  |

- ENVIRONMENTAL REQUIREMENTS

  - Environmental Condition Ranges

  - Operator Variability

- PROCEDURE

  - Manufacturing Equipment Description

    - [Vendor Name]

    - [Serial Number]

    - [Model Number]

    - [Location]

  - Raw Materials

    - Identification of Materials

  - Critical to Quality (CTQ) Attributes

    - Variable Attributes

    - Attribute Data (Pass/Fail)

  - Process Variables

    - Critical Machine Settings & Ranges

    - Raw Material Variability

    - Identification of Process Window (as determined by the OQ)

    - Describe each of the PQ Runs to be run

      - Run #1

      - Run #2

      - Run #3

  - Test Equipment

  - Test Methods

- LOT SIZE & SAMPLING PLAN

  - Lot Size

  - Sampling Plan & Statistical Rationale

- ACCEPTANCE CRITERIA

  - Acceptance Criteria for Success (Process Capability, Test Results, etc.)

- MATERIAL DISPOSITION

- DOCUMENTATION REQUIREMENTS

  - Results shall be summarized in an PQ Report.

  - Various Control Plans, including validated process settings for critical process parameters and the re-validation criteria shall be documented in the final report.

  - Raw data sheets shall be included in the appendix of the PQ report and/or a reference to the location of the raw data shall be included in the PQ report.

  - Signed PQ Protocol Training forms will be attached to the PQ report (in the Appendix).

- DEVIATIONS

  - Deviations to the PQ will be pre-approved using FRM020-012 Protocol Deviation Form and explained in the PQ report.

  - Copies of signed Protocol Deviation Forms will be attached to the PQ report (in the Appendix).

- PRE-REQUISITE STEPS

  - Approved Process Validation Plan (PVP), IQ, and OQ Reports (if required) or justification for omission.

  - Calibration certificates for equipment and measurement devices (gauges)

  - Approval and release of raw materials

  - Verification of environmental requirements

  - Completed Operator Training

  - Work instructions

  - The PQ Protocol must be signed before validation activities may begin.

- APPENDIX

  - Blank training form (or use approved form)

  - Blank data collection sheets