- PURPOSE

  - The purpose of this validation is to provide individuals who’s responsibility is to approve the performance qualification (PQ) protocol that did not approve REV 01 for pouch sealers to demonstrate that the seal applied to inner and outer pouches meets design specifications when processed at nominal parameters of the pouch sealing procedure.

  - This validation is a deliverable for PCR21-025 New Cleanroom Construction and Support Lab Expansion.

- SCOPE

  - This validation applies to the Van der Stahl heat sealer.  A heat sealer is used to seal inner and outer product pouches.  This validation includes the qualification of the heat sealer equipment and heat sealing process.  The inner pouch will maintain a sterile barrier.  Sterility maintenance was validated under VAP/VAR16-002.

  - This equipment is located in the new cleanroom at the StimLabs Roswell, GA facility.

  - Inner and outer pouch are constructed of different material, thus necessitating a different set of parameters to achieve an acceptable seal.  The sealers will be validated to define the process parameters for sealing both inner and outer pouches using the validated process setting.

- BACKGROUND

  - StimLabs is a tissue bank involved in the processing of human derived tissues into products for transplantation.  In order to provide the tissue products in a manner that complies with sterile transplantation procedures, StimLabs packages the tissue product in a dual pouch configuration, including both an inner and outer pouch.

  - To ensure the pouch sealer is fit for the intended purpose, StimLabs will perform an performance qualification and continue to monitor the performance of the pouch sealer while it is being used to seal product pouches.

- REFERENCES

  - Food and Drug Administration 21 CFR Part 1271

  - American Association for Tissue Banks Standards for Tissue Banking

  - Van Der Stahl Heat Sealer Equipment Manual

  - FDA Guidance for Industry – Process Validation: General Principles and Practices

  - GHTF/SG3/N99-10:2004 Quality Management Systems – Process Validation Guidance

  - FRM017-003 Seal Integrity Peel Test Form

  - SOP017-003 Physical Testing of Heat Deals Test Method

  - VAP21-014-01 Van Der Stahl Heat Sealer Installation Qualification

  - VAP21-014-02 Van Der Stahl Heat Sealer Operational Qualification

- DEFINITIONS

  - Acceptance Criteria:  Pre-determined requirements upon which pass/fail decisions are based.  Acceptance criteria are specific and measureable.

  - Critical to Quality:  Attribute of a part, assembly, sub-assembly, product or process that has direct and significant impact on its actual or perceived quality to the customer or end user.

  - Deviation:  The action of departing from an established method, procedure, or acceptance criteria.

  - LSL:  Lower Specification Limit; used to determine process capability.

  - Normal Distribution:  A function that represents the distribution of many random variables as a symmetrical, bell-shaped graph.

  - Performance Qualification:  Establishing by objective evidence the process control limits and action levels which result in a product that meets predetermined requirements.

  - PpK/CpK:  Process capability; an estimate of the process capability of a process during its initial set-up before it has been brought into a state of statistical control.

  - Process Validation:  A mechanism to ensure that a process whose output is not fully verifiable by inspection and test is capable of consistently (repeatedly and reliably) producing a result or product that meets pre-determined requirements.  Process validation proves that a process is under control and is used to determine the process variables, acceptable limits for those variables, and to establish appropriate in-process controls.

  - Protocol:  A pre-approved and prospective document that states the requirements and procedures for conducting a validation.

  - USL:  Upper specification limit; used in determining process capability.

  - Validation/Study Leader: The individual who has the responsibility and authority over the operation of the equipment or process being validated.

  - Verification:  Confirmation by examination with a provision of documented evidence that a specific requirement has been fulfilled.

- RESPONSIBILITIES

| Function | Name | Author Protocol | Review Protocol | Execute Protocol | Execute Testing |
| --- | --- | --- | --- | --- | --- |
| Validation/Study Leader | Michael Baird | X | X |  |  |
| MwER/MR/HOQ/TBD | Melissa O’Connor |  | X |  |  |
| Operations Management | Michael Baird |  | X |  |  |
| Quality Management | Stacy Harmon |  | X |  |  |
| Other | Rob Frye | X |  |  |  |
| Other | Trained Technician |  |  | X |  |
| Testing Lab | Nelson Laboratories |  |  |  | X |

- ENVIRONMENTAL REQUIREMENTS

  - Environmental Condition Ranges

    - N/A

  - Operator Variability

    - Three (3) Operators will be performing this validation using alternating heat sealer cycles.

- PROCEDURE

  - Manufacturing Equipment Description

    - Manufacturer: Van der Stahl

    - Serial number: will be identified in the PQ data sheets

    - Model Number: MS350

    - Location: New Cleanroom

  - Raw Materials

    - Inner Pouches

      - Sizes: 4.5” x 6.75” and 3.5” x 5.0”

      - Steriflex-101 (48g PET/0.002LDPE)

      - Steriflex-112 (48g PET/0.002LDPE Peelable)

    - Outer Pouches

      - Size: 6”x 9.25”

      - SteriFlex 301

      - SteriFlex 301-P

  - Process Variables

    - Nominal temperature and seal time setting will be evaluated.  Following is a list of nominal conditions during normal operations which are within the validated range established during the OQ:

| Outer Pouch |  |  | 4.5” x 6.75” Inner Pouch |  |  | 3.5” x 5” Inner Pouch |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Factor | Units | Nominal (N) | Factor | Units | Nominal (N) | Factor | Units | Nominal (N) |
| Sealing Temp. | F° | 325 | Sealing Temp. | F° | 285 | Sealing Temp. | F° | 285 |
| Dwell Time | Sec. | 3.5 | Dwell Time | Sec. | 3.5 | Dwell Time | Sec. | 3.5 |
| Cooling Temp. | F° | 150 | Cooling Temp. | F° | 150 | Cooling Temp. | F° | 150 |

    - The following describes the method to be used for producing and testing the test units.

      - Seal the samples per standard Amniotic Membrane processing procedure .

      - Three runs at nominal setting will be performed by three different alternating operators.  Reference the schematic below for an example of operator alternation:

| Trial Number | Performed By |
| --- | --- |
| 1 | Operator 1 |
| 2 | Operator 2 |
| 3 | Operator 3 |
| 4 | Operator 1 |
| 5 | Operator 2 |
| 6 | Operator 3 |
| 7 | Operator 1 |
| 8 | Operator 2 |
| 9 | Operator 3 |

      - Label each Pouch with parameter Trial #, Test Method, and “VAP21-014-03”

      - Record pouch information and sealing process parameters within the test report.

| Outer Pouch |  |  | 4.5” x 6.75” Inner Pouch |  |  | 3.5” x 5” Inner Pouch |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Trial # | Condition | # of Sample | Trial # | Condition | # of Sample | Trial # | Condition | # of Sample |
| 1 | N | 89 | 4 | N | 89 | 7 | N | 89 |
| 2 | N | 89 | 5 | N | 89 | 8 | N | 89 |
| 3 | N | 89 | 6 | N | 89 | 9 | N | 89 |

  - Test Equipment

    - Van der Stahl PPT-100V Simple Peel Pouch Tester

    - 178 - 4.5” x 6.75” inner pouches

    - 178 - 3.5” x 5” inner pouches

    - 178 – outer pouches

  - Test Methods

    - Pouch Peel Strength Test

      - A peel test is performed by a trained StimLabs technician. (30) sealed pouches from each trial will be tested.  Note and record all results a Seal Integrity Peel Test Form.

      - During the peel test process, observe peeled samples for sign of fiber tear. Note and record all results on the associated datasheet.

    - Bubble Emission Test

      - A bubble emissions test is performed by Nelson Laboratories per ASTM Standard F 2096-04 on (59) sealed pouches from each trial.

      - This test method applies to the detection of gross leaks in nonporous tray and pouch packages. Record results using the associated datasheet.

- LOT SIZE & SAMPLING PLAN

  - Sampling Plan & Statistical Rationale

    - Variable Data (95/95):  A minimum of thirty (30) samples will be pulled from each lot for peel strength. A sample size of thirty will make it possible to reliably determine if the data collected can be represented by the normal distribution. An accurate assessment of data normality is necessary to choose the proper statistical analysis method.

    - Attribute Data (95/95):  A sample size of at least fifty-nine (59), with zero non-conformances, provides a 95% confidence level that 95% of the population is within specification. This sample size is based on a nonparametric attribute plan derived from the binomial distribution.

- ACCEPTANCE CRITERIA

  - Acceptance Criteria for Success

    - Variable data:  All test samples must be accepted for seal verification to pass. Seal verification passes if all results are ≥ 1.41 lb/in.

    - Variable Data: Seal width -  minimum 0.25ʺ

    - Attribute data:  All test samples must be accepted for visual inspection to pass.

- MATERIAL DISPOSITION

  - All samples produced for this PQ will be consumed during testing or discarded.

- DOCUMENTATION REQUIREMENTS

  - Results shall be summarized in an PQ Report.

  - Various Control Plans, including validated process settings for critical process parameters and the re-validation criteria shall be documented in the final report.

  - Raw data sheets shall be included in the appendix of the PQ report and/or a reference to the location of the raw data shall be included in the PQ report.

  - Signed PQ Protocol Training forms will be attached to the PQ report (in the Appendix).

- DEVIATIONS

  - Deviations to the PQ will be pre-approved using FRM020-012 Protocol Deviation Form and explained in the PQ report.

  - Copies of signed Protocol Deviation Forms will be attached to the PQ report (in the Appendix).

- PRE-REQUISITE STEPS

  - Approved Process Validation Plan (PVP), IQ, and OQ Reports (if required) or justification for omission.

  - Calibration certificates for equipment and measurement devices (gauges)

  - Approval and release of raw materials

  - Verification of environmental requirements

  - Completed Operator Training

  - Work instructions

  - The PQ Protocol must be signed before validation activities may begin.

- APPENDIX

  - Blank data collection sheets

APPENDIX 16.1 – DATA COLLECTION SHEETS

STOP!  Read Instructions.

Instructions:

- Complete one data collection sheet for each Van der Stahl Heat Sealer.

- Verify the Pre-requisite criteria have been met on this data collection sheet prior to testing.  Note that the Pre-requisite criteria only need to be verified once for the validation.

- For attribute data: Circle “Pass” or “Fail” or write the words “Pass” or “Fail”

- For variable data: Record actual results in the units specified or add units if the units are not specified.

- Verification must include signature of the individual performing the step and the date performed.  Verification must be completed concurrently with performance.

- Follow Good Documentation Practice (GDP).

Pre-requisite Criteria (must be complete and verified prior to start of IQ activities):

| Pre-requisite | Complete (Circle One) | Verification Signature/Date |
| --- | --- | --- |
| Process Validation Plan is approved. | Yes No |  |
| IQ Protocol is completed | Yes No |  |
| OQ Protocol is completed | Yes No |  |
| PQ Protocol is approved | Yes No |  |
| Unit is Calibrated | Yes No |  |
| Operator Training is complete and documented. | Yes No |  |

Pre-Requisites Verified By (signature & date): _____________________________________________________________

Quality Reviewed By (signature & date): ______________________________________________________________

Section A – Equipment Setup

| Equipment Name | Model Number | Manufacturer | Serial Number | Stimlabs ID Number | Verification Signature/Date |
| --- | --- | --- | --- | --- | --- |

Section A Completed By (signature & date): ________________________________________________________________

Quality Reviewed By (signature & date): ______________________________________________________________

Section B – Trial 1 Outer Pouch – Operator 1

| Specification or Acceptance Criteria | Reading/Results | Result | Verification Signature/Date |
| --- | --- | --- | --- |
| FRM017-003 will be filled out peel strength test and trial number will be documented on the form |  |  |  |
| Operating Technician | Print Name: | N/A |  |
| Lot number of outer pouch |  | N/A |  |
| Set Seal Temperature. 325℉ | N/A | Pass / Fail |  |
| Set Dwell Time. 3.5 second | N/A | Pass / Fail |  |
| Set Cooling Temperature. 150℉ | N/A | Pass / Fail |  |
| 89 Outer Pouches sealed and labeled for appropriate test. 30 – Seal Peel Strength 59 – Bubble Emission | Total Pouches Sealed & Width Measured | Pass / Fail |  |
|  | Total Pouches Sealed and labeled for Peel Strength ________________ | Pass / Fail |  |
|  | Total Pouches Sealed and labeled for Bubble Emission _______________ | Pass / Fail |  |

Section B Completed By (signature & date): ________________________________________________________________

Quality Reviewed By (signature & date): ______________________________________________________________

Section C – Trial 2 Outer Pouch Operator 2

| Specification or Acceptance Criteria | Reading/Results | Result | Verification Signature/Date |
| --- | --- | --- | --- |
| FRM017-003 will be filled out peel strength test and trial number will be documented on the form |  |  |  |
| Operating Technician | Print Name: | N/A |  |
| Lot number of outer pouch |  | N/A |  |
| Set Seal Temperature. 325℉ | N/A | Pass / Fail |  |
| Set Dwell Time. 3.5 second | N/A | Pass / Fail |  |
| Set Cooling Temperature. 150℉ | N/A | Pass / Fail |  |
| 89 Outer Pouches sealed and labeled for appropriate test. 30 – Seal Peel Strength 59 – Bubble Emission | Total Pouches Sealed & Width Measured _________________ | Pass / Fail |  |
|  | Total Pouches Sealed and labeled for Peel Strength ________________ | Pass / Fail |  |
|  | Total Pouches Sealed and labeled for Bubble Emission ________________ | Pass / Fail |  |

Section C Completed By (signature & date): ________________________________________________________________

Quality Reviewed By (signature & date): ______________________________________________________________

Section D – Trial 3 Outer Pouch Operator 3

| Specification or Acceptance Criteria | Reading/Results | Result | Verification Signature/Date |
| --- | --- | --- | --- |
| FRM017-003 will be filled out peel strength test and trial number will be documented on the form |  |  |  |
| Operating Technician | Print Name: | N/A |  |
| Lot number of outer pouch |  | N/A |  |
| Set Seal Temperature. 325℉ | N/A | Pass / Fail |  |
| Set Dwell Time. 3.5 second | N/A | Pass / Fail |  |
| Set Cooling Temperature. 150℉ | N/A | Pass / Fail |  |
| 89 Outer pouches sealed and labeled for appropriate test. 30 – Seal Peel Strength 59 – Bubble Emission | Total Pouches Sealed & Width Measured | Pass / Fail |  |
|  | Total Pouches Sealed and labeled for Peel Strength ________________ | Pass / Fail |  |
|  | Total Pouches Sealed and labeled for Bubble Emission ________________ | Pass / Fail |  |

Section D Completed By (signature & date): ________________________________________________________________

Quality Reviewed By (signature & date): ______________________________________________________________

Section E – Trial 4 4.5” x 6.75” inner pouches Operator 1

| Specification or Acceptance Criteria | Reading/Results | Result | Verification Signature/Date |
| --- | --- | --- | --- |
| FRM017-003 will be filled out peel strength test and trial number will be documented on the form |  |  |  |
| Operating Technician | Print Name: | N/A |  |
| Lot number of 4.5” x 6.75” inner pouches |  | N/A |  |
| Set Seal Temperature. 285℉ | N/A | Pass / Fail |  |
| Set Dwell Time. 3.5 second | N/A | Pass / Fail |  |
| Set Cooling Temperature. 150℉ | N/A | Pass / Fail |  |
| 89 4.5” x 6.75” inner pouches sealed and labeled for appropriate test. 30 – Seal Peel Strength 59 – Bubble Emission | Total Pouches Sealed & Width Measured _________________ | Pass / Fail |  |
|  | Total Pouches Sealed and labeled for Peel Strength ________________ | Pass / Fail |  |
|  | Total Pouches Sealed and labeled for Bubble Emission ________________ | Pass / Fail |  |

Section E Completed By (signature & date): ________________________________________________________________

Quality Reviewed By (signature & date): ______________________________________________________________

Section F – Trial 5 4.5” x 6.75” inner pouches Operator 2

| Specification or Acceptance Criteria | Reading/Results | Result | Verification Signature/Date |
| --- | --- | --- | --- |
| FRM017-003 will be filled out peel strength test and trial number will be documented on the form |  |  |  |
| Operating Technician | Print Name: | N/A |  |
| Lot number of 4.5” x 6.75” inner pouches |  | N/A |  |
| Set Seal Temperature. 285℉ | N/A | Pass / Fail |  |
| Set Dwell Time. 3.5 second | N/A | Pass / Fail |  |
| Set Cooling Temperature. 150℉ | N/A | Pass / Fail |  |
| 89 4.5” x 6.75” inner pouches sealed and labeled for appropriate test. 30 – Seal Peel Strength 59 – Bubble Emission | Total Pouches Sealed & Width Measured | Pass / Fail |  |
|  | Total Pouches Sealed and labeled for Peel Strength ________________ | Pass / Fail |  |
|  | Total Pouches Sealed and labeled for Bubble Emission ________________ | Pass / Fail |  |

Section F Completed By (signature & date): ________________________________________________________________

Quality Reviewed By (signature & date): ______________________________________________________________

Section G – Trial 6 4.5” x 6.75” inner pouches Operator 3

| Specification or Acceptance Criteria | Reading/Results | Result | Verification Signature/Date |
| --- | --- | --- | --- |
| FRM017-003 will be filled out peel strength test and trial number will be documented on the form |  |  |  |
| Operating Technician | Print Name: | N/A |  |
| Lot number of 4.5” x 6.75” inner pouches |  | N/A |  |
| Set Seal Temperature. 285℉ | N/A | Pass / Fail |  |
| Set Dwell Time. 3.5 second | N/A | Pass / Fail |  |
| Set Cooling Temperature. 150℉ | N/A | Pass / Fail |  |
| 89 4.5” x 6.75” inner pouches sealed and labeled for appropriate test. 30 – Seal Peel Strength 59 – Bubble Emission | Total Pouches Sealed & Width Measured | Pass / Fail |  |
|  | Total Pouches Sealed and labeled for Peel Strength ________________ | Pass / Fail |  |
|  | Total Pouches Sealed and labeled for Bubble Emission ________________ | Pass / Fail |  |

Section G Completed By (signature & date): ________________________________________________________________

Quality Reviewed By (signature & date): ______________________________________________________________

Section H – Trial 7 3.5” x 5” inner pouches Operator 1

| Specification or Acceptance Criteria | Reading/Results | Result | Verification Signature/Date |
| --- | --- | --- | --- |
| FRM017-003 will be filled out peel strength test and trial number will be documented on the form |  |  |  |
| Operating Technician | Print Name: | N/A |  |
| Lot number of 3.5” x 5” inner pouches |  | N/A |  |
| Set Seal Temperature. 285℉ | N/A | Pass / Fail |  |
| Set Dwell Time. 3.5 second | N/A | Pass / Fail |  |
| Set Cooling Temperature. 150℉ | N/A | Pass / Fail |  |
| 89 3.5” x 5” inner pouches sealed and labeled for appropriate test. 30 – Seal Peel Strength 59 – Bubble Emission | Total Pouches Sealed & Width Measured | Pass / Fail |  |
|  | Total Pouches Sealed and labeled for Peel Strength ________________ | Pass / Fail |  |
|  | Total Pouches Sealed and labeled for Bubble Emission ________________ | Pass / Fail |  |

Section H Completed By (signature & date): ________________________________________________________________

Quality Reviewed By (signature & date): ______________________________________________________________

Section I – Trial 8 3.5” x 5” inner pouches Operator 1

| Specification or Acceptance Criteria | Reading/Results | Result | Verification Signature/Date |
| --- | --- | --- | --- |
| FRM017-003 will be filled out peel strength test and trial number will be documented on the form |  |  |  |
| Operating Technician | Print Name: | N/A |  |
| Lot number of 3.5” x 5” inner pouches |  | N/A |  |
| Set Seal Temperature. 285℉ | N/A | Pass / Fail |  |
| Set Dwell Time. 3.5 second | N/A | Pass / Fail |  |
| Set Cooling Temperature. 150℉ | N/A | Pass / Fail |  |
| 89 3.5” x 5” inner pouches sealed and labeled for appropriate test. 30 – Seal Peel Strength 59 – Bubble Emission | Total Pouches Sealed & Width Measured | Pass / Fail |  |
|  | Total Pouches Sealed and labeled for Peel Strength ________________ | Pass / Fail |  |
|  | Total Pouches Sealed and labeled for Bubble Emission ________________ | Pass / Fail |  |

Section I Completed By (signature & date): ________________________________________________________________

Quality Reviewed By (signature & date): ______________________________________________________________

Section J – Trial 9 3.5” x 5” inner pouches Operator 2

| Specification or Acceptance Criteria | Reading/Results | Result | Verification Signature/Date |
| --- | --- | --- | --- |
| FRM017-003 will be filled out peel strength test and trial number will be documented on the form |  |  |  |
| Operating Technician | Print Name: | N/A |  |
| Lot number of 3.5” x 5” inner pouches |  | N/A |  |
| Set Seal Temperature. 285℉ | N/A | Pass / Fail |  |
| Set Dwell Time. 3.5 second | N/A | Pass / Fail |  |
| Set Cooling Temperature. 150℉ | N/A | Pass / Fail |  |
| 89 3.5” x 5” inner pouches sealed and labeled for appropriate test. 30 – Seal Peel Strength 59 – Bubble Emission | Total Pouches Sealed & Width Measured | Pass / Fail |  |
|  | Total Pouches Sealed and labeled for Peel Strength ________________ | Pass / Fail |  |
|  | Total Pouches Sealed and labeled for Bubble Emission ________________ | Pass / Fail |  |

Section J Completed By (signature & date): ________________________________________________________________

Quality Reviewed By (signature & date): ______________________________________________________________